## chl
Cruise #: | Date | LTER Station | Cast # | Niskin # | Time In | Time Out | Replicate | Vol Filt | Filter Size | Vol Extracted | Sample | 90% Acetone | Dilution During Reading | Chl_Cal_Filename | tau_Calibration | Fd_Calibration | Rb | Ra | blank | Rb-blank | Ra-blank | Chl (ug/l) | Phaeo (ug/l) | Cal_Date | Lab notebook and Page number | Date_analyzed | cruise_project_type | Personnel Filter | Personnel Read | Fluorometer | quality_flag | QC_d | Comments | Comments2
AR22 | 20170902 |  | 1 | 1 |  |  | a | 535 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 163.3 | 104.2 | 0.78917 | 162.5108 | 103.4108 | 1.2175 | 0.63207 | 20181101 | 9-58 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 1 | 1 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 355.8 | 222 | 0.66033 | 355.1397 | 221.3397 | 1.2768 | 0.62841 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 1 | 2 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 345.8 | 212.1 | 0.66033 | 345.1397 | 211.4397 | 1.2759 | 0.54415 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | "leak at beginning"
AR22 | 20170902 |  | 1 | 2 |  |  | b | 535 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 198.3 | 125.3 | 0.78917 | 197.5108 | 124.5108 | 1.5039 | 0.7231 | 20181101 | 9-58 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR22 | 20170902 |  | 1 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 414 | 248.5 | 0.66033 | 413.3397 | 247.8397 | 1.5793 | 0.55401 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 1 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 397.5 | 243.2 | 0.66033 | 396.8397 | 242.5397 | 1.4724 | 0.61527 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 1 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 304.3 | 184.6 | 0.66033 | 303.6397 | 183.9397 | 1.1423 | 0.44103 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 1 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 297.6 | 180.2 | 0.66033 | 296.9397 | 179.5397 | 1.1203 | 0.42511 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 1 | 4 |  |  | a | 535 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 249.8 | 149.3 | 0.87933 | 248.9207 | 148.4207 | 0.95905 | 0.31852 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication, was this pre-filtered correctly?
AR22 | 20170902 |  | 1 | 4 |  |  | b | 535 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 182.6 | 114.1 | 0.66033 | 181.9397 | 113.4397 | 0.65368 | 0.32278 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication
AR22 | 20170902 |  | 2 | 1 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 267.4 | 163.7 | 0.66033 | 266.7397 | 163.0397 | 0.98958 | 0.41382 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 2 | 1 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 263.3 | 160.7 | 0.66033 | 262.6397 | 160.0397 | 0.97909 | 0.39849 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 2 | 2 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 146.6 | 95.89 | 0.493 | 146.107 | 95.397 | 0.48391 | 0.33724 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | a?
AR22 | 20170902 |  | 2 | 2 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 166.3 | 104.5 | 0.66033 | 165.6397 | 103.8397 | 0.58974 | 0.30408 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 2 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 371.7 | 234.1 | 0.66033 | 371.0397 | 233.4397 | 1.3131 | 0.6963 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 2 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 384.7 | 239.3 | 0.66033 | 384.0397 | 238.6397 | 1.3875 | 0.66663 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 2 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 335 | 208.4 | 0.66033 | 334.3397 | 207.7397 | 1.2081 | 0.58005 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 2 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 330.9 | 208.5 | 0.66033 | 330.2397 | 207.8397 | 1.168 | 0.62099 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 2 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 217.4 | 134 | 0.66033 | 216.7397 | 133.3397 | 0.79587 | 0.35188 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 2 | 5 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 187.4 | 118.4 | 0.78917 | 186.6108 | 117.6108 | 0.71074 | 0.34105 | 20181101 | 9-58 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 2 | 6 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 192.6 | 116.7 | 0.78917 | 191.8108 | 115.9108 | 0.71759 | 0.2709 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | looks like C2N6a accidentally switched with C2N6 10b. Switch values
AR22 | 20170902 |  | 2 | 6 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 185.6 | 112.8 | 0.87933 | 184.7207 | 111.9207 | 0.68828 | 0.26618 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR22 | 20170902 |  | 2 | 6 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 165 | 100.7 | 0.493 | 164.507 | 100.207 | 0.60792 | 0.24665 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR22 | 20170902 |  | 2 | 6 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 158.4 | 100.4 | 0.87933 | 157.5207 | 99.5207 | 0.5919 | 0.28987 | 20181101 | 9-58 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | looks like C2N6a accidentally switched with C2N6 10a. Switch values
AR22 | 20170902 |  | 3 | 1 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 103.9 | 71.56 | 0.87933 | 103.0207 | 70.6807 | 0.30861 | 0.29979 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | tissue capsule was labeled 03-02a but does not match the 02b rep well, and the questionable repc does. This matched the 01b rep better and 01a was never found
AR22 | 20170902 |  | 3 | 1 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 108 | 77.8 | 0.66033 | 107.3397 | 77.1397 | 0.28819 | 0.37581 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 3 | 2 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 124.7 | 83.43 | 0.66033 | 124.0397 | 82.7697 | 0.39383 | 0.31863 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 0 | tissue capsule labeled 03-02c. Decide it was supposed to say a. Matched rep b well
AR22 | 20170902 |  | 3 | 2 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 123.4 | 82.22 | 0.87933 | 122.5207 | 81.3407 | 0.39297 | 0.30719 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | replaced replicate a with c? as provided in the lab notebook 8, pg 69
AR22 | 20170902 |  | 3 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 92.73 | 61.74 | 0.66033 | 92.0697 | 61.0797 | 0.29573 | 0.23003 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 3 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 88.81 | 60.01 | 0.493 | 88.317 | 59.517 | 0.27483 | 0.23747 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | recorded as replicate b in lab notebook 8-69
AR22 | 20170902 |  | 3 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 208.7 | 137.7 | 0.66033 | 208.0397 | 137.0397 | 0.67754 | 0.50206 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 3 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 214.8 | 139.4 | 0.87933 | 213.9207 | 138.5207 | 0.71952 | 0.47282 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 3 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 466.1 | 301 | 0.493 | 465.607 | 300.507 | 1.5755 | 1.0112 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 3 | 5 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 465.6 | 296.9 | 0.87933 | 464.7207 | 296.0207 | 1.6099 | 0.9382 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 3 | 6 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 148.4 | 96.64 | 0.87933 | 147.5207 | 95.7607 | 0.48936 | 0.32729 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 3 | 6 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 154.8 | 99.94 | 0.66033 | 154.1397 | 99.2797 | 0.51867 | 0.32799 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 3 | 6 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 133.4 | 85.49 | 0.66033 | 132.7397 | 84.8297 | 0.45296 | 0.27047 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | There is no <10b
AR22 | 20170902 |  | 4 | 1 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 114.7 | 76.89 | 0.66033 | 114.0397 | 76.2297 | 0.36081 | 0.29535 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 4 | 1 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 117 | 80.2 | 0.66033 | 116.3397 | 79.5397 | 0.35117 | 0.33348 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 4 | 2 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 386.7 | 251.4 | 0.66033 | 386.0397 | 250.7397 | 1.2911 | 0.86716 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 4 | 2 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 437.9 | 279.9 | 0.66033 | 437.2397 | 279.2397 | 1.5078 | 0.89586 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 4 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 68.91 | 47.85 | 0.66033 | 68.2497 | 47.1897 | 0.20097 | 0.20522 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR22 | 20170902 |  | 4 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 80.24 | 53.47 | 0.87933 | 79.3607 | 52.5907 | 0.25546 | 0.19723 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR22 | 20170902 |  | 4 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 262.2 | 166.4 | 0.87933 | 261.3207 | 165.5207 | 0.9142 | 0.51056 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 4 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 253.2 | 162.2 | 0.87933 | 252.3207 | 161.3207 | 0.86839 | 0.52021 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 4 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 315.2 | 202 | 0.66033 | 314.5397 | 201.3397 | 1.0802 | 0.65283 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 4 | 5 |  |  | b | 535 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 144.2 | 92.3 | 0.78917 | 143.4108 | 91.5108 | 1.0692 | 0.56755 | 20181101 | 9-58 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 4 | 6 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 166.5 | 103.6 | 0.87933 | 165.6207 | 102.7207 | 0.59468 | 0.28132 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 4 | 6 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 175.7 | 111.3 | 0.493 | 175.207 | 110.807 | 0.60886 | 0.3361 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 4 | 6 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 145.6 | 91.95 | 0.66033 | 144.9397 | 91.2897 | 0.50723 | 0.27129 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 4 | 6 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 140.7 | 89.33 | 0.66033 | 140.0397 | 88.6697 | 0.48567 | 0.2705 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 5 | 1 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 61.21 | 43.55 | 0.493 | 60.717 | 43.057 | 0.16853 | 0.2021 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | 30% percent range, could this have been switched with C5N2a?
AR22 | 20170902 |  | 5 | 1 |  |  | b | 535 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 27.05 | 20.55 | 0.78917 | 26.2608 | 19.7608 | 0.13391 | 0.21953 | 20181101 | 9-58 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | 30% percent range
AR22 | 20170902 |  | 5 | 2 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 51.99 | 39.48 | 0.87933 | 51.1107 | 38.6007 | 0.11938 | 0.21288 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | 23% percent range, could this have been switched with C5N1a?
AR22 | 20170902 |  | 5 | 2 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 57.48 | 41.61 | 0.493 | 56.987 | 41.117 | 0.15144 | 0.20248 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | 23% percent range
AR22 | 20170902 |  | 5 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 114.4 | 73.1 | 0.87933 | 113.5207 | 72.2207 | 0.39412 | 0.22754 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 5 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 110.7 | 70.73 | 0.66033 | 110.0397 | 70.0697 | 0.38142 | 0.22172 | 20180126 | 8-69 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 5 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 819.2 | 532.8 | 0.493 | 818.707 | 532.307 | 2.733 | 1.8489 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 5 | 4 |  |  | b | 535 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 479.8 | 308.5 | 0.66033 | 479.1397 | 307.8397 | 3.2694 | 2.0302 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 5 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 126.8 | 81.98 | 0.493 | 126.307 | 81.487 | 0.42771 | 0.27371 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 5 | 5 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 120.8 | 79.57 | 0.66033 | 120.1397 | 78.9097 | 0.39345 | 0.28578 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 5 | 6 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 45.17 | 28.07 | 0.87933 | 44.2907 | 27.1907 | 0.16167 | 0.070212 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR22 | 20170902 |  | 5 | 6 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 44.24 | 27.67 | 0.66033 | 43.5797 | 27.0097 | 0.15666 | 0.07368 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR22 | 20170902 |  | 5 | 6 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 36.32 | 22.99 | 0.66033 | 35.6597 | 22.3297 | 0.12603 | 0.064401 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 5 | 6 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 37.5 | 23.34 | 0.87933 | 36.6207 | 22.4607 | 0.13387 | 0.057671 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 6 | 1 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 25.72 | 18.83 | 0.87933 | 24.8407 | 17.9507 | 0.06575 | 0.088765 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 6 | 1 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 26.58 | 19.1 | 0.66033 | 25.9197 | 18.4397 | 0.07138 | 0.087344 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 6 | 2 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 32.26 | 23.76 | 0.493 | 31.767 | 23.267 | 0.081113 | 0.11916 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 6 | 2 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 30.06 | 22.24 | 0.493 | 29.567 | 21.747 | 0.074624 | 0.11257 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 6 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 44.14 | 31.91 | 0.66033 | 43.4797 | 31.2497 | 0.11671 | 0.15228 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 6 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 45.8 | 33.3 | 0.493 | 45.307 | 32.807 | 0.11928 | 0.16311 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 6 | 4 |  |  | a | 535 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 523 | 335.4 | 0.493 | 522.507 | 334.907 | 3.5804 | 2.1851 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 6 | 4 |  |  | b | 535 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 528.5 | 340.2 | 0.493 | 528.007 | 339.707 | 3.5938 | 2.2544 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 6 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 104.9 | 66.38 | 0.66033 | 104.2397 | 65.7197 | 0.36759 | 0.19811 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 6 | 5 |  |  | b | 535 | 0 | 5 |  | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 98.87 |  | 0.66033 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 9 | 0 | Ra was left blank in lab notebook
AR22 | 20170902 |  | 6 | 6 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 43.17 | 27.67 | 0.66033 | 42.5097 | 27.0097 | 0.14654 | 0.083796 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 6 | 6 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 44.48 | 28.33 | 0.493 | 43.987 | 27.837 | 0.15269 | 0.084706 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 6 | 6 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 35.76 | 22.97 | 0.66033 | 35.0997 | 22.3097 | 0.12092 | 0.069335 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170902 |  | 6 | 6 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 38.89 | 25.06 | 0.493 | 38.397 | 24.567 | 0.13075 | 0.078753 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | tissue capsule labeled 07-10a but assume it's N6 <10
AR22 | 20170903 |  | 7 | 2 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 26.51 | 18.68 | 0.66033 | 25.8497 | 18.0197 | 0.07472 | 0.080388 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | Cast 7 no bottle one- misfire
AR22 | 20170903 |  | 7 | 2 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 27.1 | 18.83 | 0.493 | 26.607 | 18.337 | 0.078919 | 0.078921 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 7 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 48.26 | 35.32 | 0.493 | 47.767 | 34.827 | 0.12348 | 0.1763 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 7 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 40.9 | 29.51 | 0.493 | 40.407 | 29.017 | 0.10869 | 0.14108 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 7 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 485.3 | 298.7 | 0.87933 | 484.4207 | 297.8207 | 1.7807 | 0.78288 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 7 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 522.4 | 317.5 | 0.66033 | 521.7397 | 316.8397 | 1.9553 | 0.77196 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 7 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 71.15 | 45.84 | 0.493 | 70.657 | 45.347 | 0.24153 | 0.14881 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 7 | 5 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 70.15 | 44.9 | 0.493 | 69.657 | 44.407 | 0.24095 | 0.14129 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 7 | 6 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 47.3 | 29.83 | 0.493 | 46.807 | 29.337 | 0.16517 | 0.085018 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 7 | 6 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 44.06 | 28.34 | 0.87933 | 43.1807 | 27.4607 | 0.14862 | 0.085562 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 7 | 6 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 36.92 | 23.5 | 0.493 | 36.427 | 23.007 | 0.12688 | 0.069326 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 7 | 6 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 35.17 | 22.66 | 0.66033 | 34.5097 | 21.9997 | 0.11827 | 0.069339 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | recorded as sample 07-10b (missing niskin # 6) in lab notebook
AR22 | 20170903 |  | 8 | 1 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 18.16 | 14.35 | 0.87933 | 17.2807 | 13.4707 | 0.036358 | 0.079594 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | 26% percent range
AR22 | 20170903 |  | 8 | 1 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 20.27 | 15.3 | 0.493 | 19.777 | 14.807 | 0.047428 | 0.080027 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | 26% percent range
AR22 | 20170903 |  | 8 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 17.34 | 12.6 | 0.493 | 16.847 | 12.107 | 0.045233 | 0.058981 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | chl#02 cruise logsheet says "#2 same as #1" | ignoring "chl" numbers, entering data using numbers on tissue capsules as nisken numbers
AR22 | 20170903 |  | 8 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 19.78 | 14.97 | 0.493 | 19.287 | 14.477 | 0.045901 | 0.078713 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | chl#02 | ignoring "chl" numbers, entering data using numbers on tissue capsules as nisken numbers
AR22 | 20170903 |  | 8 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 93.14 | 60.52 | 0.493 | 92.647 | 60.027 | 0.31128 | 0.20541 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | chl#03 | ignoring "chl" numbers, entering data using numbers on tissue capsules as nisken numbers
AR22 | 20170903 |  | 8 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 90.36 | 58.34 | 0.493 | 89.867 | 57.847 | 0.30556 | 0.19237 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | chl#03 | ignoring "chl" numbers, entering data using numbers on tissue capsules as nisken numbers
AR22 | 20170903 |  | 8 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 446.6 | 290.4 | 0.493 | 446.107 | 289.907 | 1.4906 | 1.0049 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | chl#04 | ignoring "chl" numbers, entering data using numbers on tissue capsules as nisken numbers
AR22 | 20170903 |  | 8 | 5 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 487.7 | 298.9 | 0.493 | 487.207 | 298.407 | 1.8017 | 0.76693 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | chl#04 | ignoring "chl" numbers, entering data using numbers on tissue capsules as nisken numbers
AR22 | 20170903 |  | 8 | 6 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 73.67 | 46.87 | 0.493 | 73.177 | 46.377 | 0.25575 | 0.14345 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | chl#05 | ignoring "chl" numbers, entering data using numbers on tissue capsules as nisken numbers
AR22 | 20170903 |  | 8 | 6 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 63.43 | 40.62 | 0.493 | 62.937 | 40.127 | 0.21767 | 0.12773 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | chl#05 | ignoring "chl" numbers, entering data using numbers on tissue capsules as nisken numbers
AR22 | 20170903 |  | 8 | 7 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 51.37 | 32.79 | 0.493 | 50.877 | 32.297 | 0.17566 | 0.099767 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | chl#06 | ignoring "chl" numbers, entering data using numbers on tissue capsules as nisken numbers
AR22 | 20170903 |  | 8 | 7 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 53.98 | 34.3 | 0.493 | 53.487 | 33.807 | 0.18606 | 0.10224 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | chl#06 | ignoring "chl" numbers, entering data using numbers on tissue capsules as nisken numbers
AR22 | 20170903 |  | 8 | 7 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 39.96 | 25.13 | 0.493 | 39.467 | 24.637 | 0.14021 | 0.069896 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | chl#06, Spilled some of one of the <10 | ignoring "chl" numbers, entering data using numbers on tissue capsules as nisken numbers
AR22 | 20170903 |  | 8 | 7 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 35.17 | 22.66 | 0.66033 | 34.5097 | 21.9997 | 0.11827 | 0.069339 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | chl#06 | ignoring "chl" numbers, entering data using numbers on tissue capsules as nisken numbers
AR22 | 20170903 |  | 9 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 7.557 | 5.931 | 0.87933 | 6.6777 | 5.0517 | 0.015517 | 0.027967 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | chl#01, poor replication but within noise of fluorometer | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 4.777 | 3.942 | 0.66033 | 4.1167 | 3.2817 | 0.0079682 | 0.020279 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | chl#01, poor replication but within noise of fluorometer | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 4.915 | 4.257 | 0.493 | 4.422 | 3.764 | 0.0062791 | 0.02612 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | chl#02 | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 4.678 | 4.101 | 0.493 | 4.185 | 3.608 | 0.0055062 | 0.02555 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | chl#02 | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 23.09 | 16.15 | 0.87933 | 22.2107 | 15.2707 | 0.066227 | 0.065219 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | chl#03 | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 5 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 26.51 | 17.75 | 0.66033 | 25.8497 | 17.0897 | 0.083595 | 0.063509 | 20180126 | 8-70 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | chl#03 | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 6 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 98.04 | 69.12 | 0.493 | 97.547 | 68.627 | 0.27342 | 0.31183 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | chl#04 | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 6 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 101.5 | 72.03 | 0.493 | 101.007 | 71.537 | 0.27862 | 0.33145 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | chl#04 | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 7 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 182.8 | 123.6 | 0.493 | 182.307 | 123.107 | 0.5597 | 0.49016 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | chl#05 | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 7 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 189.4 | 126.8 | 0.493 | 188.907 | 126.307 | 0.59185 | 0.4853 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | chl#05 | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 8 |  |  | a | 537 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 94.95 | 61.21 | 0.493 | 94.457 | 60.717 | 0.32077 | 0.19991 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | chl#06 | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 8 |  |  | b | 539 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 93.35 | 60.52 | 0.493 | 92.857 | 60.027 | 0.31096 | 0.2019 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | chl#06 | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 9 |  |  | a | 539 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 60.22 | 38.6 | 0.493 | 59.727 | 38.107 | 0.20478 | 0.1208 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | chl#07 | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 9 |  |  | b | 539 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 56.95 | 36.31 | 0.493 | 56.457 | 35.817 | 0.1955 | 0.11051 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | chl#07 | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 9 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 47.52 | 29.94 | 0.493 | 47.027 | 29.447 | 0.16621 | 0.084916 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | chl#07, written 10-10A but no niskin 10 sampled so assume it's <10um from surf | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 9 | 9 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 44.49 | 27.79 | 0.493 | 43.997 | 27.297 | 0.15789 | 0.074901 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | chl#07, written 10-10B but no niskin 10 sampled so assume it's <10um from surf | using numbers on tissue capsules as "chl" numbers
AR22 | 20170903 |  | 10 | 2 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 5.259 | 4.467 | 0.493 | 4.766 | 3.974 | 0.0075579 | 0.026649 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 10 | 2 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 6.121 | 5.162 | 0.493 | 5.628 | 4.669 | 0.0091515 | 0.031038 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 10 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 4.721 | 4.211 | 0.493 | 4.228 | 3.718 | 0.0048668 | 0.027137 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poore replication but values basically zero, within noise of flurometer
AR22 | 20170903 |  | 10 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 5.035 | 4.313 | 0.493 | 4.542 | 3.82 | 0.0068899 | 0.025992 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poore replication but values basically zero, within noise of flurometer
AR22 | 20170903 |  | 10 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 6.09 | 5.099 | 0.493 | 5.597 | 4.606 | 0.0094569 | 0.03019 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poore replication but values basically zero, within noise of flurometer
AR22 | 20170903 |  | 10 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 4.639 | 4.032 | 0.493 | 4.146 | 3.539 | 0.0057925 | 0.02467 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poore replication but values basically zero, within noise of flurometer
AR22 | 20170903 |  | 10 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 6.508 | 5.005 | 0.493 | 6.015 | 4.512 | 0.014343 | 0.024495 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 10 | 5 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 6.603 | 5.159 | 0.493 | 6.11 | 4.666 | 0.01378 | 0.026384 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 10 | 6 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 15.4 | 10.73 | 0.493 | 14.907 | 10.237 | 0.044152 | 0.043149 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 10 | 6 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 15.61 | 10.89 | 0.493 | 15.117 | 10.397 | 0.044625 | 0.044041 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 10 | 7 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 188.9 | 128.4 | 0.493 | 188.407 | 127.907 | 0.57199 | 0.5188 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 10 | 7 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 199.1 | 135.5 | 0.493 | 198.607 | 135.007 | 0.6013 | 0.55004 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 10 | 8 |  |  | a | 537 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 44.1 | 27.48 | 0.493 | 43.607 | 26.987 | 0.15801 | 0.073421 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 10 | 8 |  |  | b | 539 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 42.13 | 25.97 | 0.493 | 41.637 | 25.477 | 0.15307 | 0.064605 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 10 | 8 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 39.27 | 24.84 | 0.493 | 38.777 | 24.347 | 0.13643 | 0.071205 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | tissue capsule labled 11-10A but no niskin 10 sampled so assume it's <10 from surf
AR22 | 20170903 |  | 10 | 8 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 36.64 | 22.98 | 0.493 | 36.147 | 22.487 | 0.12915 | 0.062622 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | tissue capsule labled 11-10B but no niskin 10 sampled so assume it's <10 from surf
AR22 | 20170903 |  | 11 | 1 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 85.68 | 60.57 | 0.493 | 85.187 | 60.077 | 0.23962 | 0.27751 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 11 | 1 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 86.3 | 58.59 | 0.493 | 85.807 | 58.097 | 0.26443 | 0.23565 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 11 | 2 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 85.45 | 58.64 | 0.493 | 84.957 | 58.147 | 0.25584 | 0.24467 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 11 | 2 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 89.87 | 61.94 | 0.493 | 89.377 | 61.447 | 0.26653 | 0.26239 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 11 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 101.4 | 70.3 | 0.87933 | 100.5207 | 69.4207 | 0.29678 | 0.30077 | 20180126 | 8-76 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 11 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 100.7 | 67.37 | 0.493 | 100.207 | 66.877 | 0.31806 | 0.2576 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 11 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 275.3 | 174.5 | 0.493 | 274.807 | 174.007 | 0.96191 | 0.53589 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 11 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 291.5 | 188.6 | 0.493 | 291.007 | 188.107 | 0.98195 | 0.63722 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 11 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 53.23 | 33.36 | 0.493 | 52.737 | 32.867 | 0.18961 | 0.093295 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 11 | 5 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 50.08 | 30.9 | 0.493 | 49.587 | 30.407 | 0.18303 | 0.078705 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 11 | 6 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 48.8 | 30.25 | 0.493 | 48.307 | 29.757 | 0.17538 | 0.078389 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR22 | 20170903 |  | 11 | 6 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 20.9 | 13.21 | 0.493 | 20.407 | 12.717 | 0.072704 | 0.035746 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | These numbers seem very questionable so toss
AR22 | 20170903 |  | 11 | 6 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 39.17 | 24.78 | 0.493 | 38.677 | 24.287 | 0.13605 | 0.071071 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 11 | 6 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 43.44 | 26.96 | 0.493 | 42.947 | 26.467 | 0.15581 | 0.069903 | 20180126 | 8-64 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 12 | 1 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 93.99 | 68.09 | 0.493 | 93.497 | 67.597 | 0.24716 | 0.3347 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 12 | 1 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 90.96 | 65.74 | 0.493 | 90.467 | 65.247 | 0.24067 | 0.32096 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 12 | 2 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 119.4 | 81.54 | 0.493 | 118.907 | 81.047 | 0.36129 | 0.33634 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 12 | 2 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 106.8 | 72.3 | 0.493 | 106.307 | 71.807 | 0.32923 | 0.28887 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 12 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 187 | 122.7 | 0.493 | 186.507 | 122.207 | 0.6136 | 0.43832 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 12 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 216.2 | 137.7 | 0.493 | 215.707 | 137.207 | 0.74911 | 0.43193 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 12 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 635.3 | 387.3 | 0.493 | 634.807 | 386.807 | 2.3666 | 0.96292 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 12 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 647.1 | 392.9 | 0.493 | 646.607 | 392.407 | 2.4258 | 0.95196 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 12 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 178.2 | 111.8 | 0.493 | 177.707 | 111.307 | 0.63364 | 0.32446 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 12 | 5 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 180.3 | 113.6 | 0.493 | 179.807 | 113.107 | 0.6365 | 0.33709 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 12 | 6 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 45.41 | 27.94 | 0.493 | 44.917 | 27.447 | 0.16517 | 0.0689 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 12 | 6 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 45.95 | 28.34 | 0.493 | 45.457 | 27.847 | 0.16649 | 0.070988 | 20180126 | 8-66 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 12 | 6 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 39.08 | 24.39 | 0.493 | 38.587 | 23.897 | 0.13889 | 0.064909 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR22 | 20170903 |  | 12 | 6 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 42.14 | 26.05 | 0.493 | 41.647 | 25.557 | 0.15212 | 0.065829 | 20180126 | 8-65 |  | 3 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 1 | 2 |  |  | a | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 351.6 | 228.3 | 0.784 | 350.816 | 227.516 | 2.2322 | 1.4831 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 1 | 2 |  |  | b | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 359.9 | 237.5 | 0.784 | 359.116 | 236.716 | 2.216 | 1.6497 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 1 | 3 |  |  | a | 282 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 382.4 | 227.3 | 0.784 | 381.616 | 226.516 | 5.6159 | 1.7822 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 1 | 3 |  |  | b | 282 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 403.7 | 239.7 | 0.784 | 402.916 | 238.916 | 5.9382 | 1.8649 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 1 | 5 |  |  | a | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 730.5 | 439.2 | 0.784 | 729.716 | 438.416 | 5.2737 | 1.8857 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 1 | 5 |  |  | b | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 693.1 | 415.5 | 0.784 | 692.316 | 414.716 | 5.0257 | 1.7467 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 1 | 7 |  |  | a | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 438.3 | 258.7 | 0.784 | 437.516 | 257.916 | 3.2515 | 0.96032 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 1 | 7 |  |  | b | 282 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 4 |  |  |  |  | 0
AR24A | 20171022 |  | 1 | 7 |  |  | a | 283 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 107.9 | 66.75 | 0.784 | 107.116 | 65.966 | 0.74235 | 0.33108 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 1 | 7 |  |  | b | 283 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 4 |  |  |  |  | 0
AR24A | 20171022 |  | 3 | 1 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 94.92 | 67.92 | 0.784 | 94.136 | 67.136 | 0.25765 | 0.32023 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 3 | 1 |  |  | b | 540 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 48.09 | 33.94 | 0.784 | 47.306 | 33.156 | 0.26756 | 0.29795 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 3 | 2 |  |  | a | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 79.04 | 47.91 | 0.784 | 78.256 | 47.126 | 0.56358 | 0.206 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 3 | 2 |  |  | b | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 74.3 | 45.18 | 0.784 | 73.516 | 44.396 | 0.52719 | 0.1978 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 3 | 3 |  |  | a | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 225.7 | 145.2 | 0.784 | 224.916 | 144.416 | 1.4574 | 0.90096 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 3 | 3 |  |  | b | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 236.7 | 151 | 0.784 | 235.916 | 150.216 | 1.5515 | 0.90154 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 3 | 6 |  |  | a | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 120.9 | 78.2 | 0.784 | 120.116 | 77.416 | 0.77305 | 0.49117 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 3 | 6 |  |  | b | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 116.8 | 74.23 | 0.66033 | 116.1397 | 73.5697 | 0.77069 | 0.43072 | 20180126 | 8-70 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 3 | 6 |  |  | a | 283 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 118.4 | 74.45 | 0.784 | 117.616 | 73.666 | 0.79287 | 0.40587 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171022 |  | 3 | 6 |  |  | b | 283 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 119.3 | 74.3 | 0.784 | 118.516 | 73.516 | 0.81181 | 0.38448 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171023 |  | 5 | 1 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 113.1 | 81.37 | 0.784 | 112.316 | 80.586 | 0.30279 | 0.39087 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171023 |  | 5 | 1 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 121.2 | 87.28 | 0.784 | 120.416 | 86.496 | 0.32369 | 0.42084 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171023 |  | 5 | 2 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 138.9 | 90.79 | 0.784 | 138.116 | 90.006 | 0.4591 | 0.31564 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171023 |  | 5 | 2 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 143 | 95.86 | 0.784 | 142.216 | 95.076 | 0.44985 | 0.36854 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171023 |  | 5 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 151.1 | 92.25 | 0.784 | 150.316 | 91.466 | 0.56159 | 0.22572 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171023 |  | 5 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 150.3 | 95.61 | 0.784 | 149.516 | 94.826 | 0.52189 | 0.29434 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171023 |  | 5 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 695.2 | 408.2 | 0.784 | 694.416 | 407.416 | 2.7388 | 0.76815 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171023 |  | 5 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 730 | 428.8 | 0.784 | 729.216 | 428.016 | 2.8743 | 0.80996 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171023 |  | 5 | 6 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 537.8 | 329.2 | 0.784 | 537.016 | 328.416 | 1.9906 | 0.83629 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171023 |  | 5 | 6 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 538.6 | 329 | 0.784 | 537.816 | 328.216 | 2.0002 | 0.82503 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171023 |  | 5 | 8 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 253 | 158.2 | 0.784 | 252.216 | 157.416 | 0.90465 | 0.45034 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171023 |  | 5 | 8 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 249.4 | 153.6 | 0.784 | 248.616 | 152.816 | 0.9142 | 0.4012 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171023 |  | 5 | 8 |  |  | a | 595 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 272.6 | 169 | 0.784 | 271.816 | 168.216 | 0.88894 | 0.41301 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (21% range) but good phaeo replication
AR24A | 20171023 |  | 5 | 8 |  |  | b | 595 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 230 | 146.7 | 0.784 | 229.216 | 145.916 | 0.71475 | 0.4146 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (21% range) but good phaeo replication
AR24A | 20171024 |  | 6 | 1 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 541.5 | 334.7 | 0.784 | 540.716 | 333.916 | 1.9734 | 0.90081 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 6 | 1 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 532.1 | 322.2 | 0.66033 | 531.4397 | 321.5397 | 2.003 | 0.7647 | 20180126 | 8-70 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 6 | 2 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 530 | 331.6 | 0.66033 | 529.3397 | 330.9397 | 1.8933 | 0.95535 | 20180126 | 8-70 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 6 | 2 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 546.5 | 343.7 | 0.66033 | 545.8397 | 343.0397 | 1.9353 | 1.0175 | 20180126 | 8-70 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 6 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 311.6 | 199.7 | 0.784 | 310.816 | 198.916 | 1.0678 | 0.64438 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 6 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 328.4 | 207 | 0.784 | 327.616 | 206.216 | 1.1585 | 0.61656 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 6 | 3 |  |  | a | 595 | 10 | 5 |  | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 247.6 |  | 0.784 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 9 | 0 | Ra left blank in lab notebook
AR24A | 20171024 |  | 6 | 3 |  |  | b | 595 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 286.8 | 186.3 | 0.66033 | 286.1397 | 185.6397 | 0.86234 | 0.57446 | 20180126 | 8-70 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 8 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 150.2 | 101.5 | 0.784 | 149.416 | 100.716 | 0.46473 | 0.4022 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (41% range)
AR24A | 20171024 |  | 8 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 208.8 | 134.6 | 0.784 | 208.016 | 133.816 | 0.70807 | 0.44378 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (41% range)
AR24A | 20171024 |  | 8 | 4 |  |  | a | 535 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 355.5 | 219.9 | 0.784 | 354.716 | 219.116 | 2.588 | 1.1842 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 8 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 633.1 | 383.3 | 0.784 | 632.316 | 382.516 | 2.3838 | 0.90881 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 8 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 240.5 | 151.1 | 0.784 | 239.716 | 150.316 | 0.85312 | 0.44076 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 8 | 5 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 253.4 | 156.2 | 0.784 | 252.616 | 155.416 | 0.92756 | 0.41022 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 8 | 5 |  |  | a | 595 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 171.4 | 107.8 | 0.66033 | 170.7397 | 107.1397 | 0.54572 | 0.28351 | 20180126 | 8-70 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | cruise note "every cast from CTD8 onward <10a funnel filtering slow, <10b filtering fast, one must be wrong mesh size"
AR24A | 20171024 |  | 8 | 5 |  |  | b | 595 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 191.7 | 123 | 0.784 | 190.916 | 122.216 | 0.58948 | 0.35644 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | cruise note "every cast from CTD8 onward <10a funnel filtering slow, <10b filtering fast, one must be wrong mesh size"
AR24A | 20171024 |  | 10 | 1 |  |  | a | 535 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 4 |  |  |  |  | 0
AR24A | 20171024 |  | 10 | 1 |  |  | b | 535 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 4 |  |  |  |  | 0
AR24A | 20171024 |  | 10 | 2 |  |  | a | 535 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 4 |  |  |  |  | 0
AR24A | 20171024 |  | 10 | 2 |  |  | b | 535 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 4 |  |  |  |  | 0
AR24A | 20171024 |  | 10 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 123.1 | 79.98 | 0.784 | 122.316 | 79.196 | 0.41148 | 0.27021 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 10 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 132.4 | 86.94 | 0.784 | 131.616 | 86.156 | 0.43381 | 0.30779 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 10 | 4 |  |  | a | 535 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 4 |  |  |  |  | 0
AR24A | 20171024 |  | 10 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 267.9 | 163.4 | 0.784 | 267.116 | 162.616 | 0.99722 | 0.40253 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 10 | 5 |  |  | a | 535 | 0 | 5 |  | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 |  | 66.78 | 0.784 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20180126 | 8-73 |  | 9 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 9 | 0 | Rb left blank in lab notebook
AR24A | 20171024 |  | 10 | 5 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 110 | 68.5 | 0.784 | 109.216 | 67.716 | 0.39602 | 0.18686 | 20180126 | 8-72 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 10 | 5 |  |  | a | 595 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 84.3 | 52.67 | 0.66033 | 83.6397 | 52.0097 | 0.2714 | 0.13114 | 20180126 | 8-70 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | cruise note "every cast from CTD8 onward <10a funnel filtering slow, <10b filtering fast, one must be wrong mesh size"
AR24A | 20171024 |  | 10 | 5 |  |  | b | 595 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 93.96 | 58.59 | 0.784 | 93.176 | 57.806 | 0.30349 | 0.14391 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | cruise note "every cast from CTD8 onward <10a funnel filtering slow, <10b filtering fast, one must be wrong mesh size"
AR24A | 20171024 |  | 11 | 1 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 26.1 | 21.98 | 0.784 | 25.316 | 21.196 | 0.039316 | 0.14313 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor phaeo replication (38% range) but extreme depth of 442m
AR24A | 20171024 |  | 11 | 1 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 19.92 | 16.16 | 0.784 | 19.136 | 15.376 | 0.035881 | 0.096472 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor phaeo replication (38% range) but extreme depth of 442m
AR24A | 20171024 |  | 11 | 2 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 139.4 | 99.32 | 0.784 | 138.616 | 98.536 | 0.38247 | 0.4657 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 11 | 2 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 150.8 | 108.4 | 0.784 | 150.016 | 107.616 | 0.40461 | 0.52172 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 11 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 93.92 | 57.58 | 0.784 | 93.136 | 56.796 | 0.34678 | 0.1421 | 20180126 | 8-74 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 11 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 88.82 | 54 | 0.66033 | 88.1597 | 53.3397 | 0.33228 | 0.12685 | 20180126 | 8-70 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 11 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 87.8 | 52.77 | 0.66033 | 87.1397 | 52.1097 | 0.33428 | 0.11426 | 20180126 | 8-70 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 11 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 92.38 | 56.41 | 0.66033 | 91.7197 | 55.7497 | 0.34325 | 0.13662 | 20180126 | 8-70 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 11 | 4 |  |  | a | 595 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 73.36 | 45.25 | 0.66033 | 72.6997 | 44.5897 | 0.2412 | 0.10391 | 20180126 | 8-70 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | cruise note "every cast from CTD8 onward <10a funnel filtering slow, <10b filtering fast, one must be wrong mesh size"
AR24A | 20171024 |  | 11 | 4 |  |  | b | 595 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 78.17 | 48.11 | 0.784 | 77.386 | 47.326 | 0.25793 | 0.10836 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | cruise note "every cast from CTD8 onward <10a funnel filtering slow, <10b filtering fast, one must be wrong mesh size"
AR24A | 20171024 |  | 14 | 2 |  |  | a | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 108.3 | 79.11 | 0.784 | 107.516 | 78.326 | 0.14126 | 0.20064 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 14 | 2 |  |  | b | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 122.7 | 89.24 | 0.66033 | 122.0397 | 88.5797 | 0.16192 | 0.22473 | 20180126 |  |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 14 | 3 |  |  | a | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 242.4 | 159.3 | 0.66033 | 241.7397 | 158.6397 | 0.40214 | 0.29033 | 20180126 | 8-70 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 14 | 3 |  |  | b | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 228.6 | 148 | 0.784 | 227.816 | 147.216 | 0.39004 | 0.25256 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 14 | 4 |  |  | a | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 248.2 | 152 | 0.784 | 247.416 | 151.216 | 0.46553 | 0.19453 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 14 | 4 |  |  | b | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 240.6 | 149.1 | 0.784 | 239.816 | 148.316 | 0.44279 | 0.20462 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171024 |  | 14 | 4 |  |  | a | 1055 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 155.6 | 96.29 | 0.784 | 154.816 | 95.506 | 0.28701 | 0.12987 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | cruise note "every cast from CTD8 onward <10a funnel filtering slow, <10b filtering fast, one must be wrong mesh size"
AR24A | 20171024 |  | 14 | 4 |  |  | b | 1055 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 172.3 | 105.5 | 0.66033 | 171.6397 | 104.8397 | 0.32326 | 0.13437 | 20180126 | 8-70 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | cruise note "every cast from CTD8 onward <10a funnel filtering slow, <10b filtering fast, one must be wrong mesh size"
AR24A | 20171024 |  | 17 | 1 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 98.75 | 75.05 | 0.784 | 97.966 | 74.266 | 0.22895 | 0.41818 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | added this row to log sheet
AR24A | 20171025 |  | 17 | 1 |  |  | b | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 182.5 | 138.4 | 0.66033 | 181.8397 | 137.7397 | 0.21341 | 0.38783 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171025 |  | 17 | 2 |  |  | a | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 216.8 | 146.6 | 0.66033 | 216.1397 | 145.9397 | 0.33971 | 0.29732 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171025 |  | 17 | 2 |  |  | b | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 240.1 | 162.7 | 0.66033 | 239.4397 | 162.0397 | 0.37456 | 0.33276 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171025 |  | 17 | 3 |  |  | a | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 384.8 | 238.3 | 0.66033 | 384.1397 | 237.6397 | 0.70895 | 0.32836 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171025 |  | 17 | 3 |  |  | b | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 381.1 | 233.3 | 0.66033 | 380.4397 | 232.6397 | 0.71524 | 0.30025 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171025 |  | 17 | 5 |  |  | a | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 235.2 | 144.3 | 0.66033 | 234.5397 | 143.6397 | 0.43989 | 0.18711 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171025 |  | 17 | 5 |  |  | b | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 242.6 | 150.6 | 0.784 | 241.816 | 149.816 | 0.44521 | 0.20875 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24A | 20171025 |  | 17 | 6 |  |  | a | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 213.1 | 131.5 | 0.784 | 212.316 | 130.716 | 0.39488 | 0.1757 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR24A | 20171025 |  | 17 | 6 |  |  | b | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 205 | 126.4 | 0.66033 | 204.3397 | 125.7397 | 0.38036 | 0.1685 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR24A | 20171025 |  | 17 | 6 |  |  | a | 1130 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 167 | 102.7 | 0.66033 | 166.3397 | 102.0397 | 0.29051 | 0.12534 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | cruise note "every cast from CTD8 onward <10a funnel filtering slow, <10b filtering fast, one must be wrong mesh size"
AR24A | 20171025 |  | 17 | 6 |  |  | b | 1130 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 172.9 | 108.7 | 0.66033 | 172.2397 | 108.0397 | 0.29006 | 0.15024 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | cruise note "every cast from CTD8 onward <10a funnel filtering slow, <10b filtering fast, one must be wrong mesh size"
AR24B | 20171027 | L1 | 1 | 1 |  |  | a | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 493.2 | 308.1 | 0.66033 | 492.5397 | 307.4397 | 3.3511 | 1.6695 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24B | 20171027 | L1 | 1 | 1 |  |  | b | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 500.4 | 318.1 | 0.66033 | 499.7397 | 317.4397 | 3.3004 | 1.8835 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24B | 20171027 | L1 | 1 | 2 |  |  | a | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 629.3 | 371.6 | 0.66033 | 628.6397 | 370.9397 | 4.6654 | 1.3921 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24B | 20171027 | L1 | 1 | 2 |  |  | b | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 691.5 | 410.2 | 0.784 | 690.716 | 409.416 | 5.0927 | 1.5932 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | tissue capsule was labled C1N3b but log sheet says N2
AR24B | 20171027 | L1 | 1 | 4 |  |  | a | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 648.2 | 379.5 | 0.66033 | 647.5397 | 378.8397 | 4.8646 | 1.322 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | originally entered as C7N4a but no N4 sampled and values extremely high, at first thought was bucket but those tisscap were labeled very differently. C1N4 a/b are missing. More likely C1 was mistaken for C7
AR24B | 20171027 | L1 | 1 | 4 |  |  | b | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 672.2 | 405.8 | 0.66033 | 671.5397 | 405.1397 | 4.823 | 1.7931 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | originally entered as C7N4a but no N4 sampled and values extremely high, at first thought was bucket but those tisscap were labeled very differently. C1N4 a/b are missing. More likely C1 was mistaken for C7
AR24B | 20171027 | L1 | 1 | 5 |  |  | a | 282 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 4 |  |  |  |  | 0
AR24B | 20171027 | L1 | 1 | 5 |  |  | b | 282 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 4 |  |  |  |  | 0
AR24B | 20171027 | L1 | 1 | 5 |  |  | a | 283 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 66.59 | 42.03 | 0.784 | 65.806 | 41.246 | 0.44307 | 0.22811 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (23% range)
AR24B | 20171027 | L1 | 1 | 5 |  |  | b | 283 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 84.64 | 53.66 | 0.784 | 83.856 | 52.876 | 0.55889 | 0.30154 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (23% range)
AR24B | 20171028 | L3 | 2 | 1 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 106.1 | 71.21 | 0.63317 | 105.4668 | 70.5768 | 0.33295 | 0.27456 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24B | 20171028 | L3 | 2 | 1 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 107.9 | 74.4 | 0.63317 | 107.2668 | 73.7668 | 0.31968 | 0.31528 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24B | 20171028 | L3 | 2 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 326.5 | 210.7 | 0.63317 | 325.8668 | 210.0668 | 1.1051 | 0.70314 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 4/5/6
AR24B | 20171028 | L3 | 2 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 313.1 | 201.7 | 0.66033 | 312.4397 | 201.0397 | 1.0631 | 0.66743 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 4/5/6
AR24B | 20171028 | L3 | 2 | 5 |  |  | a | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 185.5 | 118.9 | 0.63317 | 184.8668 | 118.2668 | 1.2057 | 0.72559 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24B | 20171028 | L3 | 2 | 5 |  |  | b | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 178 | 111.7 | 0.63317 | 177.3668 | 111.0668 | 1.2003 | 0.61344 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24B | 20171028 | L3 | 2 | 7 |  |  | a | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 205.9 | 130.2 | 0.63317 | 205.2668 | 129.5668 | 1.3705 | 0.74537 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24B | 20171028 | L3 | 2 | 7 |  |  | b | 282 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 109.7 | 70.4 | 0.63317 | 109.0668 | 69.7668 | 1.423 | 0.85563 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24B | 20171028 | L3 | 2 | 7 |  |  | a | 283 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 86.71 | 54.58 | 0.63317 | 86.0768 | 53.9468 | 0.57963 | 0.29822 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR24B | 20171028 | L3 | 2 | 7 |  |  | c | 283 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 194.9 | 120.4 | 0.63317 | 194.2668 | 119.7668 | 1.344 | 0.60492 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | forgot to pre-filter <10 so even though labeled C2N7 10b it's actually a 3rd WSW replicate,
AR24B | 20171028 | L10 | 3 | 3 |  |  | a | 1060 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 11.08 | 9.465 | 0.87933 | 10.2007 | 8.5857 | 0.0077785 | 0.029522 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | OOI ID# 03/01 | extreme phaeo:chl ratio but very deep sample so makes sense
AR24B | 20171028 | L10 | 3 | 5 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 272.4 | 188.2 | 0.63317 | 271.7668 | 187.5668 | 0.40669 | 0.4105 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24B | 20171028 | L10 | 3 | 5 |  |  | b | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 254 | 175.9 | 0.66033 | 253.3397 | 175.2397 | 0.37723 | 0.38625 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24B | 20171028 | L10 | 3 | 7 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 413.5 | 252.2 | 0.784 | 412.716 | 251.416 | 0.77835 | 0.31598 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID# 03/02
AR24B | 20171028 | L10 | 3 | 8 |  |  | a | 1053 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 413.2 | 253.5 | 0.87933 | 412.3207 | 252.6207 | 0.77429 | 0.33051 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID# 03/03
AR24B | 20171028 | L10 | 3 | 10 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 381.8 | 233 | 0.784 | 381.016 | 232.216 | 0.71871 | 0.293 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID# 03/04
AR24B | 20171028 | L10 | 3 | 11 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 353.8 | 219.8 | 0.87933 | 352.9207 | 218.9207 | 0.64723 | 0.30656 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID# 03/05
AR24B | 20171028 | L10 | 3 | 12 |  |  | a | 1060 | 0 | 10 | 1 | 0 | 0.66667 | Cal_012618_12mm | 1.902 | 0.48424 | 419.7 | 264.2 | 0.87933 | 418.8207 | 263.3207 | 2.2468 | 1.1851 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24B | 20171028 | L10 | 3 | 12 |  |  | b | 1057 | 0 | 5 |  | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 348.6 |  | 0.66033 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 9 | 0 | Ra left blank in lab notebook
AR24B | 20171028 | L10 | 3 | 12 |  |  | a | 1052 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 320 | 193.5 | 0.63317 | 319.3668 | 192.8668 | 0.61391 | 0.23037 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24B | 20171028 | L10 | 3 | 12 |  |  | b | 1057 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 317.3 | 194.3 | 0.63317 | 316.6668 | 193.6668 | 0.5941 | 0.24967 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR24B | 20171028 |  | 4 | 4 |  |  | a | 1060 | 0 | 10 | 1 | 0 | 0.66667 | Cal_012618_12mm | 1.902 | 0.48424 | 419.7 | 264.2 | 0.87933 | 418.8207 | 263.3207 | 2.2468 | 1.1851 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 0 | OOD ID# 04/01; 4 samples from this depth - 2 pairs of relplicates with very different values. Which are correct? Flag as 2 | added 5ml acetone to extract at measure
AR24B | 20171028 |  | 4 | 4 |  |  | b | 1058 | 0 | 10 | 1 | 0 | 0.66667 | Cal_012618_12mm | 1.902 | 0.48424 | 421.1 | 263.6 | 0.87933 | 420.2207 | 262.7207 | 2.28 | 1.1506 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 0 | OOI ID# 04/02; 4 samples from this depth - 2 pairs of relplicates with very different values. Which are correct? Flag as 2 | added 5ml acetone to extract at measure
AR24B | 20171028 |  | 4 | 4/6 |  |  | a | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 235.2 | 147.6 | 0.63317 | 234.5668 | 146.9668 | 1.5859 | 0.81408 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 0 | niskin # from tissue sample, changed from log sheet recorded as 4/5/6; 4 samples from this depth - 2 pairs of relplicates with very different values. Which are correct? Flag as 2
AR24B | 20171028 |  | 4 | 4/6 |  |  | b | 282 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 240.2 | 151.8 | 0.63317 | 239.5668 | 151.1668 | 1.6004 | 0.86818 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 0 | niskin # from tissue sample, changed from log sheet recorded as 4/5/6; 4 samples from this depth - 2 pairs of relplicates with very different values. Which are correct?
AR24B | 20171028 |  | 4 | 7 |  |  | a | 1053 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 228.1 | 142.9 | 0.87933 | 227.2207 | 142.0207 | 0.41308 | 0.20802 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID# 04/03
AR24B | 20171028 |  | 4 | 9 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 172.8 | 108.1 | 0.784 | 172.016 | 107.316 | 0.31251 | 0.15505 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | OOI ID# 04/04 there were 2 tissue capsules labeled AR24C, this sample was never found so assume it was either a typo on tiss cap or lab NB
AR24B | 20171028 |  | 4 | 9/10 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 173.2 | 108.7 | 0.63317 | 172.5668 | 108.0668 | 0.31154 | 0.15929 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | tissue capsule literally had N9/10 on it
AR24B | 20171028 |  | 4 | 9/10 |  |  | b | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 164.1 | 102.5 | 0.66033 | 163.4397 | 101.8397 | 0.29753 | 0.14616 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | tissue capsule literally had N9/10 on it
AR24B | 20171028 |  | 4 | 10 |  |  | a | 1052 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 159.1 | 99.9 | 0.63317 | 158.4668 | 99.2668 | 0.2873 | 0.14724 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 9/10
AR24B | 20171028 |  | 4 | 10 |  |  | b | 1057 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 148.4 | 94.87 | 0.66033 | 147.7397 | 94.2097 | 0.25855 | 0.1519 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 9/10
AR24B | 20171028 |  | 5 | 2 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 46.67 | 35.6 | 0.87933 | 45.7907 | 34.7207 | 0.053469 | 0.097802 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 05/01
AR24B | 20171028 |  | 5 | 3 |  |  | a | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 503.1 | 326.3 | 0.87933 | 502.2207 | 325.4207 | 0.85557 | 0.5649 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 05/02
AR24B | 20171028 |  | 5 | 4 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 163.6 | 103.8 | 0.87933 | 162.7207 | 102.9207 | 0.28856 | 0.15942 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 05/03
AR24B | 20171028 |  | 5 | 5 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 154.1 | 98.26 | 0.784 | 153.316 | 97.476 | 0.26946 | 0.15483 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 05/04
AR24B | 20171029 |  | 6 | 2 |  |  | a | 1060 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 49.72 | 37.16 | 0.87933 | 48.8407 | 36.2807 | 0.060494 | 0.097126 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 06/01
AR24B | 20171029 |  | 6 | 3 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 272 | 189.8 | 0.784 | 271.216 | 189.016 | 0.39666 | 0.42607 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 06/02
AR24B | 20171029 |  | 6 | 4 |  |  | a | 1053 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 318.5 | 199.8 | 0.784 | 317.716 | 199.016 | 0.57551 | 0.29486 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 06/03
AR24B | 20171029 |  | 6 | 5 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 430.5 | 264.6 | 0.87933 | 429.6207 | 263.7207 | 0.80131 | 0.34767 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 06/04
AR24B | 20171029 | L8 | 7 | 3 |  |  | a | 1060 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 93.22 | 71.03 | 0.87933 | 92.3407 | 70.1507 | 0.10688 | 0.19789 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 07/01
AR24B | 20171029 | L8 | 7 | 4 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 88.8 | 65.39 | 0.784 | 88.016 | 64.606 | 0.11296 | 0.16824 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 07/02
AR24B | 20171029 | L8 | 7 | 5 |  |  | a | 1053 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 334 | 211 | 0.87933 | 333.1207 | 210.1207 | 0.59635 | 0.32258 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 07/03
AR24B | 20171029 | L8 | 7 | 6 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 323.1 | 207.6 | 0.87933 | 322.2207 | 206.7207 | 0.55787 | 0.34277 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 07/04
AR24B | 20171029 | L8 | 7 | 8 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 315.2 | 197.7 | 0.784 | 314.416 | 196.916 | 0.56753 | 0.29039 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 07/05
AR24B | 20171029 | L8 | 7 | 9 |  |  | a | 1055 | 0 | 5 |  | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 328.9 |  | 0.87933 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 07/06 Ra was blank in lab NB
AR24B | 20171029 | L8 | 7 | 10 |  |  | a | 1052 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 341.3 | 215.4 | 0.63317 | 340.6668 | 214.7668 | 0.611 | 0.32915 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 10
AR24B | 20171029 | L8 | 7 | 10 |  |  | b | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 327.1 | 205.2 | 0.784 | 326.316 | 204.416 | 0.58878 | 0.30181 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 10
AR24B | 20171029 | L8 | 7 | 10 |  |  | a | 1055 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 286.1 | 180.3 | 0.63317 | 285.4668 | 179.6668 | 0.51199 | 0.27227 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 10
AR24B | 20171029 | L8 | 7 | 10 |  |  | b | 1055 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 279.3 | 178.1 | 0.63317 | 278.6668 | 177.4668 | 0.48973 | 0.28492 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 10
AR24B | 20171030 |  | 0 | 0 |  |  | a | 283 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 381.9 | 233.7 | 0.66033 | 381.2397 | 233.0397 | 5.3471 | 2.2372 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | By Menemsha
AR24B | 20171101 |  | 8 | 4 |  |  | a | 1060 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 38.27 | 27.54 | 0.784 | 37.486 | 26.756 | 0.05168 | 0.064561 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 08/01
AR24B | 20171101 |  | 8 | 6 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 338.1 | 215.5 | 0.784 | 337.316 | 214.716 | 0.59161 | 0.34298 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 08/02
AR24B | 20171101 |  | 8 | 6 |  |  | b | 1053 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 341.7 | 220.6 | 0.87933 | 340.8207 | 219.7207 | 0.58714 | 0.37377 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 08/03
AR24B | 20171101 |  | 8 | 8 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 305.5 | 191.3 | 0.784 | 304.716 | 190.516 | 0.55159 | 0.27845 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 08/04
AR24B | 20171101 |  | 8 | 9 |  |  | a | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 337 | 208.6 | 0.66033 | 336.3397 | 207.9397 | 0.62136 | 0.28631 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 8/9
AR24B | 20171101 |  | 8 | 9 |  |  | b | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 300 | 192.9 | 0.784 | 299.216 | 192.116 | 0.5173 | 0.31971 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 8/9
AR24B | 20171101 |  | 8 | 9 |  |  | a | 1052 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 275.3 | 175.7 | 0.63317 | 274.6668 | 175.0668 | 0.48336 | 0.28299 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 8/9
AR24B | 20171101 |  | 8 | 9 |  |  | b | 1057 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 265.1 | 166.5 | 0.66033 | 264.4397 | 165.8397 | 0.47624 | 0.24628 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 8/9
AR24B | 20171101 | L6 | 9 | 3 |  |  | a | 1060 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 313.9 | 200.6 | 0.784 | 313.116 | 199.816 | 0.5457 | 0.3224 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 09/01
AR24B | 20171101 | L6 | 9 | 4 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 335.3 | 212.6 | 0.784 | 334.516 | 211.816 | 0.59209 | 0.32988 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 09/02
AR24B | 20171101 | L6 | 9 | 6 |  |  | a | 1053 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 332 | 200.8 | 0.784 | 331.216 | 200.016 | 1.2722 | 0.47725 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 09/03
AR24B | 20171101 | L6 | 9 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 678 | 408.3 | 0.784 | 677.216 | 407.516 | 1.3027 | 0.47279 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 09/04
AR24B | 20171101 | L6 | 9 | 9 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 364.8 | 221.2 | 0.87933 | 363.9207 | 220.3207 | 1.3872 | 0.53259 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 09/05
AR24B | 20171101 | L6 | 9 | 10 |  |  | a | 1055 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 397.2 | 242.2 | 0.87933 | 396.3207 | 241.3207 | 1.5002 | 0.6066 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID # 09/06
AR24B | 20171101 | L6 | 9 | 12 |  |  | a | 1052 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 472.7 | 285.1 | 0.66033 | 472.0397 | 284.4397 | 1.8209 | 0.66942 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 11/12
AR24B | 20171101 | L6 | 9 | 12 |  |  | b | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 482.8 | 291.6 | 0.66033 | 482.1397 | 290.9397 | 1.847 | 0.68811 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 11/12
AR24B | 20171101 | L6 | 9 | 12 |  |  | a | 1055 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 567.3 | 344.4 | 0.66033 | 566.6397 | 343.7397 | 1.0787 | 0.42178 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 11/12
AR24B | 20171101 | L6 | 9 | 12 |  |  | b | 1057 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 594.1 | 370.8 | 0.63317 | 593.4668 | 370.1668 | 1.0786 | 0.53419 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 11/12
AR24B | 20171102 |  | 10 | 4 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 246.1 | 153 | 0.66033 | 245.4397 | 152.3397 | 0.4501 | 0.21424 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 0 | N4 not recorded on cruise log sheet. Also N4 was 250m depth and 0.45 seems very unlikely. CTD fluor trace has no fluor at 250m. Flag 3 and throw out
AR24B | 20171102 |  | 10 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 527.6 | 337.3 | 0.63317 | 526.9668 | 336.6668 | 0.92003 | 0.54815 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 5/6/7
AR24B | 20171102 |  | 10 | 7 |  |  | b | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 504.1 | 324.2 | 0.784 | 503.316 | 323.416 | 0.86975 | 0.54064 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 5/6/7
AR24B | 20171102 |  | 10 | 9 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 445 | 291.7 | 0.66033 | 444.3397 | 291.0397 | 0.74045 | 0.52755 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 8/9
AR24B | 20171102 |  | 10 | 9 |  |  | b | 1065 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 452.3 | 291.8 | 0.66033 | 451.6397 | 291.1397 | 0.7694 | 0.48951 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 8/9
AR24B | 20171102 |  | 10 | 10 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 489.1 | 316.6 | 0.63317 | 488.4668 | 315.9668 | 0.83319 | 0.54342 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 10/11
AR24B | 20171102 |  | 10 | 10 |  |  | b | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 490.7 | 312.9 | 0.63317 | 490.0668 | 312.2668 | 0.85879 | 0.5017 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 10/11
AR24B | 20171102 |  | 10 | 10 |  |  | a | 1052 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 427.3 | 272.6 | 0.63317 | 426.6668 | 271.9668 | 0.75076 | 0.43977 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 10/11
AR24B | 20171102 |  | 10 | 10 |  |  | b | 1057 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 400.6 | 256.9 | 0.63317 | 399.9668 | 256.2668 | 0.69408 | 0.42242 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 10/11
AR24B | 20171103 | L13 | 11 | 2 |  |  | a | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 104.9 | 83.72 | 0.66033 | 104.2397 | 83.0597 | 0.10249 | 0.26007 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | niskin # from tissue sample, changed from log sheet recorded as 1/2
AR24B | 20171103 | L13 | 11 | 2 |  |  | b | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 116.9 | 90.56 | 0.63317 | 116.2668 | 89.9268 | 0.12722 | 0.26457 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | niskin # from tissue sample, changed from log sheet recorded as 1/2
AR24B | 20171103 | L13 | 11 | 5 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 466.7 | 305.4 | 0.63317 | 466.0668 | 304.7668 | 0.77909 | 0.54872 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 3/4/5
AR24B | 20171103 | L13 | 11 | 5 |  |  | b | 1065 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 457.8 | 295.7 | 0.66033 | 457.1397 | 295.0397 | 0.77707 | 0.4987 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 3/4/5
AR24B | 20171103 | L13 | 11 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 640.5 | 403.9 | 0.66033 | 639.8397 | 403.2397 | 1.1428 | 0.61404 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 6/7
AR24B | 20171103 | L13 | 11 | 7 |  |  | b | 1065 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 347 | 217.4 | 0.63317 | 346.3668 | 216.7668 | 1.2425 | 0.63208 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | lab NB said one of several samples was extracted in 10ml, definitely this one. True niskin taken from tissue capsule, recorded on log sheet as N6/7
AR24B | 20171103 | L13 | 11 | 11 |  |  | a | 253 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 282 | 176.9 | 0.63317 | 281.3668 | 176.2668 | 2.1209 | 1.0876 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | are these vol filter really 253?, niskin # from tissue sample, changed from log sheet recorded as 10/11
AR24B | 20171103 | L13 | 11 | 11 |  |  | b | 253 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 275.3 | 170.1 | 0.66033 | 274.6397 | 169.4397 | 2.1229 | 0.96129 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | are these vol filter really 253?, niskin # from tissue sample, changed from log sheet recorded as 10/11
AR24B | 20171103 | L13 | 11 | 11 |  |  | a | 253 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 213.2 | 131.9 | 0.63317 | 212.5668 | 131.2668 | 1.6406 | 0.74875 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | are these vol filter really 253?, niskin # from tissue sample, changed from log sheet recorded as 10/11
AR24B | 20171103 | L13 | 11 | 11 |  |  | b | 253 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 219.9 | 138.6 | 0.63317 | 219.2668 | 137.9668 | 1.6406 | 0.8707 | 20180126 | 8-67 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | are these vol filter really 253?, niskin # from tissue sample, changed from log sheet recorded as 10/11
AR24B | 20171103 |  | 13 | 4 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 81.68 | 59.94 | 0.784 | 80.896 | 59.156 | 0.10501 | 0.15272 | 20180126 | 8-73 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 3/4
AR24B | 20171103 |  | 13 | 4 |  |  | b | 1065 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 87.18 | 64.25 | 0.63317 | 86.5468 | 63.6168 | 0.10992 | 0.16516 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 3/4
AR24B | 20171103 |  | 13 | 7 |  |  | a | 283 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 324.8 | 200.9 | 0.63317 | 324.1668 | 200.2668 | 2.2352 | 1.0237 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 5/6/7
AR24B | 20171103 |  | 13 | 7 |  |  | b | 283 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 282.4 | 176.2 | 0.66033 | 281.7397 | 175.5397 | 1.9159 | 0.94061 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 0 | AR24B C13N7b never found, but C7N13b sample was run but that doesn't exist in log sheet, assuem it is C13N7b, reasonable agreement with rep a
AR24B | 20171103 |  | 13 | 10 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 475.9 | 293 | 0.63317 | 475.2668 | 292.3668 | 1.7454 | 0.77124 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 8/9/10
AR24B | 20171103 |  | 13 | 10 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 491.4 | 303.6 | 0.63317 | 490.7668 | 302.9668 | 1.7755 | 0.80817 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 8/9/10
AR24B | 20171103 |  | 13 | 10 |  |  | a | 535 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 409.1 | 250.5 | 0.63317 | 408.4668 | 249.8668 | 1.5135 | 0.6373 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 8/9/10
AR24B | 20171103 |  | 13 | 10 |  |  | b | 535 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 386.7 | 239.7 | 0.63317 | 386.0668 | 239.0668 | 1.4028 | 0.65503 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 8/9/10
AR24B | 20171103 |  | 15 | 4 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 55.33 | 42.6 | 0.63317 | 54.6968 | 41.9668 | 0.061487 | 0.12135 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 3/4
AR24B | 20171103 |  | 15 | 4 |  |  | b | 1065 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 52.88 | 41.18 | 0.63317 | 52.2468 | 40.5468 | 0.056087 | 0.11924 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 3/4
AR24B | 20171103 |  | 15 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 588.3 | 377.1 | 0.63317 | 587.6668 | 376.4668 | 1.0211 | 0.62067 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 5/6/7
AR24B | 20171103 |  | 15 | 7 |  |  | b | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 573.7 | 366.5 | 0.63317 | 573.0668 | 365.8668 | 1.0008 | 0.59322 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 5/6/7
AR24B | 20171103 |  | 15 | 10 |  |  | a | 1052 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 301.8 | 188.7 | 0.63317 | 301.1668 | 188.0668 | 1.0978 | 0.54877 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 8/9/10
AR24B | 20171103 |  | 15 | 10 |  |  | b | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 295.5 | 186.6 | 0.63317 | 294.8668 | 185.9668 | 1.052 | 0.56845 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 8/9/10
AR24B | 20171103 |  | 15 | 10 |  |  | a | 1055 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 385.5 | 243.7 | 0.66033 | 384.8397 | 243.0397 | 0.6862 | 0.37468 | 20180126 | 8-70 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 8/9/10
AR24B | 20171103 |  | 15 | 10 |  |  | b | 1057 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 457.2 | 287.4 | 0.63317 | 456.5668 | 286.7668 | 0.82014 | 0.42924 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 8/9/10
AR24B | 20171103 | L12 | 17 | 5 |  |  | a | 1057 | 0 | 5 |  | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 68.81 |  | 0.63317 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 9 | 0 | lab notebook: forgot to read Ra
AR24B | 20171103 | L12 | 17 | 5 |  |  | b | 1065 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 71.86 | 54.76 | 0.63317 | 71.2268 | 54.1268 | 0.081974 | 0.15207 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 9/10
AR24B | 20171103 | L12 | 17 | 8 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 572.9 | 375.7 | 0.66033 | 572.2397 | 375.0397 | 0.95249 | 0.68148 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 6/7/8
AR24B | 20171103 | L12 | 17 | 8 |  |  | b | 1065 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 569.5 | 364.9 | 0.66033 | 568.8397 | 364.2397 | 0.98081 | 0.59419 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 6/7/8
AR24B | 20171103 | L12 | 17 | 10 |  |  | a | 1052 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 492 | 314.4 | 0.63317 | 491.3668 | 313.7668 | 0.8619 | 0.51162 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 9/10
AR24B | 20171103 | L12 | 17 | 10 |  |  | b | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 495.7 | 313.4 | 0.63317 | 495.0668 | 312.7668 | 0.88052 | 0.48214 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 9/10
AR24B | 20171103 | L12 | 17 | 10 |  |  | a | 1055 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 332.3 | 219.4 | 0.63317 | 331.6668 | 218.7668 | 0.54635 | 0.40858 | 20180126 | 8-68 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 9/10
AR24B | 20171103 | L12 | 17 | 10 |  |  | b | 1057 | 10 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 341.3 | 220.4 | 0.66033 | 340.6397 | 219.7397 | 0.58395 | 0.37341 | 20180126 | 8-71 |  | 4 |  | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | niskin # from tissue sample, changed from log sheet recorded as 9/10
AR24C | 20171106 |  | 2 | 3 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 426 | 268.7 | 0.784 | 425.216 | 267.916 | 0.75905 | 0.4071 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 02/01 | volume not recorded in sample log, infer from brown bottle # volume from another cast
AR24C | 20171106 |  | 2 | 4 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 397.6 | 251.8 | 0.784 | 396.816 | 251.016 | 0.70356 | 0.38903 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 02/02 | volume not recorded in sample log, infer from brown bottle # volume from another cast
AR24C | 20171106 |  | 2 | 5 |  |  | a | 1053 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 531.1 | 328.8 | 0.784 | 530.316 | 328.016 | 0.98083 | 0.45369 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 02/03 | volume not recorded in sample log, infer from brown bottle # volume from another cast
AR24C | 20171106 |  | 2 | 6 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 573.5 | 357.1 | 0.87933 | 572.6207 | 356.2207 | 1.0452 | 0.50675 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 02/04 | volume not recorded in sample log, infer from brown bottle # volume from another cast
AR24C | 20171106 |  | 2 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 463.1 | 284.8 | 0.87933 | 462.2207 | 283.9207 | 0.8612 | 0.37578 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 02/05 | volume not recorded in sample log, infer from brown bottle # volume from another cast
AR24C | 20171106 |  | 2 | 8 |  |  | a | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 451.5 | 276.7 | 0.784 | 450.716 | 275.916 | 0.8459 | 0.35849 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 02/06 | volume not recorded in sample log, infer from brown bottle # volume from another cast
AR24C | 20171106 |  | 3 | 5 |  |  | a | 1060 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 471 | 286.4 | 0.784 | 470.216 | 285.616 | 0.88911 | 0.35174 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 03/01 | volume not recorded in sample log, infer from brown bottle # volume from another cast
AR24C | 20171106 |  | 3 | 6 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 496 | 305.2 | 0.87933 | 495.1207 | 304.3207 | 0.92071 | 0.4039 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 03/02 | volume not recorded in sample log, infer from brown bottle # volume from another cast
AR24C | 20171106 |  | 3 | 7 |  |  | a | 1053 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 372.8 | 230.6 | 0.784 | 372.016 | 229.816 | 0.68944 | 0.31562 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 03/03 | volume not recorded in sample log, infer from brown bottle # volume from another cast
AR24C | 20171106 |  | 3 | 8 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 424.5 | 260.7 | 0.784 | 423.716 | 259.916 | 0.79116 | 0.34124 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 03/04 | volume not recorded in sample log, infer from brown bottle # volume from another cast
AR24C | 20171107 |  | 4 | 5 |  |  | a | 1060 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 534.9 | 336.2 | 0.784 | 534.116 | 335.416 | 0.95702 | 0.50018 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 04/01
AR24C | 20171107 |  | 4 | 6 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 411.9 | 255.5 | 0.87933 | 411.0207 | 254.6207 | 0.75471 | 0.35357 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 04/02
AR24C | 20171107 |  | 4 | 7 |  |  | a | 1053 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 541 | 328.9 | 0.87933 | 540.1207 | 328.0207 | 1.0283 | 0.4062 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 04/03
AR24C | 20171107 |  | 4 | 8 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 489.3 | 295.5 | 0.784 | 488.516 | 294.716 | 0.93607 | 0.34795 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 04/04 | two 04/04 entries in lab notebook, second entry is Rb 172.8 Ra 108.1, these values match the duplicate depth better, maybe typo and other sample was AR24B 04/004 because that sample missing
AR24C | 20171109 |  | 9 | 4 |  |  | a | 1060 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 160.3 | 103.3 | 0.87933 | 159.4207 | 102.4207 | 0.27453 | 0.17043 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 09/01
AR24C | 20171109 |  | 9 | 5 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 643.2 | 391.4 | 0.784 | 642.416 | 390.616 | 1.2151 | 0.48516 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 09/02
AR24C | 20171109 |  | 9 | 6 |  |  | a | 1053 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 697.8 | 425.9 | 0.87933 | 696.9207 | 425.0207 | 1.3183 | 0.54048 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 09/03
AR24C | 20171109 |  | 9 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 735.5 | 441.1 | 0.87933 | 734.6207 | 440.2207 | 1.422 | 0.49598 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 09/04
AR24C | 20171109 |  | 9 | 8 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 370.6 | 228 | 0.784 | 369.816 | 227.216 | 1.3775 | 0.60233 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 09/05
AR24C | 20171110 |  | 10 | 5 |  |  | a | 1060 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 263.4 | 174.2 | 0.784 | 262.616 | 173.416 | 0.42962 | 0.32378 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 10/01
AR24C | 20171110 |  | 10 | 6 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 264.5 | 174.5 | 0.784 | 263.716 | 173.716 | 0.43429 | 0.32183 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 10/02
AR24C | 20171110 |  | 10 | 7 |  |  | a | 1053 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 268.8 | 176.6 | 0.784 | 268.016 | 175.816 | 0.44702 | 0.32188 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 10/03
AR24C | 20171110 |  | 10 | 8 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 269.1 | 178.8 | 0.784 | 268.316 | 178.016 | 0.43615 | 0.33942 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 10/04
AR24C | 20171111 |  | 11 | 5 |  |  | a | 1060 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 349.9 | 216 | 0.87933 | 349.0207 | 215.1207 | 0.64492 | 0.28967 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 11/01
AR24C | 20171111 |  | 11 | 6 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 343.7 | 210 | 0.87933 | 342.8207 | 209.1207 | 0.64517 | 0.26507 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 11/02
AR24C | 20171111 |  | 11 | 7 |  |  | a | 1053 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 351.2 | 216.2 | 0.87933 | 350.3207 | 215.3207 | 0.65454 | 0.28714 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 11/03
AR24C | 20171111 |  | 11 | 8 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 340.4 | 209.8 | 0.87933 | 339.5207 | 208.9207 | 0.63081 | 0.27942 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 11/04
AR24C | 20171112 |  | 12 | 3 |  |  | a | 1060 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 356.5 | 218.5 | 0.784 | 355.716 | 217.716 | 0.66466 | 0.2812 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 12/01
AR24C | 20171112 |  | 12 | 4 |  |  | a | 1058 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 388.3 | 235.5 | 0.784 | 387.516 | 234.716 | 0.73734 | 0.28431 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 12/02
AR24C | 20171112 |  | 12 | 5 |  |  | a | 1053 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 458.4 | 270.7 | 0.784 | 457.616 | 269.916 | 0.91005 | 0.27039 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 12/03
AR24C | 20171112 |  | 12 | 6 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 410 | 248.8 | 0.87933 | 409.1207 | 247.9207 | 0.77861 | 0.30153 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 12/04
AR24C | 20171112 |  | 12 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 479.6 | 287.4 | 0.784 | 478.816 | 286.616 | 0.92834 | 0.32039 | 20180126 | 8-74 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 12/05
AR24C | 20171112 |  | 12 | 8 |  |  | a | 1055 | 0 | 5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 425.7 | 254 | 0.87933 | 424.8207 | 253.1207 | 0.83089 | 0.27399 | 20180126 | 8-75 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | OOI ID 12/06
EN608 | 20180131 | L1 | 1 | 2 | 2018-01-31 22:00 | 2018-02-02 00:06 | a | 282 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 516.8 | 298.9 | 1.0527 | 515.7473 | 297.8473 | 5.5229 | 1.2866 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180131 | L1 | 1 | 2 | 2018-01-31 22:00 | 2018-02-02 00:06 | b | 282 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 561.5 | 326.7 | 1.0527 | 560.4473 | 325.6473 | 5.9512 | 1.4939 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180131 | L1 | 1 | 5 | 2018-01-31 22:00 | 2018-02-02 00:06 | a | 287 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 493.5 | 295.2 | 1.0527 | 492.4473 | 294.1473 | 4.9385 | 1.6692 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180131 | L1 | 1 | 5 | 2018-01-31 22:00 | 2018-02-02 00:06 | b | 285 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 514.4 | 298.7 | 1.0527 | 513.3473 | 297.6473 | 5.4096 | 1.3237 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180131 | L1 | 1 | 5 | 2018-01-31 22:00 | 2018-02-02 00:06 | a | 285 | 10 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 89.21 | 56.63 | 1.0527 | 88.1573 | 55.5773 | 0.81708 | 0.44018 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180131 | L1 | 1 | 5 | 2018-01-31 22:00 | 2018-02-02 00:06 | b | 283 | 10 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 98.44 | 62.24 | 1.0527 | 97.3873 | 61.1873 | 0.91428 | 0.47967 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180131 | L1 | 1 | 9 | 2018-01-31 22:00 | 2018-02-02 00:06 | a | 288 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 508.7 | 296.8 | 1.0527 | 507.6473 | 295.7473 | 5.2589 | 1.3617 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180131 | L1 | 1 | 9 | 2018-01-31 22:00 | 2018-02-02 00:06 | b | 288 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 530 | 307.7 | 1.0527 | 528.9473 | 306.6473 | 5.517 | 1.3476 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180131 | L1 | 1 | 19 | 2018-01-31 22:00 | 2018-02-02 00:06 | a | 283 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 520.7 | 311.9 | 1.0527 | 519.6473 | 310.8473 | 5.2735 | 1.8081 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180131 | L1 | 1 | 19 | 2018-01-31 22:00 | 2018-02-02 00:06 | b | 283 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 524.3 | 309.9 | 1.0527 | 523.2473 | 308.8473 | 5.4149 | 1.6211 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180131 | L1 | 1 | 19 | 2018-01-31 22:00 | 2018-02-02 00:06 | a | 282 | 10 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 73 | 47.69 | 1.0527 | 71.9473 | 46.6373 | 0.6415 | 0.42474 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180131 | L1 | 1 | 19 | 2018-01-31 22:00 | 2018-02-02 00:06 | b | 282 | 10 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 69.83 | 46.56 | 1.0527 | 68.7773 | 45.5073 | 0.5898 | 0.45061 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L2 | 4 | 1 | 2018-02-01 07:30 | 2018-02-02 07:00 | a | 282 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 374 | 219.8 | 1.0527 | 372.9473 | 218.7473 | 3.9083 | 1.0927 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L2 | 4 | 1 | 2018-02-01 07:30 | 2018-02-02 07:00 | b | 282 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 382.6 | 222.1 | 1.0527 | 381.5473 | 221.0473 | 4.068 | 0.98565 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L2 | 4 | 4 | 2018-02-01 07:30 | 2018-02-02 07:00 | a | 283 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 350.6 | 203.7 | 1.0527 | 349.5473 | 202.6473 | 3.7101 | 0.90647 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L2 | 4 | 4 | 2018-02-01 07:30 | 2018-02-02 07:00 | b | 285 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 359.7 | 208.9 | 1.0527 | 358.6473 | 207.8473 | 3.7819 | 0.91994 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L2 | 4 | 4 | 2018-02-01 07:30 | 2018-02-02 07:00 | a | 286 | 10 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 104.2 | 64.78 | 1.0527 | 103.1473 | 63.7273 | 0.98516 | 0.45142 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L2 | 4 | 4 | 2018-02-01 07:30 | 2018-02-02 07:00 | b | 288 | 10 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 95.71 | 60.52 | 1.0527 | 94.6573 | 59.4673 | 0.87334 | 0.4579 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L2 | 4 | 6 | 2018-02-01 07:30 | 2018-02-02 07:00 | a | 287 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 368.7 | 213.7 | 1.0527 | 367.6473 | 212.6473 | 3.8602 | 0.91676 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L2 | 4 | 6 | 2018-02-01 07:30 | 2018-02-02 07:00 | b | 285 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 357.3 | 207.8 | 1.0527 | 356.2473 | 206.7473 | 3.7493 | 0.92766 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L2 | 4 | 12 | 2018-02-01 07:30 | 2018-02-02 07:00 | a | 282 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 367.3 | 212.6 | 1.0527 | 366.2473 | 211.5473 | 3.921 | 0.91547 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L2 | 4 | 12 | 2018-02-01 07:30 | 2018-02-02 07:00 | b | 282 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 351.5 | 203.7 | 1.0527 | 350.4473 | 202.6473 | 3.7461 | 0.88688 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L2 | 4 | 12 | 2018-02-01 07:30 | 2018-02-02 07:00 | a | 200 | 5 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 236.7 | 146.8 | 1.0527 | 235.6473 | 145.7473 | 3.2128 | 1.4855 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | Also read on 10AU Rynearson fluorometer | Taylor previously QC'ed
EN608 | 20180201 | L2 | 4 | 12 | 2018-02-01 07:30 | 2018-02-02 07:00 | b | 200 | 5 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 225.3 | 134.2 | 1.0527 | 224.2473 | 133.1473 | 3.2557 | 1.0364 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | Also read on 10AU Rynearson fluorometer | Taylor previously QC'ed
EN608 | 20180201 | L2 | 4 | 12 | 2018-02-01 07:30 | 2018-02-02 07:00 | a | 283 | 10 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 71.95 | 45.61 | 1.0527 | 70.8973 | 44.5573 | 0.66525 | 0.34984 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (31% range)
EN608 | 20180201 | L2 | 4 | 12 | 2018-02-01 07:30 | 2018-02-02 07:00 | b | 283 | 10 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 97.49 | 61.12 | 1.0527 | 96.4373 | 60.0673 | 0.91857 | 0.44986 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (31% range)
EN608 | 20180201 | L2 | 4 | 12 | 2018-02-01 07:30 | 2018-02-02 07:00 | a | 200 | 20 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 228.3 | 133.6 | 1.0527 | 227.2473 | 132.5473 | 3.3844 | 0.88843 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | Also read on 10AU Rynearson fluorometer | value larger than either of the 5um values, but reasonable variation, 10AU fluor also has wide variation with a >20 being higher than>5
EN608 | 20180201 | L2 | 4 | 12 | 2018-02-01 07:30 | 2018-02-02 07:00 | b | 200 | 20 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 206.3 | 124.5 | 1.0527 | 205.2473 | 123.4473 | 2.9233 | 1.0561 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | Also read on 10AU Rynearson fluorometer | Taylor previously QC'ed
EN608 | 20180201 | L3 | 5 | 1 | 2018-02-01 11:25 | 2018-02-02 20:30 | a | 282 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 364.7 | 208.9 | 1.3165 | 363.3835 | 207.5835 | 3.9489 | 0.79696 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Heidi | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L3 | 5 | 1 | 2018-02-01 11:25 | 2018-02-02 20:30 | b | 282 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 342.2 | 194.3 | 1.3165 | 340.8835 | 192.9835 | 3.7487 | 0.66341 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Heidi | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L3 | 5 | 5 | 2018-02-01 11:25 | 2018-02-02 20:30 | a | 287 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 327.7 | 186.1 | 1.3165 | 326.3835 | 184.7835 | 3.5264 | 0.62454 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Heidi | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L3 | 5 | 5 | 2018-02-01 11:25 | 2018-02-02 20:30 | b | 285 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 328.6 | 188.2 | 1.3165 | 327.2835 | 186.8835 | 3.5211 | 0.70652 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Heidi | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L3 | 5 | 9 | 2018-02-01 11:25 | 2018-02-02 20:30 | a | 285 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 320.6 | 182.5 | 1.3165 | 319.2835 | 181.1835 | 3.4634 | 0.63526 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Heidi | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L3 | 5 | 9 | 2018-02-01 11:25 | 2018-02-02 20:30 | b | 286 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 317.4 | 181.2 | 1.3165 | 316.0835 | 179.8835 | 3.4038 | 0.65122 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Heidi | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L3 | 5 | 13 | 2018-02-01 11:25 | 2018-02-02 20:30 | a | 283 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 311.9 | 175.4 | 1.3165 | 310.5835 | 174.0835 | 3.4475 | 0.51841 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Heidi | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L3 | 5 | 13 | 2018-02-01 11:25 | 2018-02-02 20:30 | b | 282 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 313.3 | 178.7 | 1.3165 | 311.9835 | 177.3835 | 3.4116 | 0.64385 | 20180126 | en608_0202 | 20180202 | 1 | Emily | Heidi | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L4 | 8 | 2 | 2018-02-01 21:47 | 2018-02-02 20:55 | a | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 430.5 | 252.4 | 1.224 | 429.276 | 251.176 | 3.8692 | 1.0529 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L4 | 8 | 2 | 2018-02-01 21:47 | 2018-02-02 20:55 | b | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 468.5 | 272 | 1.224 | 467.276 | 270.776 | 4.269 | 1.0372 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L4 | 8 | 4 | 2018-02-01 21:47 | 2018-02-02 20:55 | a | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 402.7 | 229 | 1.224 | 401.476 | 227.776 | 3.7736 | 0.68993 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L4 | 8 | 4 | 2018-02-01 21:47 | 2018-02-02 20:55 | b | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 394.8 | 225.3 | 1.224 | 393.576 | 224.076 | 3.6824 | 0.70867 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L4 | 8 | 9 | 2018-02-01 21:47 | 2018-02-02 20:55 | a | 287 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 398.3 | 231.1 | 1.224 | 397.076 | 229.876 | 3.5691 | 0.8571 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L4 | 8 | 9 | 2018-02-01 21:47 | 2018-02-02 20:55 | b | 285 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 356.5 | 212.1 | 1.224 | 355.276 | 210.876 | 3.1041 | 0.98482 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L4 | 8 | 9 | 2018-02-01 21:47 | 2018-02-02 20:55 | a | 283 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 82.37 | 50.43 | 1.224 | 81.146 | 49.206 | 0.69144 | 0.2694 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L4 | 8 | 9 | 2018-02-01 21:47 | 2018-02-02 20:55 | b | 283 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 89.84 | 53.98 | 1.224 | 88.616 | 52.756 | 0.77631 | 0.25386 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L4 | 8 | 12 | 2018-02-01 21:47 | 2018-02-02 20:55 | a | 286 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 391.8 | 234.2 | 1.224 | 390.576 | 232.976 | 3.376 | 1.1256 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor phaeo replication (73% range)
EN608 | 20180201 | L4 | 8 | 12 | 2018-02-01 21:47 | 2018-02-02 20:55 | b | 286 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 404.9 | 226.3 | 1.224 | 403.676 | 225.076 | 3.8258 | 0.52315 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor phaeo replication (73% range)
EN608 | 20180201 | L4 | 8 | 12 | 2018-02-01 21:47 | 2018-02-02 20:55 | a | 288 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 56.46 | 35.6 | 1.224 | 55.236 | 34.376 | 0.44374 | 0.21587 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN608 | 20180201 | L4 | 8 | 12 | 2018-02-01 21:47 | 2018-02-02 20:55 | b | 283 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 66.5 | 42.77 | 1.224 | 65.276 | 41.546 | 0.51371 | 0.29756 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN608 | 20180201 | L4 | 8 | 16 | 2018-02-01 21:47 | 2018-02-02 20:55 | a | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 399.1 | 227.9 | 1.224 | 397.876 | 226.676 | 3.7193 | 0.72269 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN608 | 20180201 | L4 | 8 | 16 | 2018-02-01 21:47 | 2018-02-02 20:55 | b | 225 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 258.8 | 145.1 | 1.224 | 257.576 | 143.876 | 3.0959 | 0.43779 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | log sheet has vol filt of "?225?" somewhat poor replication (esp phaeo at 49% range) so toss this one
EN608 | 20180201 | L4 | 8 | 16 | 2018-02-01 21:47 | 2018-02-02 20:55 | a | 175 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 53.2 | 32.84 | 1.224 | 51.976 | 31.616 | 0.71277 | 0.2856 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L4 | 8 | 16 | 2018-02-01 21:47 | 2018-02-02 20:55 | b | 175 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 50.22 | 30.63 | 1.224 | 48.996 | 29.406 | 0.68581 | 0.24277 | 20180126 | en608_0202 | 20180202 | 1 | Taylor | Heidi & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L5 | 9 | 2 | 2018-02-02 01:50 | 2018-02-03 06:45 | a | 282 | 0 | 4.5 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 404.6 | 235 | 1.3482 | 403.2518 | 233.6518 | 2.7634 | 0.67061 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L5 | 9 | 2 | 2018-02-02 01:50 | 2018-02-03 06:45 | b | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 314.6 | 184.9 | 1.3482 | 313.2518 | 183.5518 | 2.8177 | 0.7792 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L5 | 9 | 4 | 2018-02-02 01:50 | 2018-02-03 06:45 | a | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 334.9 | 196 | 1.3482 | 333.5518 | 194.6518 | 3.0176 | 0.79685 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L5 | 9 | 4 | 2018-02-02 01:50 | 2018-02-03 06:45 | b | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 313.7 | 182.3 | 1.3482 | 312.3518 | 180.9518 | 2.8547 | 0.69132 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L5 | 9 | 9 | 2018-02-02 01:50 | 2018-02-03 06:45 | a | 287 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 339.9 | 196.6 | 1.3482 | 338.5518 | 195.2518 | 3.059 | 0.7006 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L5 | 9 | 9 | 2018-02-02 01:50 | 2018-02-03 06:45 | b | 285 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 309 | 180.5 | 1.3482 | 307.6518 | 179.1518 | 2.7623 | 0.71148 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L5 | 9 | 9 | 2018-02-02 01:50 | 2018-02-03 06:45 | a | 283 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 70 | 43.65 | 1.3482 | 68.6518 | 42.3018 | 0.57043 | 0.2556 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L5 | 9 | 9 | 2018-02-02 01:50 | 2018-02-03 06:45 | b | 285 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 71.58 | 44.99 | 1.3482 | 70.2318 | 43.6418 | 0.57159 | 0.27463 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L5 | 9 | 12 | 2018-02-02 01:50 | 2018-02-03 06:45 | a | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 316.7 | 182.7 | 1.3482 | 315.3518 | 181.3518 | 2.9112 | 0.64267 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L5 | 9 | 12 | 2018-02-02 01:50 | 2018-02-03 06:45 | b | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 295.4 | 171.2 | 1.3482 | 294.0518 | 169.8518 | 2.6982 | 0.63022 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L5 | 9 | 16 | 2018-02-02 01:50 | 2018-02-03 06:45 | a | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 277.6 | 163.4 | 1.3482 | 276.2518 | 162.0518 | 2.481 | 0.69462 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L5 | 9 | 16 | 2018-02-02 01:50 | 2018-02-03 06:45 | b | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 284.7 | 165.5 | 1.3482 | 283.3518 | 164.1518 | 2.5896 | 0.62715 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180201 | L5 | 9 | 16 | 2018-02-02 01:50 | 2018-02-03 06:45 | a | 283 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 70.99 | 44.59 | 1.3482 | 69.6418 | 43.2418 | 0.57151 | 0.27287 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poore replication (60% range) could filter have been clogged?
EN608 | 20180201 | L5 | 9 | 16 | 2018-02-02 01:50 | 2018-02-03 06:45 | b | 283 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 122.6 | 73.37 | 1.3482 | 121.2518 | 72.0218 | 1.0657 | 0.34063 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | poore replication (60% range) could filter have not been on all the way/leaking? Judging from this value being ~40% of WSW and surrounding water samples being ~20% decide to eliminate this sample
EN608 | 20180202 | L6 | 12 | 1 | 2018-02-02 12:00 | 2018-02-03 14:15 | a | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 162 | 98.17 | 1.122 | 160.878 | 97.048 | 1.3867 | 0.51507 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (20% range)
EN608 | 20180202 | L6 | 12 | 1 | 2018-02-02 12:00 | 2018-02-03 14:15 | b | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 191.1 | 112.5 | 1.122 | 189.978 | 111.378 | 1.7076 | 0.47501 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (20% range)
EN608 | 20180202 | L6 | 12 | 4 | 2018-02-02 12:00 | 2018-02-03 14:15 | a | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 177.2 | 105.3 | 1.122 | 176.078 | 104.178 | 1.562 | 0.47947 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L6 | 12 | 4 | 2018-02-02 12:00 | 2018-02-03 14:15 | b | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 166.8 | 99.22 | 1.122 | 165.678 | 98.098 | 1.4682 | 0.45418 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L6 | 12 | 8 | 2018-02-02 12:00 | 2018-02-03 14:15 | a | 287 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 165.2 | 97.34 | 1.122 | 164.078 | 96.218 | 1.4486 | 0.40409 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L6 | 12 | 8 | 2018-02-02 12:00 | 2018-02-03 14:15 | b | 285 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 184.3 | 107.7 | 1.122 | 183.178 | 106.578 | 1.6466 | 0.41993 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L6 | 12 | 8 | 2018-02-02 12:00 | 2018-02-03 14:15 | a | 200 | 5 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 111.2 | 67.15 | 1.122 | 110.078 | 66.028 | 1.3494 | 0.47505 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L6 | 12 | 8 | 2018-02-02 12:00 | 2018-02-03 14:15 | b | 200 | 5 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 112.1 | 69.42 | 1.122 | 110.978 | 68.298 | 1.3074 | 0.57974 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L6 | 12 | 8 | 2018-02-02 12:00 | 2018-02-03 14:15 | a | 285 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 69.25 | 42.91 | 1.122 | 68.128 | 41.788 | 0.56621 | 0.24406 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (37% range)
EN608 | 20180202 | L6 | 12 | 8 | 2018-02-02 12:00 | 2018-02-03 14:15 | b | 283 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 50.41 | 32.47 | 1.122 | 49.288 | 31.348 | 0.38837 | 0.22376 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (37% range)
EN608 | 20180202 | L6 | 12 | 8 | 2018-02-02 12:00 | 2018-02-03 14:15 | b | 200 | 20 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 124.2 | 72.47 | 1.122 | 123.078 | 71.348 | 1.5846 | 0.38679 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | value higher than both 5's and between WSW values, 10AU readings also highly variable, flag as 2 but is entire depth just highly variable?
EN608 | 20180202 | L6 | 12 | 11 | 2018-02-02 12:00 | 2018-02-03 14:15 | a | 283 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 187.1 | 110.3 | 1.122 | 185.978 | 109.178 | 1.6626 | 0.46934 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | Given high variability and of 5's and 20's and repeatedly higher than WSW, these WSW values seem to be suspiciously low
EN608 | 20180202 | L6 | 12 | 11 | 2018-02-02 12:00 | 2018-02-03 14:15 | b | 283 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 164.9 | 96.93 | 1.122 | 163.778 | 95.808 | 1.4714 | 0.39942 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | Given high variability and of 5's and 20's and repeatedly higher than WSW, these WSW values seem to be suspiciously low
EN608 | 20180202 | L6 | 12 | 11 | 2018-02-02 12:00 | 2018-02-03 14:15 | a | 200 | 5 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 152.8 | 92.57 | 1.122 | 151.678 | 91.448 | 1.845 | 0.6818 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | this value higher than both WSW samples and both 10AU >5 readings
EN608 | 20180202 | L6 | 12 | 11 | 2018-02-02 12:00 | 2018-02-03 14:15 | a | 288 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 67.39 | 42.09 | 1.122 | 66.268 | 40.968 | 0.53819 | 0.2479 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (75% range) interesting that phaeo match
EN608 | 20180202 | L6 | 12 | 11 | 2018-02-02 12:00 | 2018-02-03 14:15 | b | 288 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 134.6 | 77.68 | 1.122 | 133.478 | 76.558 | 1.2108 | 0.25817 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | eliminated <10 in previous cast because high value and assume pre-filter nitex leaking, other <10 from this cast much lower % of WSW value, assume same for this sample; interesting that phaeo match
EN608 | 20180202 | L6 | 12 | 14 | 2018-02-02 12:00 | 2018-02-03 14:15 | a | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 148.9 | 89.29 | 1.122 | 147.778 | 88.168 | 1.295 | 0.43274 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (24 % range) phaeo ok though, could it be the diff of 1 big cell?
EN608 | 20180202 | L6 | 12 | 14 | 2018-02-02 12:00 | 2018-02-03 14:15 | b | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 217.3 | 124.7 | 1.122 | 216.178 | 123.578 | 2.0117 | 0.40994 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (24 % range) phaeo ok though, could it be the diff of 1 big cell?
EN608 | 20180202 | L6 | 12 | 18 | 2018-02-02 12:00 | 2018-02-03 14:15 | a | 283 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 235.4 | 137.1 | 1.122 | 234.278 | 135.978 | 2.128 | 0.52723 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L6 | 12 | 18 | 2018-02-02 12:00 | 2018-02-03 14:15 | b | 282 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 226.2 | 132 | 1.122 | 225.078 | 130.878 | 2.0465 | 0.51823 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L6 | 12 | 18 | 2018-02-02 12:00 | 2018-02-03 14:15 | a | 283 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 68.63 | 43.02 | 1.122 | 67.508 | 41.898 | 0.55441 | 0.26373 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (chl 43% range)
EN608 | 20180202 | L6 | 12 | 18 | 2018-02-02 12:00 | 2018-02-03 14:15 | b | 282 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 45.44 | 29.12 | 1.122 | 44.318 | 27.998 | 0.35455 | 0.1941 | 20180126 | en608_0203 | 20180203 | 1 | Emily | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (chl 43% range)
EN608 | 20180202 | L7 | 14 | 1 | 2018-02-02 17:30 | 2018-02-03 18:30 | a | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 69.85 | 48.89 | 1.0727 | 68.7773 | 47.8173 | 0.24002 | 0.2539 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication (24% range)
EN608 | 20180202 | L7 | 14 | 1 | 2018-02-02 17:30 | 2018-02-03 18:30 | b | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 83.39 | 56.26 | 1.0727 | 82.3173 | 55.1873 | 0.3078 | 0.25697 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication (24% range)
EN608 | 20180202 | L7 | 14 | 5 | 2018-02-02 17:30 | 2018-02-03 18:30 | a | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 54.51 | 36.31 | 1.0727 | 53.4373 | 35.2373 | 0.20648 | 0.15412 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication (24% range)
EN608 | 20180202 | L7 | 14 | 5 | 2018-02-02 17:30 | 2018-02-03 18:30 | b | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 44.49 | 30.33 | 1.0727 | 43.4173 | 29.2573 | 0.16215 | 0.14006 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication (24% range)
EN608 | 20180202 | L7 | 14 | 9 | 2018-02-02 17:30 | 2018-02-03 18:30 | a | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 160.4 | 100.3 | 1.0727 | 159.3273 | 99.2273 | 0.68822 | 0.33672 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | no rep b, bottle wasn't filled
EN608 | 20180202 | L7 | 14 | 13 | 2018-02-02 17:30 | 2018-02-03 18:30 | a | 547 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 266.9 | 158.6 | 1.0727 | 265.8273 | 157.5273 | 1.213 | 0.37847 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L7 | 14 | 13 | 2018-02-02 17:30 | 2018-02-03 18:30 | b | 547 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 235.4 | 143.7 | 1.0727 | 234.3273 | 142.6273 | 1.027 | 0.41387 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L7 | 14 | 17 | 2018-02-02 17:30 | 2018-02-03 18:30 | a | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 476.5 | 287.2 | 1.0727 | 475.4273 | 286.1273 | 2.1677 | 0.78775 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L7 | 14 | 17 | 2018-02-02 17:30 | 2018-02-03 18:30 | b | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 500.7 | 292.4 | 1.0727 | 499.6273 | 291.3273 | 2.3853 | 0.62389 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L7 | 14 | 22 | 2018-02-02 17:30 | 2018-02-03 18:30 | a | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 578.9 | 337.8 | 1.0727 | 577.8273 | 336.7273 | 2.7353 | 0.71059 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L7 | 14 | 22 | 2018-02-02 17:30 | 2018-02-03 18:30 | b | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 517.3 | 304.6 | 1.0727 | 516.2273 | 303.5273 | 2.4131 | 0.69304 | 20180126 | en608_0203 | 20180203 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L8 | 15 | 1 | 2018-02-02 22:00 | 2018-02-04 06:55 | a | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 61.29 | 40.63 | 0.9985 | 60.2915 | 39.6315 | 0.23439 | 0.17118 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN608 | 20180202 | L8 | 15 | 1 | 2018-02-02 22:00 | 2018-02-04 06:55 | b | 535 | 0 | 7 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 43.84 | 29.25 | 0.9985 | 42.8415 | 28.2515 | 0.19492 | 0.14553 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | not sure why just this sample says vol extracted 7, but says that on cruise logsheet and makes replicate matchup better
EN608 | 20180202 | L8 | 15 | 3 | 2018-02-02 22:00 | 2018-02-04 06:55 | a | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 81.55 | 50.63 | 0.9985 | 80.5515 | 49.6315 | 0.3508 | 0.15711 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L8 | 15 | 3 | 2018-02-02 22:00 | 2018-02-04 06:55 | b | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 87.58 | 54.85 | 0.9985 | 86.5815 | 53.8515 | 0.3748 | 0.18144 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L8 | 15 | 5 | 2018-02-02 22:00 | 2018-02-04 06:55 | a | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 249.7 | 151.1 | 0.9985 | 248.7015 | 150.1015 | 1.1186 | 0.41744 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN608 | 20180202 | L8 | 15 | 5 | 2018-02-02 22:00 | 2018-02-04 06:55 | b | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 273.2 | 164.6 | 0.9985 | 272.2015 | 163.6015 | 1.2436 | 0.44627 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | C15 N5b and N8a appear to have been switched so switch them in this document
EN608 | 20180202 | L8 | 15 | 8 | 2018-02-02 22:00 | 2018-02-04 06:55 | a | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 345.5 | 206.2 | 0.9985 | 344.5015 | 205.2015 | 1.5952 | 0.52441 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | C15 N5b and N8a appear to have been switched so switch them in this document
EN608 | 20180202 | L8 | 15 | 8 | 2018-02-02 22:00 | 2018-02-04 06:55 | b | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 317.1 | 191.7 | 0.9985 | 316.1015 | 190.7015 | 1.436 | 0.53381 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN608 | 20180202 | L8 | 15 | 11 | 2018-02-02 22:00 | 2018-02-04 06:55 | a | 549 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 360.3 | 213.8 | 0.9985 | 359.3015 | 212.8015 | 1.6348 | 0.50719 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L8 | 15 | 11 | 2018-02-02 22:00 | 2018-02-04 06:55 | b | 549 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 305.5 | 184.9 | 0.9985 | 304.5015 | 183.9015 | 1.3458 | 0.50532 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L8 | 15 | 11 | 2018-02-02 22:00 | 2018-02-04 06:55 | a | 532 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 104.8 | 66.72 | 0.9985 | 103.8015 | 65.7215 | 0.43853 | 0.24416 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L8 | 15 | 11 | 2018-02-02 22:00 | 2018-02-04 06:55 | b | 532 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 113.1 | 72.59 | 0.9985 | 112.1015 | 71.5915 | 0.46651 | 0.27715 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L8 | 15 | 15 | 2018-02-02 22:00 | 2018-02-04 06:55 | a | 547 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 530.2 | 307.7 | 0.9985 | 529.2015 | 306.7015 | 2.492 | 0.60648 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L8 | 15 | 15 | 2018-02-02 22:00 | 2018-02-04 06:55 | b | 547 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 509.9 | 294.6 | 0.9985 | 508.9015 | 293.6015 | 2.4114 | 0.55478 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L8 | 15 | 15 | 2018-02-02 22:00 | 2018-02-04 06:55 | a | 532 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 160.7 | 98.41 | 0.9985 | 159.7015 | 97.4115 | 0.71732 | 0.29454 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L8 | 15 | 15 | 2018-02-02 22:00 | 2018-02-04 06:55 | b | 547 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 143.3 | 90.01 | 0.9985 | 142.3015 | 89.0115 | 0.59685 | 0.3024 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L8 | 15 | 19 | 2018-02-02 22:00 | 2018-02-04 06:55 | a | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 500.6 | 291.5 | 0.9985 | 499.6015 | 290.5015 | 2.3723 | 0.60058 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L8 | 15 | 19 | 2018-02-02 22:00 | 2018-02-04 06:55 | b | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 528.1 | 308.9 | 0.9985 | 527.1015 | 307.9015 | 2.4869 | 0.66406 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L8 | 15 | 19 | 2018-02-02 22:00 | 2018-02-04 06:55 | a | 547 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 137.7 | 85.75 | 0.9985 | 136.7015 | 84.7515 | 0.58185 | 0.27437 | 20180126 | en608_0204 | 20180204 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180202 | L8 | 15 | 19 | 2018-02-02 22:00 | 2018-02-04 06:55 | b | 547 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 149.7 | 94.32 | 0.9985 | 148.7015 | 93.3215 | 0.62026 | 0.32253 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L11 | 19 | 1 | freeze | freeze | a | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 7.957 | 6.36 | 1.0192 | 6.9378 | 5.3408 | 0.018288 | 0.036879 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high perce range (46%) but close to noise fo fluorometer and deep sampled so basically zero chl; not sure if switched a & b sample volumes
EN608 | 20180203 | L11 | 19 | 1 | freeze | freeze | b | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 6.672 | 5.665 | 1.0192 | 5.6528 | 4.6458 | 0.011425 | 0.036119 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high perce range (46%) but close to noise fo fluorometer and deep sampled so basically zero chl; not sure if switched a & b sample volumes
EN608 | 20180203 | L11 | 19 | 3 | freeze | freeze | a | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 15.72 | 11.23 | 1.0192 | 14.7008 | 10.2108 | 0.05094 | 0.053553 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication(33% range) chose volumes to make poor matchup less; cruise logsheet says not sure if switched a & b sample volumes
EN608 | 20180203 | L11 | 19 | 3 | freeze | freeze | b | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 13.21 | 9.989 | 1.0192 | 12.1908 | 8.9698 | 0.036885 | 0.055767 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication(33% range) chose volumes to make poor matchup less; cruise logsheet says not sure if switched a & b sample volumes
EN608 | 20180203 | L11 | 19 | 5 | freeze | freeze | a | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 15.2 | 11.97 | 1.0192 | 14.1808 | 10.9508 | 0.036988 | 0.076126 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L11 | 19 | 5 | freeze | freeze | b | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 16.31 | 12.77 | 1.0192 | 15.2908 | 11.7508 | 0.040538 | 0.08084 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L11 | 19 | 7 | freeze | freeze | a | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 53.34 | 37.23 | 1.0192 | 52.3208 | 36.2108 | 0.18448 | 0.18955 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L11 | 19 | 7 | freeze | freeze | b | 525 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 49.37 | 33.75 | 1.0192 | 48.3508 | 32.7308 | 0.18228 | 0.16225 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L11 | 19 | 9 | freeze | freeze | a | 549 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 291 | 175.1 | 1.0192 | 289.9808 | 174.0808 | 1.2934 | 0.45891 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L11 | 19 | 9 | freeze | freeze | b | 549 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 272.1 | 166.1 | 1.0192 | 271.0808 | 165.0808 | 1.1829 | 0.47879 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L11 | 19 | 11 | freeze | freeze | a | 547 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 514.8 | 306.1 | 1.0192 | 513.7808 | 305.0808 | 2.3375 | 0.74467 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN608 | 20180203 | L11 | 19 | 11 | freeze | freeze | b | 547 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 564 | 326.4 | 1.0192 | 562.9808 | 325.3808 | 2.6611 | 0.62607 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN608 | 20180203 | L11 | 19 | 11 | freeze | freeze | a | 547 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 166.4 | 106.3 | 1.0192 | 165.3808 | 105.2808 | 0.67313 | 0.39049 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | possible <10 mixup between N18 & N11 judging from readings
EN608 | 20180203 | L11 | 19 | 11 | freeze | freeze | b | 532 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 242.1 | 147 | 1.0192 | 241.0808 | 145.9808 | 1.0952 | 0.42122 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | possible <10 mixup between N18 & N11 judging from readings
EN608 | 20180203 | L11 | 19 | 15 | freeze | freeze | a | 535 | 0 | 6 | 1 | 0 | 0.5 | Cal_012618_12mm | 1.902 | 0.48424 | 254.2 | 149.1 | 1.0192 | 253.1808 | 148.0808 | 2.4071 | 0.65207 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN608 | 20180203 | L11 | 19 | 15 | freeze | freeze | b | 525 | 0 | 6 | 1 | 0 | 0.5 | Cal_012618_12mm | 1.902 | 0.48424 | 267.4 | 160.7 | 1.0192 | 266.3808 | 159.6808 | 2.4903 | 0.87135 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN608 | 20180203 | L11 | 19 | 18 | freeze | freeze | a | 540 | 0 | 6 | 1 | 0 | 0.5 | Cal_012618_12mm | 1.902 | 0.48424 | 286.5 | 171.4 | 1.0192 | 285.4808 | 170.3808 | 2.6117 | 0.87554 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN608 | 20180203 | L11 | 19 | 18 | freeze | freeze | b | 540 | 0 | 6 | 1 | 0 | 0.5 | Cal_012618_12mm | 1.902 | 0.48424 | 282.7 | 168.8 | 1.0192 | 281.6808 | 167.7808 | 2.5845 | 0.84956 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN608 | 20180203 | L11 | 19 | 18 | freeze | freeze | a | 547 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 223.9 | 140.6 | 1.0192 | 222.8808 | 139.5808 | 0.93297 | 0.47717 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | possible <10 mixup between N18 & N11 judging from readings
EN608 | 20180203 | L11 | 19 | 18 | freeze | freeze | b | 520 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 166.1 | 109.1 | 1.0192 | 165.0808 | 108.0808 | 0.67155 | 0.47705 | 20180126 | 8-94 |  | 1 | Taylor | Emily P | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | possible <10 mixup between N18 & N11 judging from readings
EN608 | 20180203 | L10 | 21 | 1 | 2018-02-04 02:06 | 2018-02-05 00:30 | a | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 8.295 | 6.825 | 0.85433 | 7.4407 | 5.9707 | 0.016678 | 0.044424 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L10 | 21 | 1 | 2018-02-04 02:06 | 2018-02-05 00:30 | b | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 7.835 | 6.551 | 0.85433 | 6.9807 | 5.6967 | 0.014703 | 0.044139 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L10 | 21 | 3 | 2018-02-04 02:06 | 2018-02-05 00:30 | a | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 7.114 | 6.03 | 0.85433 | 6.2597 | 5.1757 | 0.012413 | 0.041048 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor phaeo:chl ratio agreement
EN608 | 20180203 | L10 | 21 | 3 | 2018-02-04 02:06 | 2018-02-05 00:30 | b | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 6.721 | 5.487 | 0.85433 | 5.8667 | 4.6327 | 0.014 | 0.033409 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor phaeo:chl ratio agreement
EN608 | 20180203 | L10 | 21 | 5 | 2018-02-04 02:06 | 2018-02-05 00:30 | a | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 13.72 | 11.17 | 0.85433 | 12.8657 | 10.3157 | 0.029201 | 0.077352 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L10 | 21 | 5 | 2018-02-04 02:06 | 2018-02-05 00:30 | b | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 14.91 | 12.27 | 0.85433 | 14.0557 | 11.4157 | 0.029952 | 0.086872 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L10 | 21 | 9 | 2018-02-04 02:06 | 2018-02-05 00:30 | a | 532 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 49.26 | 33.65 | 0.85433 | 48.4057 | 32.7957 | 0.17976 | 0.1609 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | <10 and WSW tubes must have been switched because originially, both <10 were greater than both WSW
EN608 | 20180203 | L10 | 21 | 9 | 2018-02-04 02:06 | 2018-02-05 00:30 | b | 532 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 52.83 | 35.54 | 0.85433 | 51.9757 | 34.6857 | 0.19911 | 0.16119 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | <10 and WSW tubes must have been switched because originially, both <10 were greater than both WSW
EN608 | 20180203 | L10 | 21 | 9 | 2018-02-04 02:06 | 2018-02-05 00:30 | a | 549 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 106.3 | 64.86 | 0.85433 | 105.4457 | 64.0057 | 0.46244 | 0.18183 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (33% range) <10 and WSW tubes must have been switched because originially, both <10 were greater than both WSW
EN608 | 20180203 | L10 | 21 | 9 | 2018-02-04 02:06 | 2018-02-05 00:30 | b | 549 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 78.62 | 49.05 | 0.85433 | 77.7657 | 48.1957 | 0.32998 | 0.15515 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (33% range) <10 and WSW tubes must have been switched because originially, both <10 were greater than both WSW
EN608 | 20180203 | L10 | 21 | 11 | 2018-02-04 02:06 | 2018-02-05 00:30 | a | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 314.4 | 188.7 | 0.85433 | 313.5457 | 187.8457 | 1.4394 | 0.50088 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L10 | 21 | 11 | 2018-02-04 02:06 | 2018-02-05 00:30 | b | 525 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 289.4 | 174.7 | 0.85433 | 288.5457 | 173.8457 | 1.3385 | 0.49142 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L10 | 21 | 14 | 2018-02-04 02:06 | 2018-02-05 00:30 | a | 547 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 352.6 | 213.7 | 0.85433 | 351.7457 | 212.8457 | 1.5557 | 0.59461 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (33% range)
EN608 | 20180203 | L10 | 21 | 14 | 2018-02-04 02:06 | 2018-02-05 00:30 | b | 547 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 476.4 | 281.9 | 0.85433 | 475.5457 | 281.0457 | 2.1784 | 0.66089 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (33% range)
EN608 | 20180203 | L10 | 21 | 14 | 2018-02-04 02:06 | 2018-02-05 00:30 | a | 547 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 174.8 | 107.8 | 0.85433 | 173.9457 | 106.9457 | 0.75041 | 0.33003 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (42% range)
EN608 | 20180203 | L10 | 21 | 14 | 2018-02-04 02:06 | 2018-02-05 00:30 | b | 532 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 116 | 73.64 | 0.85433 | 115.1457 | 72.7857 | 0.48781 | 0.26825 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (42% range)
EN608 | 20180203 | L10 | 21 | 16 | 2018-02-04 02:06 | 2018-02-05 00:30 | a | 535 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 395.8 | 230.9 | 0.85433 | 394.9457 | 230.0457 | 1.8883 | 0.48788 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L10 | 21 | 16 | 2018-02-04 02:06 | 2018-02-05 00:30 | b | 531 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 367.2 | 216.4 | 0.85433 | 366.3457 | 215.5457 | 1.7399 | 0.50333 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L10 | 21 | 19 | 2018-02-04 02:06 | 2018-02-05 00:30 | a | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 392.5 | 232.2 | 0.85433 | 391.6457 | 231.3457 | 1.8186 | 0.54886 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L10 | 21 | 19 | 2018-02-04 02:06 | 2018-02-05 00:30 | b | 540 | 0 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 435.9 | 259.2 | 0.85433 | 435.0457 | 258.3457 | 2.0047 | 0.6391 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L10 | 21 | 19 | 2018-02-04 02:06 | 2018-02-05 00:30 | a | 547 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 131.8 | 83.36 | 0.85433 | 130.9457 | 82.5057 | 0.54253 | 0.29099 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180203 | L10 | 21 | 19 | 2018-02-04 02:06 | 2018-02-05 00:30 | b | 547 | 10 | 6 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 115.9 | 73.31 | 0.85433 | 115.0457 | 72.4557 | 0.47701 | 0.25498 | 20180126 | en608_0205 | 20180205 | 1 | Taylor | Emily & Ian | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | L9 | 22 | 1 | freeze | freeze | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 48.67 | 33.62 | 0.627 | 48.043 | 32.993 | 0.15359 | 0.13873 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication
EN608 | 20180204 | L9 | 22 | 1 | freeze | freeze | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 32.33 | 24.3 | 0.627 | 31.703 | 23.673 | 0.082713 | 0.12899 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication
EN608 | 20180204 | L9 | 22 | 3 | freeze | freeze | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 25.3 | 19.03 | 0.627 | 24.673 | 18.403 | 0.064584 | 0.099992 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | L9 | 22 | 3 | freeze | freeze | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 31.89 | 24.25 | 0.627 | 31.263 | 23.623 | 0.077967 | 0.13134 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | L9 | 22 | 5 | freeze | freeze | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 60.81 | 40.63 | 0.627 | 60.183 | 40.003 | 0.20594 | 0.14849 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | originally written as 22-15a in lab notebook?? No 22-15 and 5a was unfilled
EN608 | 20180204 | L9 | 22 | 5 | freeze | freeze | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 62.51 | 42.51 | 0.627 | 61.883 | 41.883 | 0.20601 | 0.16855 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | L9 | 22 | 8 | freeze | freeze | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 283.1 | 173 | 0.627 | 282.473 | 172.373 | 1.1236 | 0.40366 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | L9 | 22 | 8 | freeze | freeze | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 306.4 | 188.6 | 0.627 | 305.773 | 187.973 | 1.2134 | 0.46762 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | L9 | 22 | 10 | freeze | freeze | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 291.1 | 177.5 | 0.627 | 290.473 | 176.873 | 1.1701 | 0.41162 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | L9 | 22 | 10 | freeze | freeze | b | 531 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 143.7 | 88.09 | 0.627 | 143.073 | 87.463 | 1.1543 | 0.42188 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | L9 | 22 | 13 | freeze | freeze | a | 535 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 229.8 | 136.7 | 0.627 | 229.173 | 136.073 | 1.918 | 0.51582 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | L9 | 22 | 13 | freeze | freeze | b | 535 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 241.3 | 144 | 0.627 | 240.673 | 143.373 | 2.0045 | 0.55986 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | L9 | 22 | 16 | freeze | freeze | a | 549 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 233.5 | 139.2 | 0.627 | 232.873 | 138.573 | 1.8931 | 0.52215 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | L9 | 22 | 16 | freeze | freeze | b | 549 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 228.2 | 136.4 | 0.627 | 227.573 | 135.773 | 1.8429 | 0.52354 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | L9 | 22 | 19 | freeze | freeze | a | 530 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 235.2 | 141.9 | 0.627 | 234.573 | 141.273 | 1.9402 | 0.61041 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | L9 | 22 | 19 | freeze | freeze | b | 530 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 253.7 | 151.7 | 0.627 | 253.073 | 151.073 | 2.1211 | 0.60643 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | MVCO | 30 | 1 | freeze | freeze | a | 282 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 329.7 | 200.7 | 0.627 | 329.073 | 200.073 | 5.0418 | 1.7472 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | MVCO | 30 | 1 | freeze | freeze | b | 282 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 344.9 | 210.1 | 0.627 | 344.273 | 209.473 | 5.2685 | 1.8394 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | MVCO | 30 | 5 | freeze | freeze | a | 282 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 267.2 | 162.9 | 0.627 | 266.573 | 162.273 | 4.0764 | 1.4299 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | MVCO | 30 | 5 | freeze | freeze | b | 282 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 284.1 | 170.5 | 0.627 | 283.473 | 169.873 | 4.4399 | 1.3243 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | MVCO | 30 | 9 | freeze | freeze | a | 282 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 241.9 | 147.9 | 0.627 | 241.273 | 147.273 | 3.6739 | 1.3235 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | MVCO | 30 | 9 | freeze | freeze | b | 282 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 263.3 | 159.5 | 0.627 | 262.673 | 158.873 | 4.0569 | 1.3341 | 20181101 | 9-86 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN608 | 20180204 | MVCO | 30 | 9 | freeze | freeze | a | 283 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 90.01 | 61.71 | 0.627 | 89.383 | 61.083 | 0.55108 | 0.4816 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (29% range)
EN608 | 20180204 | MVCO | 30 | 9 | freeze | freeze | b | 283 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 112.8 | 74.65 | 0.627 | 112.173 | 74.023 | 0.74288 | 0.50856 | 20181101 | 9-84 |  | 1 | Taylor |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (29% range)
EN617 | 20180702 | L1 | 1 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 229.6 | 147.4 | 0.6432 | 228.9568 | 146.7568 | 1.562 | 0.85919 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L1 | 1 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 236.8 | 149.1 | 0.6432 | 236.1568 | 148.4568 | 1.6665 | 0.78272 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L1 | 1 | 5 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 85.6 | 51.22 | 0.53383 | 85.0662 | 50.6862 | 0.9473 | 0.26523 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR3 | somewhat poor phaeo replication (32% range); Taylor previously QC'ed
EN617 | 20180702 | L1 | 1 | 5 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 91.97 | 56.9 | 0.53383 | 91.4362 | 56.3662 | 0.96632 | 0.38209 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR4 | somewhat poor phaeo replication (32% range); Taylor previously QC'ed
EN617 | 20180702 | L1 | 1 | 5 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 27.4 | 16.6 | 0.508 | 26.892 | 16.092 | 0.29758 | 0.087375 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR9 | Taylor previously QC'ed
EN617 | 20180702 | L1 | 1 | 5 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 22.08 | 13.61 | 0.53383 | 21.5462 | 13.0762 | 0.23338 | 0.07943 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR10 | cruise logsheet "got more water from niskin" so not exact same sample water at same time; Taylor previously QC'ed
EN617 | 20180702 | L1 | 1 | 5 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 290.8 | 181.7 | 0.6432 | 290.1568 | 181.0568 | 2.1096 | 0.92994 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L1 | 1 | 5 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 308.4 | 192.9 | 0.6432 | 307.7568 | 192.2568 | 2.2333 | 0.99421 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L1 | 1 | 5 |  |  | a | 286 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 253.3 | 159.6 | 0.6432 | 252.6568 | 158.9568 | 1.8055 | 0.85371 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L1 | 1 | 5 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 242 | 154.3 | 0.6432 | 241.3568 | 153.6568 | 1.6958 | 0.88375 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L1 | 1 | 8 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 60.3 | 36.63 | 0.53383 | 59.7662 | 36.0962 | 0.6522 | 0.2113 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR5 | poor replication (23% range); originally flagged as 2 but upon second look variation seems reasonable
EN617 | 20180702 | L1 | 1 | 8 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 48.28 | 29.55 | 0.53383 | 47.7462 | 29.0162 | 0.51608 | 0.17805 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR6 | poor replication (23% range); originally flagged as 2 but upon second look variation seems reasonable
EN617 | 20180702 | L1 | 1 | 8 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.31 | 8.099 | 0.53383 | 12.7762 | 7.5652 | 0.14358 | 0.037393 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR11
EN617 | 20180702 | L1 | 1 | 8 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 14.29 | 8.735 | 0.53383 | 13.7562 | 8.2012 | 0.15306 | 0.043129 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR12
EN617 | 20180702 | L1 | 1 | 8 |  |  | a | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 195.5 | 121.9 | 0.6432 | 194.8568 | 121.2568 | 1.4132 | 0.60821 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L1 | 1 | 8 |  |  | b | 288 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 212.9 | 132.6 | 0.6432 | 212.2568 | 131.9568 | 1.5365 | 0.65565 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L1 | 1 | 8 |  |  | a | 287 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 155.9 | 98.22 | 0.6432 | 155.2568 | 97.5768 | 1.1075 | 0.51913 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L1 | 1 | 8 |  |  | b | 286 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 148.9 | 94.08 | 0.6432 | 148.2568 | 93.4368 | 1.0563 | 0.5068 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L1 | 1 | 12 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 31.14 | 18.91 | 0.53383 | 30.6062 | 18.3762 | 0.33698 | 0.10262 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR1
EN617 | 20180702 | L1 | 1 | 12 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 34.13 | 21.04 | 0.53383 | 33.5962 | 20.5062 | 0.36068 | 0.12987 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR2
EN617 | 20180702 | L1 | 1 | 12 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 10.91 | 6.308 | 0.508 | 10.402 | 5.8 | 0.1268 | 0.011946 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR7 | very poor phaeo replication (79% range); Taylor previously QC'ed
EN617 | 20180702 | L1 | 1 | 12 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 11.13 | 6.864 | 0.508 | 10.622 | 6.356 | 0.11754 | 0.034505 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR8 | very poor phaeo replication (79% range); Taylor previously QC'ed
EN617 | 20180702 | L1 | 1 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 147.4 | 80.56 | 0.6432 | 146.7568 | 79.9168 | 1.2701 | 0.048338 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | Not real Rb, exact same value as previous Ra in lab NB
EN617 | 20180702 | L1 | 1 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 124.1 | 77.78 | 0.6432 | 123.4568 | 77.1368 | 0.8802 | 0.39241 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN617 | 20180702 | L1 | 1 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 92.12 | 58.06 | 0.6432 | 91.4768 | 57.4168 | 0.64723 | 0.30004 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L1 | 1 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 101.9 | 64.41 | 0.6432 | 101.2568 | 63.7668 | 0.71241 | 0.33962 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L2 | 2 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 90.16 | 60.75 | 0.6432 | 89.5168 | 60.1068 | 0.55887 | 0.43278 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L2 | 2 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 87.81 | 61.52 | 0.6432 | 87.1668 | 60.8768 | 0.49958 | 0.50477 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L2 | 2 | 3 |  |  | a | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 313.9 | 199.4 | 0.6432 | 313.2568 | 198.7568 | 2.2062 | 1.1187 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L2 | 2 | 3 |  |  | b | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 330.5 | 208.8 | 0.6432 | 329.8568 | 208.1568 | 2.3368 | 1.1333 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L2 | 2 | 11 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 48.87 | 32.05 | 0.53383 | 48.3362 | 31.5162 | 0.46346 | 0.29048 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR17
EN617 | 20180702 | L2 | 2 | 11 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 51.23 | 33.82 | 0.53383 | 50.6962 | 33.2862 | 0.47971 | 0.31657 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR18
EN617 | 20180702 | L2 | 2 | 11 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 10.08 | 6.923 | 0.53383 | 9.5462 | 6.3892 | 0.086988 | 0.065856 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR23 | somewhat poor phaeo replication (37% range); Taylor previously QC'ed
EN617 | 20180702 | L2 | 2 | 11 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 11.97 | 8.541 | 0.53383 | 11.4362 | 8.0072 | 0.094482 | 0.097068 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR24 | somewhat poor phaeo replication (37% range); Taylor previously QC'ed
EN617 | 20180702 | L2 | 2 | 11 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 288.9 | 181.9 | 0.6432 | 288.2568 | 181.2568 | 2.069 | 0.9739 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L2 | 2 | 11 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 310.3 | 197.5 | 0.64317 | 309.6568 | 196.8568 | 2.1811 | 1.1236 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L2 | 2 | 11 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 270.2 | 170.1 | 0.6432 | 269.5568 | 169.4568 | 1.9355 | 0.90923 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | may be 2-10-10a om 9-51 has 248.5, 158.4, 0.6432 avg
EN617 | 20180702 | L2 | 2 | 11 |  |  | b | 286 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 267.3 | 170.6 | 0.64317 | 266.6568 | 169.9568 | 1.8633 | 0.97993 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L2 | 2 | 12 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 20.56 | 12.86 | 0.53383 | 20.0262 | 12.3262 | 0.21217 | 0.082705 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR15
EN617 | 20180702 | L2 | 2 | 12 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 17.18 | 10.75 | 0.53383 | 16.6462 | 10.2162 | 0.17717 | 0.067223 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR16
EN617 | 20180702 | L2 | 2 | 12 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 5.191 | 3.174 | 0.53383 | 4.6572 | 2.6402 | 0.055576 | 0.0075826 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR21 | poor phaeo replication (55% range); phaeo value extremely low, not really in line with other samples at this depth, this sample causing the suspect variation
EN617 | 20180702 | L2 | 2 | 12 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 4.896 | 3.166 | 0.53383 | 4.3622 | 2.6322 | 0.047668 | 0.015299 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR22 | poor phaeo replication (55% range), this phaeo more in line with surf sample (2.7 vs 9.8m)
EN617 | 20180702 | L2 | 2 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 152.5 | 94.75 | 0.64317 | 151.8568 | 94.1068 | 1.0974 | 0.45518 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L2 | 2 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 143.1 | 88.11 | 0.6432 | 142.4568 | 87.4668 | 1.045 | 0.39808 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L2 | 2 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 128 | 79.51 | 0.6432 | 127.3568 | 78.8668 | 0.92144 | 0.37971 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L2 | 2 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 121.8 | 74.58 | 0.6432 | 121.1568 | 73.9368 | 0.89731 | 0.32251 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | typo in lab NB says 2-2 10b, but investigated and feel confident it's C2N12<10b
EN617 | 20180702 | L2 | 2 | 16 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.4449 | 4.429 | 0.53383 | 6.9111 | 3.8952 | 0.0831 | 0.010082 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR13
EN617 | 20180702 | L2 | 2 | 16 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.915 | 4.248 | 0.53383 | 6.3812 | 3.7142 | 0.073486 | 0.015365 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR14
EN617 | 20180702 | L2 | 2 | 16 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.929 | 1.576 | 0.53383 | 1.3952 | 1.0422 | 0.0097265 | 0.015204 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR19 | poor replication (33% range) but within noise of fluorometer and value basically zero; ; Taylor previously QC'ed
EN617 | 20180702 | L2 | 2 | 16 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.608 | 1.357 | 0.53383 | 1.0742 | 0.82317 | 0.006916 | 0.012776 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR20 | poor replication (33% range) but within noise of fluorometer and value basically zero; ; Taylor previously QC'ed
EN617 | 20180702 | L2 | 2 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 58.5 | 35.6 | 0.6432 | 57.8568 | 34.9568 | 0.43516 | 0.14156 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN617 | 20180702 | L2 | 2 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 75.32 | 36.53 | 0.6432 | 74.6768 | 35.8868 | 0.73711 | -0.14505 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | negative phaeo
EN617 | 20180702 | L2 | 2 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 56.23 | 34.4 | 0.64317 | 55.5868 | 33.7568 | 0.41483 | 0.1421 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | lab notebook said cas2niskin16 for two values. Moved other values to cas12niskin16
EN617 | 20180702 | L2 | 2 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 57.7 | 35.07 | 0.6742 | 57.0258 | 34.3958 | 0.43003 | 0.13744 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | lab notebook said cas2niskin16 for two values. Moved other values to cas12niskin16
EN617 | 20180702 | L3 | 5 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 45.29 | 31.83 | 0.6432 | 44.6468 | 31.1868 | 0.25578 | 0.25875 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L3 | 5 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 43.4 | 30.78 | 0.63133 | 42.7687 | 30.1487 | 0.23981 | 0.25758 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L3 | 5 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 164.4 | 103.3 | 0.6432 | 163.7568 | 102.6568 | 1.1611 | 0.53258 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L3 | 5 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 176.5 | 111.8 | 0.6432 | 175.8568 | 111.1568 | 1.2295 | 0.60441 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L3 | 5 | 6 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 27.21 | 18.06 | 0.53383 | 26.6762 | 17.5262 | 0.25212 | 0.16715 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR25 | very poor replication (50% range); Taylor previously QC'ed; QF was 2, DF changed to 1 because within biological variabilty for size fractions
EN617 | 20180702 | L3 | 5 | 6 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 44.35 | 29.09 | 0.53383 | 43.8162 | 28.5562 | 0.42047 | 0.26266 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR26 | very poor replication (50% range); Taylor previously QC'ed; QF was 2, DF changed to 1 because within biological variabilty for size fractions
EN617 | 20180702 | L3 | 5 | 6 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 14.54 | 9.808 | 0.53383 | 14.0062 | 9.2742 | 0.13039 | 0.091474 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR31 | poor replication (24% range); Taylor previously QC'ed
EN617 | 20180702 | L3 | 5 | 6 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 20.11 | 14.05 | 0.53383 | 19.5762 | 13.5162 | 0.16698 | 0.15636 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR32 | poor replication (24% range); Taylor previously QC'ed
EN617 | 20180702 | L3 | 5 | 6 |  |  | a | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 296 | 192.8 | 0.6432 | 295.3568 | 192.1568 | 1.9885 | 1.2261 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L3 | 5 | 6 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 304 | 194.6 | 0.6432 | 303.3568 | 193.9568 | 2.1154 | 1.1407 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L3 | 5 | 10 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 12.68 | 8.025 | 0.53383 | 12.1462 | 7.4912 | 0.12826 | 0.050943 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR27 | poor replication (24% range); Taylor previously QC'ed
EN617 | 20180702 | L3 | 5 | 10 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 10.18 | 6.546 | 0.53383 | 9.6462 | 6.0122 | 0.10013 | 0.043694 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR28 | poor replication (24% range) could this C5N10 >5b actually be the >20a? Unlikely with filtering protocol; Taylor previously QC'ed
EN617 | 20180702 | L3 | 5 | 10 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.696 | 4.335 | 0.53383 | 6.1622 | 3.8012 | 0.065055 | 0.025878 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR33 | extremely poor replication (80% range), something wrong with one of these?; Taylor previously QC'ed
EN617 | 20180702 | L3 | 5 | 10 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 14.99 | 9.458 | 0.53383 | 14.4562 | 8.9242 | 0.15243 | 0.061058 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR34 | extremely poor replication (80% range), something wrong with one of these? Could this C5N10 >20b actually be the >5? Unlikely with filtering protocol; Taylor previously QC'ed
EN617 | 20180702 | L3 | 5 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 82.97 | 51.64 | 0.6432 | 82.3268 | 50.9968 | 0.59535 | 0.246 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L3 | 5 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 80.49 | 51.24 | 0.6432 | 79.8468 | 50.5968 | 0.55583 | 0.27892 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L3 | 5 | 12 |  |  | a | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 55.67 | 34.35 | 0.6432 | 55.0268 | 33.7068 | 0.40937 | 0.15254 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L3 | 5 | 12 |  |  | b | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 50.81 | 30.67 | 0.6432 | 50.1668 | 30.0268 | 0.38807 | 0.11425 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L3 | 5 | 16 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.731 | 4.388 | 0.53383 | 6.1972 | 3.8542 | 0.064559 | 0.027642 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR29 | Taylor previously QC'ed
EN617 | 20180702 | L3 | 5 | 16 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.774 | 4.239 | 0.53383 | 6.2402 | 3.7052 | 0.069849 | 0.018787 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR30 | Taylor previously QC'ed
EN617 | 20180702 | L3 | 5 | 16 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.272 | 1.516 | 0.53383 | 1.7382 | 0.98217 | 0.020831 | 0.0026649 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR35 | Taylor previously QC'ed
EN617 | 20180702 | L3 | 5 | 16 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.447 | 1.694 | 0.53383 | 1.9132 | 1.1602 | 0.020748 | 0.0070058 | 20181101 | 9-44 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR36 | Taylor previously QC'ed
EN617 | 20180702 | L3 | 5 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 42.55 | 25.9 | 0.6432 | 41.9068 | 25.2568 | 0.31639 | 0.1003 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L3 | 5 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 42.45 | 25.94 | 0.6432 | 41.8068 | 25.2968 | 0.31373 | 0.10362 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L3 | 5 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 35.54 | 21.93 | 0.6432 | 34.8968 | 21.2868 | 0.25863 | 0.092566 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L3 | 5 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 34.18 | 20.62 | 0.6432 | 33.5368 | 19.9768 | 0.25768 | 0.071903 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L4 | 8 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 32.72 | 23.89 | 0.6432 | 32.0768 | 23.2468 | 0.16779 | 0.21574 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L4 | 8 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 32.27 | 24.28 | 0.6742 | 31.5958 | 23.6058 | 0.15183 | 0.23762 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L4 | 8 | 4 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 44.71 | 29.72 | 0.53383 | 44.1762 | 29.1862 | 0.41303 | 0.28517 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR37
EN617 | 20180702 | L4 | 8 | 4 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 38.38 | 25.35 | 0.53383 | 37.8462 | 24.8162 | 0.35903 | 0.23463 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR38
EN617 | 20180702 | L4 | 8 | 4 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 14.63 | 10.24 | 0.53383 | 14.0962 | 9.7062 | 0.12096 | 0.11123 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR43 | poor replication (29% range) but other sample questionable;Taylor previously QC'ed
EN617 | 20180702 | L4 | 8 | 4 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 17.51 | 11.59 | 0.53383 | 16.9762 | 11.0562 | 0.16312 | 0.10137 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR44 | logsheet not *spilled! Maybe refilled w bottle 7? Due to poor replication toss sample; Taylor previously QC'ed
EN617 | 20180702 | L4 | 8 | 4 |  |  | a | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 156.5 | 100.7 | 0.6432 | 155.8568 | 100.0568 | 1.0714 | 0.59657 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L4 | 8 | 4 |  |  | b | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 157.6 | 101.5 | 0.6432 | 156.9568 | 100.8568 | 1.0772 | 0.60415 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L4 | 8 | 4 |  |  | a | 288 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 117.6 | 71.42 | 0.6432 | 116.9568 | 70.7768 | 0.88364 | 0.29216 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor phaeo replication (42% range) but chl reps ok
EN617 | 20180702 | L4 | 8 | 4 |  |  | b | 286 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 124.2 | 79.73 | 0.6432 | 123.5568 | 79.0868 | 0.85687 | 0.46617 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor phaeo replication (42% range) but chl reps ok
EN617 | 20180702 | L4 | 8 | 7 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 63.1 | 43.47 | 0.53383 | 62.5662 | 42.9362 | 0.54088 | 0.48625 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR39
EN617 | 20180702 | L4 | 8 | 7 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 59.2 | 41.58 | 0.53383 | 58.6662 | 41.0462 | 0.4855 | 0.49642 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR40
EN617 | 20180702 | L4 | 8 | 7 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 27.46 | 20.25 | 0.53383 | 26.9262 | 19.7162 | 0.19866 | 0.27299 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR45
EN617 | 20180702 | L4 | 8 | 7 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 26.25 | 19.38 | 0.508 | 25.742 | 18.872 | 0.1893 | 0.26217 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR46
EN617 | 20180702 | L4 | 8 | 7 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 270.8 | 173.3 | 0.6742 | 270.1258 | 172.6258 | 1.8853 | 1.0127 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L4 | 8 | 7 |  |  | b | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 281.7 | 185.1 | 0.6742 | 281.0258 | 184.4258 | 1.8613 | 1.2239 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L4 | 8 | 7 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 216.4 | 139.9 | 0.6432 | 215.7568 | 139.2568 | 1.4792 | 0.85857 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L4 | 8 | 7 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 220.9 | 140.9 | 0.6432 | 220.2568 | 140.2568 | 1.5469 | 0.80769 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L4 | 8 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 281.6 | 181.2 | 0.6432 | 280.9568 | 180.5568 | 1.9079 | 1.071 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L4 | 8 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 287.5 | 183.6 | 0.6432 | 286.8568 | 182.9568 | 1.9744 | 1.0441 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L4 | 8 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 122.1 | 73.8 | 0.6432 | 121.4568 | 73.1568 | 0.91783 | 0.28912 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (24% range)
EN617 | 20180702 | L4 | 8 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 98.76 | 60.88 | 0.6432 | 98.1168 | 60.2368 | 0.71982 | 0.27397 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (24% range)
EN617 | 20180702 | L4 | 8 | 16 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.136 | 3.988 | 0.53383 | 5.6022 | 3.4542 | 0.059186 | 0.023446 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR41
EN617 | 20180702 | L4 | 8 | 16 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 5.937 | 3.929 | 0.53383 | 5.4032 | 3.3952 | 0.055328 | 0.025892 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR42
EN617 | 20180702 | L4 | 8 | 16 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.736 | 1.768 | 0.508 | 1.228 | 1.26 | -0.00088172 | 0.031024 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR47 | completely in noise of fluorometer, negative chl so toss; Taylor previously QC'ed
EN617 | 20180702 | L4 | 8 | 16 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.293 | 1.137 | 0.508 | 0.785 | 0.629 | 0.0042984 | 0.010749 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR48 | completely in noise of fluorometer, signal basically zero; Taylor previously QC'ed
EN617 | 20180702 | L4 | 8 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 49.98 | 30.86 | 0.6432 | 49.3368 | 30.2168 | 0.36333 | 0.13519 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L4 | 8 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 52.93 | 31.85 | 0.6432 | 52.2868 | 31.2068 | 0.40058 | 0.11428 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L4 | 8 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 44.53 | 26.66 | 0.6432 | 43.8868 | 26.0168 | 0.33958 | 0.089651 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L4 | 8 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 47.63 | 29.01 | 0.6432 | 46.9868 | 28.3668 | 0.35383 | 0.11417 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 32.75 | 24.54 | 0.67417 | 32.0758 | 23.8658 | 0.15601 | 0.23773 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 33.33 | 25.26 | 0.67417 | 32.6558 | 24.5858 | 0.15335 | 0.25227 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 3 |  |  | a | 270 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 27.25 | 20.62 | 0.67417 | 26.5758 | 19.9458 | 0.13532 | 0.21812 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 28.42 | 20.91 | 0.6432 | 27.7768 | 20.2668 | 0.14271 | 0.19165 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 5 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 40.38 | 26.47 | 0.53383 | 39.8462 | 25.9362 | 0.38327 | 0.23718 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR49
EN617 | 20180702 | L5 | 9 | 5 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 42.21 | 27.31 | 0.53383 | 41.6762 | 26.7762 | 0.41055 | 0.22999 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR50
EN617 | 20180702 | L5 | 9 | 5 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 14.29 | 9.004 | 0.53383 | 13.7562 | 8.4702 | 0.14565 | 0.056976 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR55
EN617 | 20180702 | L5 | 9 | 5 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.72 | 8.939 | 0.53383 | 13.1862 | 8.4052 | 0.13174 | 0.069336 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR56
EN617 | 20180702 | L5 | 9 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 152.9 | 102 | 0.6432 | 152.2568 | 101.3568 | 0.96724 | 0.70496 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 155.1 | 102.6 | 0.6742 | 154.4258 | 101.9258 | 0.99764 | 0.68394 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 7 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 43.84 | 27.48 | 0.53383 | 43.3062 | 26.9462 | 0.45078 | 0.19383 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR51
EN617 | 20180702 | L5 | 9 | 7 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 49.72 | 31.09 | 0.53383 | 49.1862 | 30.5562 | 0.51333 | 0.21764 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR52
EN617 | 20180702 | L5 | 9 | 7 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.226 | 4.785 | 0.53383 | 6.6922 | 4.2512 | 0.067259 | 0.034439 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR57 | poor chl replication (40% range) but getting close to noise of fluorometer; Taylor previously QC'ed
EN617 | 20180702 | L5 | 9 | 7 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 5.547 | 3.927 | 0.53383 | 5.0132 | 3.3932 | 0.044637 | 0.036535 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR58 | poor chl replication (40% range) but getting close to noise of fluorometer; Taylor previously QC'ed
EN617 | 20180702 | L5 | 9 | 7 |  |  | a | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 255.3 | 159.6 | 0.6432 | 254.6568 | 158.9568 | 1.844 | 0.81518 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 7 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 263.4 | 167.6 | 0.6432 | 262.7568 | 166.9568 | 1.8524 | 0.9504 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 7 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 237.8 | 149.6 | 0.67417 | 237.1258 | 148.9258 | 1.7054 | 0.79466 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 7 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 246.5 | 157.1 | 0.67417 | 245.8258 | 156.4258 | 1.7286 | 0.89737 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 9 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 201.3 | 125.6 | 0.6432 | 200.6568 | 124.9568 | 1.4637 | 0.63398 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 9 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 199.5 | 124.1 | 0.67417 | 198.8258 | 123.4258 | 1.4579 | 0.61408 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 11 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 18.54 | 11.54 | 0.53383 | 18.0062 | 11.0062 | 0.19288 | 0.070415 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR53
EN617 | 20180702 | L5 | 9 | 11 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 16.17 | 9.985 | 0.53383 | 15.6362 | 9.4512 | 0.17042 | 0.055673 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR54
EN617 | 20180702 | L5 | 9 | 11 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.977 | 1.892 | 0.53383 | 2.4432 | 1.3582 | 0.029896 | 0.0025945 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR59 | extremely poor phaeo replication (155% range), in noise of fluorometer but 20a phaeo seems false;Taylor previously QC'ed
EN617 | 20180702 | L5 | 9 | 11 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.858 | 2.865 | 0.53383 | 3.3242 | 2.3312 | 0.027361 | 0.028406 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR60 | extremely poor phaeo replication (155% range), in noise of fluorometer but 20a phaeo seems false; Taylor previously QC'ed
EN617 | 20180702 | L5 | 9 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 92.86 | 56.08 | 0.6432 | 92.2168 | 55.4368 | 0.69892 | 0.21568 | 20181101 | 9-52 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 94.57 | 57.15 | 0.6432 | 93.9268 | 56.5068 | 0.71108 | 0.22118 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 78.3 | 48.33 | 0.6742 | 77.6258 | 47.6558 | 0.56951 | 0.21672 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L5 | 9 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 74.03 | 45.1 | 0.6432 | 73.3868 | 44.4568 | 0.54975 | 0.18371 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.72 | 14.97 | 0.6742 | 19.0458 | 14.2958 | 0.090263 | 0.14559 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.3 | 14.86 | 0.6742 | 18.6258 | 14.1858 | 0.084372 | 0.14967 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 23.84 | 17.74 | 0.6432 | 23.1968 | 17.0968 | 0.11592 | 0.16615 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 23.83 | 18.09 | 0.6742 | 23.1558 | 17.4158 | 0.10908 | 0.17825 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 27.29 | 19.83 | 0.6432 | 26.6468 | 19.1868 | 0.14176 | 0.17479 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 28.25 | 20.12 | 0.6742 | 27.5758 | 19.4458 | 0.15449 | 0.16633 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 43.91 | 28.71 | 0.6742 | 43.2358 | 28.0358 | 0.28884 | 0.1737 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 42.11 | 29.14 | 0.6432 | 41.4668 | 28.4968 | 0.24646 | 0.22368 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 10 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 83.63 | 51.06 | 0.53383 | 83.0962 | 50.5262 | 0.89743 | 0.31127 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR63
EN617 | 20180702 | L6 | 12 | 10 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 84.62 | 52.21 | 0.53383 | 84.0862 | 51.6762 | 0.89302 | 0.34319 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR64
EN617 | 20180702 | L6 | 12 | 10 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 26.07 | 16.15 | 0.508 | 25.562 | 15.642 | 0.27333 | 0.10086 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR69
EN617 | 20180702 | L6 | 12 | 10 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 23.5 | 14.61 | 0.508 | 22.992 | 14.102 | 0.24495 | 0.092398 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR70
EN617 | 20180702 | L6 | 12 | 10 |  |  | a | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 306.8 | 194.5 | 0.6432 | 306.1568 | 193.8568 | 2.1638 | 1.0792 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 10 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 323.6 | 205.6 | 0.6432 | 322.9568 | 204.9568 | 2.2817 | 1.1591 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 10 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 248.5 | 158.4 | 0.6432 | 247.8568 | 157.7568 | 1.7422 | 0.90617 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | typo in lab NB says C2N10<10a but feel confident it's C12N10<10a; QF was 2 but confident in this switch so changed it to a 1
EN617 | 20180702 | L6 | 12 | 10 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 256 | 162 | 0.6742 | 255.3258 | 161.3258 | 1.8176 | 0.89068 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | lab notebook said Cast2Niskin18 rep b <10 but C2N18 doesn't exist!! Moved to C12N10 <10b; QF was a 2 but confident in switch so changed to 1
EN617 | 20180702 | L6 | 12 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 234.5 | 148.8 | 0.6742 | 233.8258 | 148.1258 | 1.6285 | 0.81527 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 232.2 | 147.7 | 0.6742 | 231.5258 | 147.0258 | 1.6057 | 0.81992 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 16 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 33.33 | 20.65 | 0.53383 | 32.7962 | 20.1162 | 0.34938 | 0.13184 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR61
EN617 | 20180702 | L6 | 12 | 16 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 34.02 | 21.51 | 0.53383 | 33.4862 | 20.9762 | 0.3447 | 0.1571 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR62
EN617 | 20180702 | L6 | 12 | 16 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.829 | 4.06 | 0.508 | 6.321 | 3.552 | 0.076297 | 0.0086752 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR67 | very poor phaeo replication (88% range); Taylor previously QC'ed
EN617 | 20180702 | L6 | 12 | 16 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.009 | 4.521 | 0.508 | 6.501 | 4.013 | 0.068554 | 0.027446 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR68 | very poor phaeo replication (88% range); phaeo much higher than >5's and any of the other >20's; Taylor previously QC'ed; phaeo is not higher than the other 5s, QF was 2, but changed to 2
EN617 | 20180702 | L6 | 12 | 16 |  |  | a | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 175.7 | 108.3 | 0.6742 | 175.0258 | 107.6258 | 1.2987 | 0.50176 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 16 |  |  | b | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 171.7 | 105.8 | 0.6742 | 171.0258 | 105.1258 | 1.2654 | 0.48714 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 16 |  |  | a | 258 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 142.7 | 87.96 | 0.6742 | 142.0258 | 87.2858 | 1.1692 | 0.44944 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | lab notebook said cas2niskin16 for two set of values, moved these values to cas12niskin16
EN617 | 20180702 | L6 | 12 | 16 |  |  | b | 287 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 147.6 | 93.63 | 0.6742 | 146.9258 | 92.9558 | 1.0363 | 0.51333 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | lab notebook said cas2niskin16 for two set of values, moved these values to cas12niskin16
EN617 | 20180702 | L6 | 12 | 20 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 9.237 | 5.767 | 0.53383 | 8.7032 | 5.2332 | 0.095612 | 0.029577 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR65
EN617 | 20180702 | L6 | 12 | 20 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.984 | 5.113 | 0.53383 | 7.4502 | 4.5792 | 0.079107 | 0.030437 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR66
EN617 | 20180702 | L6 | 12 | 20 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.197 | 1.545 | 0.508 | 1.689 | 1.037 | 0.017965 | 0.0068423 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR71 | Taylor previously QC'ed
EN617 | 20180702 | L6 | 12 | 20 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.115 | 1.615 | 0.508 | 1.607 | 1.107 | 0.013777 | 0.012705 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR72 | Taylor previously QC'ed
EN617 | 20180702 | L6 | 12 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 40.76 | 25.02 | 0.6742 | 40.0858 | 24.3458 | 0.2991 | 0.10256 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 42.88 | 26.48 | 0.6742 | 42.2058 | 25.8058 | 0.31164 | 0.1141 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L6 | 12 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 35.19 | 21.98 | 0.6742 | 34.5158 | 21.3058 | 0.25103 | 0.10048 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | lab notebook orignially said cas2niskin20, moved values to cas12niskin20 as 2-20 DNE
EN617 | 20180702 | L6 | 12 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 34.39 | 21.27 | 0.6742 | 33.7158 | 20.5958 | 0.24932 | 0.090477 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | lab notebook orignially said cas2niskin20, moved values to cas12niskin20 as 2-20 DNE
EN617 | 20180702 | L7 | 13 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 11.1 | 9.246 | 0.6742 | 10.4258 | 8.5718 | 0.035231 | 0.10619 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN617 | 20180702 | L7 | 13 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 12.05 | 9.701 | 0.63133 | 11.4187 | 9.0697 | 0.044637 | 0.105 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN617 | 20180702 | L7 | 13 | 3 |  |  | a | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 12.27 | 9.314 | 0.6742 | 11.5958 | 8.6398 | 0.056759 | 0.087271 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 3 |  |  | b | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 12.71 | 9.894 | 0.6742 | 12.0358 | 9.2198 | 0.054071 | 0.099629 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 18.12 | 13.88 | 0.6742 | 17.4458 | 13.2058 | 0.080571 | 0.1373 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 18.72 | 14.33 | 0.6742 | 18.0458 | 13.6558 | 0.083422 | 0.14187 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 41.64 | 29.18 | 0.6742 | 40.9658 | 28.5058 | 0.23677 | 0.23352 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 42.78 | 30.22 | 0.6742 | 42.1058 | 29.5458 | 0.23867 | 0.24878 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 100.3 | 64.98 | 0.6742 | 99.6258 | 64.3058 | 0.67118 | 0.38975 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 95.59 | 64.07 | 0.6742 | 94.9158 | 63.3958 | 0.59896 | 0.44695 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 13 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 52.97 | 33.18 | 0.53383 | 52.4362 | 32.6462 | 0.54529 | 0.23568 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR77 | odd that these >5 ratio to WSW are a lot higher at this depth compared to shallower depths of this cast but could be at base of mixed layer with 5-20um cells accumulating?
EN617 | 20180702 | L7 | 13 | 13 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 42.35 | 26.77 | 0.53383 | 41.8162 | 26.2362 | 0.42929 | 0.19834 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR78 | odd that these >5 ratio to WSW are a lot higher at this depth compared to shallower depths of this cast but could be at base of mixed layer with 5-20um cells accumulating?
EN617 | 20180702 | L7 | 13 | 13 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.04 | 9.148 | 0.53383 | 12.5062 | 8.6142 | 0.10724 | 0.098831 | 20181101 | 9-46 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR83
EN617 | 20180702 | L7 | 13 | 13 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.47 | 9.411 | 0.53383 | 12.9362 | 8.8772 | 0.11184 | 0.10052 | 20181101 | 9-46 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR84
EN617 | 20180702 | L7 | 13 | 13 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 202.8 | 130.2 | 0.6742 | 202.1258 | 129.5258 | 1.4038 | 0.77063 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 13 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 207.4 | 131 | 0.6742 | 206.7258 | 130.3258 | 1.4773 | 0.71058 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 15 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.27 | 12.27 | 0.53383 | 18.7362 | 11.7362 | 0.19288 | 0.087879 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR75 | Taylor previously QC'ed
EN617 | 20180702 | L7 | 13 | 15 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 16.46 | 10.84 | 0.53383 | 15.9262 | 10.3062 | 0.15485 | 0.091694 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR76 | Taylor previously QC'ed
EN617 | 20180702 | L7 | 13 | 15 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.328 | 4.879 | 0.53383 | 6.7942 | 4.3452 | 0.06748 | 0.036467 | 20181101 | 9-46 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR81 | poor phaeo replication (48% range); Taylor previously QC'ed
EN617 | 20180702 | L7 | 13 | 15 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.489 | 4.138 | 0.53383 | 5.9552 | 3.6042 | 0.064779 | 0.021441 | 20181101 | 9-46 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR82 | poor phaeo replication (48% range); Taylor previously QC'ed
EN617 | 20180702 | L7 | 13 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 82.1 | 53.75 | 0.6742 | 81.4258 | 53.0758 | 0.53873 | 0.33693 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 81.31 | 52.42 | 0.6742 | 80.6358 | 51.7458 | 0.54899 | 0.30472 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 18 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 10.58 | 6.445 | 0.53383 | 10.0462 | 5.9112 | 0.11394 | 0.027473 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR73 | Taylor previously QC'ed
EN617 | 20180702 | L7 | 13 | 18 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 9.147 | 5.8 | 0.53383 | 8.6132 | 5.2662 | 0.092223 | 0.033756 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR74 | Taylor previously QC'ed
EN617 | 20180702 | L7 | 13 | 18 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.047 | 1.552 | 0.53383 | 1.5132 | 1.0182 | 0.013639 | 0.010718 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR79 | poor phaeo ratio though values in general very low; Rb/Ra in noise of fluorometer; Taylor previously QC'ed
EN617 | 20180702 | L7 | 13 | 18 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.784 | 1.197 | 0.53383 | 1.2502 | 0.66317 | 0.016174 | -0.00030967 | 20181101 | 9-45 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR80 | poor phaeo ratio, this phaeo virtually zero, RB/Ra in noise of fluorometer; Taylor previously QC'ed
EN617 | 20180702 | L7 | 13 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 44 | 26.74 | 0.6742 | 43.3258 | 26.0658 | 0.32799 | 0.10205 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 42.11 | 25.53 | 0.6742 | 41.4358 | 24.8558 | 0.31506 | 0.09501 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 39.93 | 24.79 | 0.6742 | 39.2558 | 24.1158 | 0.2877 | 0.11017 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L7 | 13 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 46.87 | 28.64 | 0.6742 | 46.1958 | 27.9658 | 0.34642 | 0.11496 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L8 | 14 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.186 | 5.937 | 0.63133 | 6.5547 | 5.3057 | 0.023734 | 0.063799 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN617 | 20180702 | L8 | 14 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 8.054 | 6.463 | 0.6742 | 7.3798 | 5.7888 | 0.030233 | 0.065271 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN617 | 20180702 | L8 | 14 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 8.72 | 6.242 | 0.6742 | 8.0458 | 5.5678 | 0.047089 | 0.04477 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (34% range)
EN617 | 20180702 | L8 | 14 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 8.27 | 6.48 | 0.6742 | 7.5958 | 5.8058 | 0.034015 | 0.06177 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (34% range)
EN617 | 20180702 | L8 | 14 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 12.9 | 9.991 | 0.6742 | 12.2258 | 9.3168 | 0.055279 | 0.098431 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L8 | 14 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.2 | 10.22 | 0.63133 | 12.5687 | 9.5887 | 0.056628 | 0.10157 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L8 | 14 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.81 | 13.78 | 0.6432 | 19.1668 | 13.1368 | 0.11459 | 0.10215 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (34% range) but phaeo:chl ratio similar between samples
EN617 | 20180702 | L8 | 14 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 29.41 | 20.83 | 0.6742 | 28.7358 | 20.1558 | 0.16304 | 0.16949 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (34% range) but phaeo:chl ratio similar between samples
EN617 | 20180702 | L8 | 14 | 10 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 50.63 | 32.7 | 0.0508 | 50.5792 | 32.6492 | 0.49404 | 0.287 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR87
EN617 | 20180702 | L8 | 14 | 10 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 60.43 | 39.37 | 0.0508 | 60.3792 | 39.3192 | 0.58029 | 0.36032 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR88
EN617 | 20180702 | L8 | 14 | 10 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 11.8 | 8.636 | 0.0508 | 11.7492 | 8.5852 | 0.087181 | 0.1182 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR93 | Taylor previously QC'ed
EN617 | 20180702 | L8 | 14 | 10 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 9.273 | 6.607 | 0.0508 | 9.2222 | 6.5562 | 0.073459 | 0.083381 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR94 | Taylor previously QC'ed
EN617 | 20180702 | L8 | 14 | 10 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 126.2 | 83.04 | 0.6432 | 125.5568 | 82.3968 | 0.83454 | 0.5487 | 20181101 | 9-51 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (37% range) and this rep is lower than both <10 values at this depth, what does HPLC say?
EN617 | 20180702 | L8 | 14 | 10 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 180 | 117.2 | 0.6742 | 179.3258 | 116.5258 | 1.2143 | 0.74188 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (37% range) compare to HPLC
EN617 | 20180702 | L8 | 14 | 10 |  |  | a | 286 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 142.4 | 92.06 | 0.6742 | 141.7258 | 91.3858 | 0.96997 | 0.55881 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN617 | 20180702 | L8 | 14 | 10 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 145.7 | 94.9 | 0.6742 | 145.0258 | 94.2258 | 0.98227 | 0.59955 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN617 | 20180702 | L8 | 14 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 169 | 109.7 | 0.6742 | 168.3258 | 109.0258 | 1.1269 | 0.67186 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L8 | 14 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 178.9 | 114 | 0.6742 | 178.2258 | 113.3258 | 1.2333 | 0.63639 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L8 | 14 | 15 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 25.1 | 15.49 | 0.0508 | 25.0492 | 15.4392 | 0.26479 | 0.10455 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR89
EN617 | 20180702 | L8 | 14 | 15 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 26.57 | 16.31 | 0.0508 | 26.5192 | 16.2592 | 0.2827 | 0.10625 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR90
EN617 | 20180702 | L8 | 14 | 15 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.242 | 1.47 | 0.0508 | 2.1912 | 1.4192 | 0.021272 | 0.012679 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR95
EN617 | 20180702 | L8 | 14 | 15 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.321 | 1.581 | 0.0508 | 2.2702 | 1.5302 | 0.02039 | 0.016216 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR96
EN617 | 20180702 | L8 | 14 | 15 |  |  | a | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 72 | 45.15 | 0.6742 | 71.3258 | 44.4758 | 0.51556 | 0.22588 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L8 | 14 | 15 |  |  | b | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 76.5 | 47.29 | 0.6742 | 75.8258 | 46.6158 | 0.56087 | 0.21624 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L8 | 14 | 15 |  |  | a | 287 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 62.23 | 38.66 | 0.63133 | 61.5987 | 38.0287 | 0.45258 | 0.18139 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L8 | 14 | 15 |  |  | b | 287 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 55.47 | 35.08 | 0.6742 | 54.7958 | 34.4058 | 0.39152 | 0.18205 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L8 | 14 | 17 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 5.33 | 3.432 | 0.0508 | 5.2792 | 3.3812 | 0.052297 | 0.028589 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR85
EN617 | 20180702 | L8 | 14 | 17 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.375 | 4.182 | 0.0508 | 6.3242 | 4.1312 | 0.060426 | 0.038402 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR86
EN617 | 20180702 | L8 | 14 | 17 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.983 | 1.307 | 0.0508 | 1.9322 | 1.2562 | 0.018626 | 0.011425 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR91
EN617 | 20180702 | L8 | 14 | 17 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.381 | 1.574 | 0.0508 | 2.3302 | 1.5232 | 0.022236 | 0.014202 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR92
EN617 | 20180702 | L8 | 14 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 32.61 | 20.32 | 0.6742 | 31.9358 | 19.6458 | 0.23354 | 0.090576 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L8 | 14 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 32.44 | 20.07 | 0.63133 | 31.8087 | 19.4387 | 0.23506 | 0.085638 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L8 | 14 | 17 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 26.99 | 17.28 | 0.63133 | 26.3587 | 16.6487 | 0.18452 | 0.090156 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L8 | 14 | 17 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 23.12 | 14.64 | 0.6742 | 22.4458 | 13.9658 | 0.16114 | 0.069267 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L9 | 17 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 5.619 | 4.421 | 0.63133 | 4.9877 | 3.7897 | 0.022765 | 0.039757 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor phaeo:chl match, but very low values
EN617 | 20180702 | L9 | 17 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 5.619 | 4.613 | 0.63133 | 4.9877 | 3.9817 | 0.019117 | 0.046573 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor phaeo:chl match, but very low values
EN617 | 20180702 | L9 | 17 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 9.027 | 6.867 | 0.63133 | 8.3957 | 6.2357 | 0.041046 | 0.061831 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L9 | 17 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 9.183 | 6.743 | 0.63133 | 8.5517 | 6.1117 | 0.046367 | 0.054465 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L9 | 17 | 5 |  |  | a | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 24.79 | 17.83 | 0.6742 | 24.1158 | 17.1558 | 0.13364 | 0.15236 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN617 | 20180702 | L9 | 17 | 5 |  |  | b | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 25.44 | 17.67 | 0.63133 | 24.8087 | 17.0387 | 0.14919 | 0.13485 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN617 | 20180702 | L9 | 17 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 82.73 | 55.84 | 0.6742 | 82.0558 | 55.1658 | 0.51098 | 0.39915 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L9 | 17 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 80.25 | 54.38 | 0.6742 | 79.5758 | 53.7058 | 0.4916 | 0.39445 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L9 | 17 | 10 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 63.01 | 39.44 | 0.0508 | 62.9592 | 39.3892 | 0.64945 | 0.29283 | 20181101 | 9-46 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR97
EN617 | 20180702 | L9 | 17 | 10 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 61.42 | 38.8 | 0.0508 | 61.3692 | 38.7492 | 0.62327 | 0.3037 | 20181101 | 9-46 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR98
EN617 | 20180702 | L9 | 17 | 10 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 4.156 | 2.843 | 0.0508 | 4.1052 | 2.7922 | 0.036178 | 0.030618 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR103
EN617 | 20180702 | L9 | 17 | 10 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 4.274 | 2.775 | 0.0508 | 4.2232 | 2.7242 | 0.041303 | 0.023866 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR104
EN617 | 20180702 | L9 | 17 | 10 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 210.3 | 139.3 | 0.63133 | 209.6687 | 138.6687 | 1.3729 | 0.95505 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L9 | 17 | 10 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 215.6 | 142.8 | 0.63133 | 214.9687 | 142.1687 | 1.4077 | 0.979 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L9 | 17 | 13 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 45.29 | 28.12 | 0.0508 | 45.2392 | 28.0692 | 0.4731 | 0.19838 | 20181101 | 9-46 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR99
EN617 | 20180702 | L9 | 17 | 13 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 43.68 | 26.27 | 0.0508 | 43.6292 | 26.2192 | 0.47971 | 0.14751 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR100
EN617 | 20180702 | L9 | 17 | 13 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 4.316 | 2.759 | 0.0508 | 4.2652 | 2.7082 | 0.042901 | 0.021885 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR105
EN617 | 20180702 | L9 | 17 | 13 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 4.325 | 2.823 | 0.0508 | 4.2742 | 2.7722 | 0.041386 | 0.024931 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR106
EN617 | 20180702 | L9 | 17 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 175.3 | 112 | 0.6742 | 174.6258 | 111.3258 | 1.2029 | 0.6338 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L9 | 17 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 161 | 101.1 | 0.6313 | 160.3687 | 100.4687 | 1.1383 | 0.51929 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L9 | 17 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 140.1 | 87.38 | 0.6742 | 139.4258 | 86.7058 | 1.0018 | 0.42866 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L9 | 17 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 127.7 | 79.91 | 0.63133 | 127.0687 | 79.2787 | 0.90814 | 0.39981 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L9 | 17 | 17 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 20.3 | 12.48 | 0.0508 | 20.2492 | 12.4292 | 0.21547 | 0.081864 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR101
EN617 | 20180702 | L9 | 17 | 17 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.44 | 12.18 | 0.0508 | 19.3892 | 12.1292 | 0.20004 | 0.090117 | 20181101 | 9-39 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR102
EN617 | 20180702 | L9 | 17 | 17 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.758 | 2.178 | 0.0508 | 2.7072 | 2.1272 | 0.015981 | 0.034906 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR107 | very poor replication (75% range) close to noise of fluroometer; Taylor previously QC'ed
EN617 | 20180702 | L9 | 17 | 17 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.302 | 2.022 | 0.0508 | 3.2512 | 1.9712 | 0.035269 | 0.011887 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR108 | very poor replication (75% range) close to noise of fluroometer; Taylor previously QC'ed
EN617 | 20180702 | L9 | 17 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 81.99 | 50.18 | 0.63133 | 81.3587 | 49.5487 | 0.60448 | 0.21299 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L9 | 17 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 81.85 | 50.19 | 0.63133 | 81.2187 | 49.5587 | 0.60163 | 0.216 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L9 | 17 | 17 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 63 | 39.38 | 0.63133 | 62.3687 | 38.7487 | 0.44884 | 0.19044 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L9 | 17 | 17 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 62.12 | 38.54 | 0.6742 | 61.4458 | 37.8658 | 0.44808 | 0.17663 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.652 | 2.915 | 0.63133 | 3.0207 | 2.2837 | 0.014005 | 0.023671 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication but basically in noise of fluorometer
EN617 | 20180702 | L10 | 18 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.463 | 3.086 | 0.6742 | 2.7888 | 2.4118 | 0.007164 | 0.032626 | 20181101 | 9-49 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication but basically in noise of fluorometer
EN617 | 20180702 | L10 | 18 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.064 | 4.535 | 0.6313 | 5.4327 | 3.9037 | 0.029055 | 0.035349 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN617 | 20180702 | L10 | 18 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.25 | 4.228 | 0.6313 | 5.6187 | 3.5967 | 0.038423 | 0.020915 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN617 | 20180702 | L10 | 18 | 5 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.83 | 4.841 | 0.0508 | 6.7792 | 4.7902 | 0.054805 | 0.059788 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR109 | Taylor previously QC'ed
EN617 | 20180702 | L10 | 18 | 5 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.252 | 4.666 | 0.0508 | 6.2012 | 4.6152 | 0.0437 | 0.066706 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR110 | Taylor previously QC'ed
EN617 | 20180702 | L10 | 18 | 5 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.835 | 1.449 | 0.0508 | 1.7842 | 1.3982 | 0.010636 | 0.022812 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR115 | poor replication but signal basically zero; Taylor previously QC'ed
EN617 | 20180702 | L10 | 18 | 5 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.611 | 1.44 | 0.0508 | 1.5602 | 1.3892 | 0.0047117 | 0.028521 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR116 | poor replication but signal basically zero; Taylor previously QC'ed
EN617 | 20180702 | L10 | 18 | 5 |  |  | a | 288 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 31.2 | 21.32 | 0.63133 | 30.5687 | 20.6887 | 0.18905 | 0.15464 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 5 |  |  | b | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 32.15 | 22.17 | 0.6742 | 31.4758 | 21.4958 | 0.19163 | 0.16672 | 20181101 | 9-50 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 5 |  |  | a | 287 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 27.18 | 17.44 | 0.63133 | 26.5487 | 16.8087 | 0.18702 | 0.09319 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 5 |  |  | b | 286 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 26.93 | 17.92 | 0.63133 | 26.2987 | 17.2887 | 0.17361 | 0.11561 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 28.9 | 20.14 | 0.6313 | 28.2687 | 19.5087 | 0.16646 | 0.15539 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 29.94 | 20.64 | 0.6313 | 29.3087 | 20.0087 | 0.17673 | 0.15338 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 9 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 33.85 | 20.91 | 0.0508 | 33.7992 | 20.8592 | 0.35655 | 0.14245 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR111
EN617 | 20180702 | L10 | 18 | 9 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 34.63 | 21.29 | 0.0508 | 34.5792 | 21.2392 | 0.36757 | 0.14052 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR112
EN617 | 20180702 | L10 | 18 | 9 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.863 | 1.931 | 0.0508 | 2.8122 | 1.8802 | 0.02568 | 0.019298 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR117 | Taylor previously QC'ed
EN617 | 20180702 | L10 | 18 | 9 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.984 | 2.146 | 0.0508 | 2.9332 | 2.0952 | 0.02309 | 0.027032 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR118 | Taylor previously QC'ed
EN617 | 20180702 | L10 | 18 | 9 |  |  | a | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 114.2 | 74.67 | 0.63133 | 113.5687 | 74.0387 | 0.76168 | 0.4769 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 9 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 105.5 | 68.4 | 0.6313 | 104.8687 | 67.7687 | 0.71737 | 0.4203 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 9 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 86.43 | 57.18 | 0.63133 | 85.7987 | 56.5487 | 0.56558 | 0.38374 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 9 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 86.51 | 56.52 | 0.6313 | 85.8787 | 55.8887 | 0.57989 | 0.35835 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 117.2 | 73.74 | 0.63133 | 116.5687 | 73.1087 | 0.82586 | 0.3803 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 119.5 | 75.27 | 0.63133 | 118.8687 | 74.6387 | 0.84049 | 0.39091 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 37.33 | 23.67 | 0.6313 | 36.6987 | 23.0387 | 0.25958 | 0.12052 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 36.3 | 22.9 | 0.63133 | 35.6687 | 22.2687 | 0.25464 | 0.11276 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 16 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.801 | 2.7 | 0.0508 | 3.7502 | 2.6492 | 0.030337 | 0.033038 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR113
EN617 | 20180702 | L10 | 18 | 16 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.404 | 2.316 | 0.0508 | 3.3532 | 2.2652 | 0.029979 | 0.02421 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR114
EN617 | 20180702 | L10 | 18 | 16 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 0.977 | 0.849 | 0.0508 | 0.9262 | 0.7982 | 0.0035269 | 0.015568 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR119 | poor replication but in noise of fluorometer, signal basically zero; Taylor previously QC'ed
EN617 | 20180702 | L10 | 18 | 16 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.136 | 0.734 | 0.0508 | 1.0852 | 0.6832 | 0.011077 | 0.005267 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR120 | poor replication but in noise of fluorometer, signal basically zero; Taylor previously QC'ed
EN617 | 20180702 | L10 | 18 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 29.38 | 18.65 | 0.6313 | 28.7487 | 18.0187 | 0.2039 | 0.093376 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 25.55 | 16.37 | 0.6313 | 24.9187 | 15.7387 | 0.17444 | 0.085215 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L10 | 18 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.61 | 12.7 | 0.63133 | 18.9787 | 12.0687 | 0.13131 | 0.067802 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | somewhat poor replication (26% range)
EN617 | 20180702 | L10 | 18 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 24.88 | 15.84 | 0.6313 | 24.2487 | 15.2087 | 0.17178 | 0.079131 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | somewhat poor replication (26% range)
EN617 | 20180702 | L11 | 21 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.39 | 2.548 | 0.63133 | 2.7587 | 1.9167 | 0.016 | 0.015621 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication but within noise of flurometer, deep 150m sample basically reading of zero
EN617 | 20180702 | L11 | 21 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.229 | 2.927 | 0.6313 | 2.5977 | 2.2957 | 0.0057388 | 0.032136 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication but within noise of flurometer, deep 150m sample basically reading of zero
EN617 | 20180702 | L11 | 21 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 5.167 | 3.894 | 0.63133 | 4.5357 | 3.2627 | 0.02419 | 0.029637 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L11 | 21 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 5.192 | 3.825 | 0.63133 | 4.5607 | 3.1937 | 0.025977 | 0.026713 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L11 | 21 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 14.78 | 10.01 | 0.63133 | 14.1487 | 9.3787 | 0.090643 | 0.064088 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L11 | 21 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 14.92 | 10.19 | 0.6313 | 14.2887 | 9.5587 | 0.089883 | 0.067818 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L11 | 21 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 33.95 | 23 | 0.63133 | 33.3187 | 22.3687 | 0.20808 | 0.16096 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L11 | 21 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 33.28 | 23.04 | 0.63133 | 32.6487 | 22.4087 | 0.19459 | 0.17511 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L11 | 21 | 9 |  |  | a | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 93.61 | 61.89 | 0.63133 | 92.9787 | 61.2587 | 0.60907 | 0.41215 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L11 | 21 | 9 |  |  | b | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 94.17 | 62.52 | 0.63133 | 93.5387 | 61.8887 | 0.60772 | 0.424 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L11 | 21 | 11 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 52.68 | 32.59 | 0.0508 | 52.6292 | 32.5392 | 0.55356 | 0.22485 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR125
EN617 | 20180702 | L11 | 21 | 11 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 52.28 | 31.54 | 0.0508 | 52.2292 | 31.4892 | 0.57147 | 0.18183 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR126
EN617 | 20180702 | L11 | 21 | 11 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 5.481 | 3.286 | 0.0508 | 5.4302 | 3.2352 | 0.060481 | 0.016913 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR131 | very poor phaeo replication (70% range); Taylor previously QC'ed
EN617 | 20180702 | L11 | 21 | 11 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.314 | 4.138 | 0.0508 | 6.2632 | 4.0872 | 0.059957 | 0.037818 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR132 | very poor phaeo replication (70% range); Taylor previously QC'ed
EN617 | 20180702 | L11 | 21 | 11 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 158.6 | 98.06 | 0.63133 | 157.9687 | 97.4287 | 1.1706 | 0.46499 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L11 | 21 | 11 |  |  | b | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 161.9 | 99.93 | 0.63133 | 161.2687 | 99.2987 | 1.1941 | 0.46709 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L11 | 21 | 16 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 15.96 | 9.966 | 0.0508 | 15.9092 | 9.9152 | 0.16516 | 0.072036 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR123
EN617 | 20180702 | L11 | 21 | 16 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 15.74 | 9.771 | 0.0508 | 15.6892 | 9.7202 | 0.16447 | 0.06806 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR124
EN617 | 20180702 | L11 | 21 | 16 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 4.976 | 3.085 | 0.0508 | 4.9252 | 3.0342 | 0.052104 | 0.020481 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR129 | poor phaeo replication (56% range); Taylor previously QC'ed
EN617 | 20180702 | L11 | 21 | 16 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 5.289 | 3.599 | 0.0508 | 5.2382 | 3.5482 | 0.046566 | 0.038315 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR130 | poor phaeo replication (56% range); Taylor previously QC'ed
EN617 | 20180702 | L11 | 21 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 69.24 | 42.73 | 0.63133 | 68.6087 | 42.0987 | 0.50376 | 0.19079 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | unsure if correct values–why is this here? QF marked at 1 but changed to 2 because of this comment. Should we double check lab NB?
EN617 | 20180702 | L11 | 21 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 70.76 | 43.7 | 0.63133 | 70.1287 | 43.0687 | 0.51421 | 0.19634 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L11 | 21 | 20 |  |  | a | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.702 | 4.787 | 0.0508 | 7.6512 | 4.7362 | 0.08032 | 0.032981 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR121
EN617 | 20180702 | L11 | 21 | 20 |  |  | b | 200 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 8.316 | 5.142 | 0.0508 | 8.2652 | 5.0912 | 0.087456 | 0.034337 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR122
EN617 | 20180702 | L11 | 21 | 20 |  |  | a | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.879 | 1.828 | 0.0508 | 2.8282 | 1.7772 | 0.028959 | 0.013556 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR127 | poor replication but basically in noise of fluorometer; Taylor previously QC'ed
EN617 | 20180702 | L11 | 21 | 20 |  |  | b | 200 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.255 | 1.635 | 0.0508 | 2.2042 | 1.5842 | 0.017083 | 0.020814 | 20181101 | 9-40 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR128 | poor replication but basically in noise of fluorometer; Taylor previously QC'ed
EN617 | 20180702 | L11 | 21 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 37.61 | 23.51 | 0.63133 | 36.9787 | 22.8787 | 0.26794 | 0.10952 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L11 | 21 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 36.89 | 22.61 | 0.63133 | 36.2587 | 21.9787 | 0.27136 | 0.091249 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L11 | 21 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 23.12 | 15.15 | 0.63133 | 22.4887 | 14.5187 | 0.15145 | 0.08808 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | L11 | 21 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 23.82 | 15.1 | 0.63133 | 23.1887 | 14.4687 | 0.1657 | 0.073003 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | MVCO | 35 | 2 |  |  | a | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 200.1 | 125.3 | 0.63133 | 199.4687 | 124.6687 | 1.4413 | 0.64429 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | MVCO | 35 | 2 |  |  | b | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 229 | 144.1 | 0.63133 | 228.3687 | 143.4687 | 1.6302 | 0.76152 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | MVCO | 35 | 4 |  |  | a | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 220.1 | 138 | 0.63133 | 219.4687 | 137.3687 | 1.5819 | 0.71608 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | MVCO | 35 | 4 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 218.1 | 138 | 0.63133 | 217.4687 | 137.3687 | 1.5488 | 0.75727 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | MVCO | 35 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 229.5 | 145.6 | 0.63133 | 228.8687 | 144.9687 | 1.5943 | 0.79739 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | MVCO | 35 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 221 | 140.7 | 0.63133 | 220.3687 | 140.0687 | 1.5259 | 0.78496 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | MVCO | 35 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 140.1 | 89.11 | 0.63133 | 139.4687 | 88.4787 | 0.96895 | 0.49079 | 20181101 | 9-47 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN617 | 20180702 | MVCO | 35 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 153 | 96.81 | 0.63133 | 152.3687 | 96.1787 | 1.0678 | 0.51901 | 20181101 | 9-48 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR32 | 20181112 | L1 | 1 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 195.2 | 117.3 | 0.79933 | 194.4007 | 116.5007 | 4.4409 | 1.3252 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181112 | L1 | 1 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 180.1 | 111.4 | 0.95867 | 179.1413 | 110.4413 | 3.9165 | 1.5498 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181112 | L1 | 1 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 204.2 | 126.4 | 0.95867 | 203.2413 | 125.4413 | 4.4352 | 1.7734 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181112 | L1 | 1 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 152.8 | 95.78 | 0.95867 | 151.8413 | 94.8213 | 3.2506 | 1.4425 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR32 | 20181112 | L1 | 1 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 98.48 | 60.86 | 0.79933 | 97.6807 | 60.0607 | 0.71488 | 0.27601 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | poor replication, judging from full cast something clearly wrong with this sample, even if forgot to write in lab notebook dilute 2a:1s still only get chl=2
AR32 | 20181112 | L1 | 1 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 81.49 | 52.96 | 0.95867 | 80.5313 | 52.0013 | 0.54215 | 0.31578 | 20181101 | 9-57 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (39% range)
AR32 | 20181112 | L1 | 1 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 118.6 | 75.92 | 0.95867 | 117.6413 | 74.9613 | 0.81103 | 0.42569 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (39% range)
AR32 | 20181112 | L3 | 2 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 185.7 | 114.8 | 0.95867 | 184.7413 | 113.8413 | 4.0419 | 1.5926 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181112 | L3 | 2 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 193.7 | 120.1 | 0.95867 | 192.7413 | 119.1413 | 4.1958 | 1.701 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181112 | L3 | 2 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 171.6 | 109.4 | 0.95867 | 170.6413 | 108.4413 | 3.5459 | 1.8213 | 20181101 | 9-57 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication (22% range), dilution often not great replication
AR32 | 20181112 | L3 | 2 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 205 | 126.8 | 0.95867 | 204.0413 | 125.8413 | 4.458 | 1.7704 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication (22% range), dilution often not great replication
AR32 | 20181112 | L3 | 2 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 142.9 | 88.55 | 0.79933 | 142.1007 | 87.7507 | 3.0984 | 1.2448 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181112 | L3 | 2 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 150.8 | 93.19 | 0.95867 | 149.8413 | 92.2313 | 3.2842 | 1.2807 | 20181101 | 9-57 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181112 | L3 | 2 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 78.31 | 49.65 | 0.95867 | 77.3513 | 48.6913 | 0.54462 | 0.2587 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication (25% range)
AR32 | 20181112 | L3 | 2 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 102.6 | 65.38 | 0.95867 | 101.6413 | 64.4213 | 0.70728 | 0.35555 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication (25% range)
AR32 | 20181112 | L5 | 3 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 142.3 | 100.8 | 0.95867 | 141.3413 | 99.8413 | 0.78861 | 0.85858 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181112 | L5 | 3 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 137.9 | 97.31 | 0.79933 | 137.1007 | 96.5107 | 0.77132 | 0.82093 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181112 | L5 | 3 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 288 | 183.8 | 0.79933 | 287.2007 | 183.0007 | 1.9801 | 1.0391 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181112 | L5 | 3 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 339 | 215.9 | 0.95867 | 338.0413 | 214.9413 | 2.3392 | 1.2069 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181112 | L5 | 3 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 216.7 | 138.1 | 0.95867 | 215.7413 | 137.1413 | 1.4936 | 0.76897 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181112 | L5 | 3 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 237.6 | 150.1 | 0.79933 | 236.8007 | 149.3007 | 1.6627 | 0.80045 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181112 | L5 | 3 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 120.5 | 77.22 | 0.79933 | 119.7007 | 76.4207 | 0.82244 | 0.43836 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181112 | L5 | 3 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 105.2 | 67.43 | 0.79933 | 104.4007 | 66.6307 | 0.71773 | 0.38155 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L6 | 4 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 95.94 | 73.13 | 0.79933 | 95.1407 | 72.3307 | 0.43345 | 0.75987 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication (24% range) but phaeo good match
AR32 | 20181113 | L6 | 4 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 110.7 | 81.53 | 0.79933 | 109.9007 | 80.7307 | 0.55431 | 0.7776 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication (24% range) but phaeo good match
AR32 | 20181113 | L6 | 4 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 167.1 | 114 | 0.79933 | 166.3007 | 113.2007 | 1.009 | 0.85856 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L6 | 4 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 163.8 | 112.2 | 0.95867 | 162.8413 | 111.2413 | 0.98054 | 0.85474 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L6 | 4 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 209 | 132.5 | 0.95867 | 208.0413 | 131.5413 | 1.4537 | 0.71648 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L6 | 4 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 222.3 | 144.1 | 0.79933 | 221.5007 | 143.3007 | 1.486 | 0.87818 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L6 | 4 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 156.2 | 96.48 | 0.79933 | 155.4007 | 95.6807 | 3.4045 | 1.3311 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR32 | 20181113 | L6 | 4 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 143.5 | 90.03 | 0.95867 | 142.5413 | 89.0713 | 3.0482 | 1.3603 | 20181101 | 9-56 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR32 | 20181113 | L6 | 4 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 172.6 | 103.8 | 0.79933 | 171.8007 | 103.0007 | 3.9222 | 1.1758 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L6 | 4 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 165.9 | 103 | 0.79933 | 165.1007 | 102.2007 | 3.5858 | 1.4726 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L6 | 4 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 320.8 | 202.5 | 0.79933 | 320.0007 | 201.7007 | 2.248 | 1.0797 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L6 | 4 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 363.5 | 227.3 | 0.79933 | 362.7007 | 226.5007 | 2.5882 | 1.1487 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L6 | 4 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 120.4 | 76.99 | 0.79933 | 119.6007 | 76.1907 | 0.82491 | 0.4321 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication, could this Rb possibly be the prvious Ra, and then this Ra is possibly the actual sample's Rb
AR32 | 20181113 | L6 | 4 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 70.96 | 46.31 | 0.79933 | 70.1607 | 45.5107 | 0.46842 | 0.28242 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication
AR32 | 20181113 | L7 | 5 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 32.08 | 25.66 | 0.79933 | 31.2807 | 24.8607 | 0.122 | 0.28816 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L7 | 5 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 28.06 | 22.16 | 0.79933 | 27.2607 | 21.3607 | 0.11212 | 0.2403 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L7 | 5 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 18.17 | 13.74 | 0.79933 | 17.3707 | 12.9407 | 0.084182 | 0.12931 | 20181101 | 9-55 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L7 | 5 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.08 | 14.51 | 0.79933 | 18.2807 | 13.7107 | 0.086842 | 0.13936 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L7 | 5 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 62.75 | 44.99 | 0.79933 | 61.9507 | 44.1907 | 0.33749 | 0.39158 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L7 | 5 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 60.31 | 41.71 | 0.79933 | 59.5107 | 40.9107 | 0.35345 | 0.3215 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L7 | 5 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 159.2 | 97.2 | 0.79933 | 158.4007 | 96.4007 | 3.5345 | 1.2368 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | poor chl replication (30% range), since this is dilution I'd say this one is more questionable since other Rb is <400
AR32 | 20181113 | L7 | 5 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 371.5 | 234.1 | 0.79933 | 370.7007 | 233.3007 | 2.611 | 1.2381 | 20181101 | 9-57 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (30% range), I'd say this one is more reliable since Rb is <400 and other is dilution
AR32 | 20181113 | L7 | 5 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 241.3 | 148 | 0.95867 | 240.3413 | 147.0413 | 5.3189 | 1.9589 | 20181101 | 9-57 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication, both samples dilution
AR32 | 20181113 | L7 | 5 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 158.9 | 97.3 | 0.79933 | 158.1007 | 96.5007 | 3.5117 | 1.2645 | 20181101 | 9-55 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication, both samples dilution
AR32 | 20181113 | L7 | 5 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 185.1 | 111.9 | 0.79933 | 184.3007 | 111.1007 | 4.173 | 1.3259 | 20181101 | 9-55 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L7 | 5 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 196.8 | 118.3 | 0.79933 | 196.0007 | 117.5007 | 4.4751 | 1.3405 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L7 | 5 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 185.9 | 116.7 | 0.79933 | 185.1007 | 115.9007 | 1.315 | 0.59716 | 20181101 | 9-55 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR32 | 20181113 | L7 | 5 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 229 | 146.5 | 0.79933 | 228.2007 | 145.7007 | 1.5677 | 0.83607 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR32 | 20181113 | L10 | 6 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 16.29 | 10.88 | 0.79933 | 15.4907 | 10.0807 | 0.1028 | 0.063508 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (49% range)
AR32 | 20181113 | L10 | 6 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 11.49 | 8.237 | 0.79933 | 10.6907 | 7.4377 | 0.061816 | 0.060892 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (49% range)
AR32 | 20181113 | L10 | 6 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 12.79 | 9.304 | 0.79933 | 11.9907 | 8.5047 | 0.066243 | 0.074068 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication (chl 27% range)
AR32 | 20181113 | L10 | 6 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.32 | 10.67 | 0.79933 | 12.5207 | 9.8707 | 0.050357 | 0.11249 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication (chl 27% range)
AR32 | 20181113 | L10 | 6 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.19 | 10.19 | 0.79933 | 12.3907 | 9.3907 | 0.057008 | 0.09792 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L10 | 6 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 15.02 | 11.52 | 0.79933 | 14.2207 | 10.7207 | 0.066509 | 0.11036 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L10 | 6 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 40.07 | 27.55 | 0.79933 | 39.2707 | 26.7507 | 0.23791 | 0.20342 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | cruise log "volume leakage on filter apparatus" for both replicates, , vol filt says 290 on logsheet, though this value higher, still don't trust how much volume leaked
AR32 | 20181113 | L10 | 6 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.62 | 10.26 | 0.79933 | 12.8207 | 9.4607 | 0.063849 | 0.092234 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | cruise log "volume leakage on filter apparatus" for both replicates, something clearly wrong with this on, vol filt says 290 on logsheet
AR32 | 20181113 | L10 | 6 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 259.4 | 160.5 | 0.79933 | 258.6007 | 159.7007 | 1.8794 | 0.7554 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L10 | 6 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 248.5 | 156.7 | 0.79933 | 247.7007 | 155.9007 | 1.7444 | 0.82762 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L10 | 6 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 233.1 | 146.1 | 0.79933 | 232.3007 | 145.3007 | 1.6532 | 0.74396 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L10 | 6 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 254.4 | 158.8 | 0.79933 | 253.6007 | 158.0007 | 1.8167 | 0.79006 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L10 | 6 | 11 |  |  | c | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 285.4 | 179.5 | 0.79933 | 284.6007 | 178.7007 | 2.0124 | 0.93584 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L10 | 6 | 11 |  |  | d | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 263.5 | 164.4 | 0.79933 | 262.7007 | 163.6007 | 1.8832 | 0.81594 | 20181101 | 9-55 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L11 | 7 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 4.703 | 3.962 | 0.79933 | 3.9037 | 3.1627 | 0.014081 | 0.038097 | 20181101 | 9-55 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR32 | 20181113 | L11 | 7 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 4.254 | 3.272 | 0.79933 | 3.4547 | 2.4727 | 0.018661 | 0.022134 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR32 | 20181113 | L11 | 7 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.928 | 5.952 | 0.79933 | 7.1287 | 5.1527 | 0.037549 | 0.04746 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication
AR32 | 20181113 | L11 | 7 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 4.705 | 4.105 | 0.79933 | 3.9057 | 3.3057 | 0.011402 | 0.043136 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | could this Rb be the previous tube's Ra reading? Poor replication
AR32 | 20181113 | L11 | 7 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 8.718 | 7.557 | 0.79933 | 7.9187 | 6.7577 | 0.022062 | 0.089427 | 20181101 | 9-55 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR32 | 20181113 | L11 | 7 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.362 | 4.765 | 0.79933 | 5.5627 | 3.9657 | 0.030347 | 0.035079 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR32 | 20181113 | L11 | 7 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.091 | 6.191 | 0.79933 | 6.2917 | 5.3917 | 0.017102 | 0.07185 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (50% range)
AR32 | 20181113 | L11 | 7 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.952 | 6.441 | 0.79933 | 7.1527 | 5.6417 | 0.028713 | 0.064364 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (50% range)
AR32 | 20181113 | L11 | 7 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.985 | 3.485 | 0.79933 | 3.1857 | 2.6857 | 0.0095013 | 0.034807 | 20181101 | 9-55 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | was this acidied for the Rb too? Very poor chl replication (72% range)
AR32 | 20181113 | L11 | 7 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 5.612 | 4.542 | 0.79933 | 4.8127 | 3.7427 | 0.020333 | 0.041414 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication (72% range) something seems off with other sample
AR32 | 20181113 | L11 | 7 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 202.7 | 126.9 | 0.79933 | 201.9007 | 126.1007 | 1.4404 | 0.64002 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR32 | 20181113 | L11 | 7 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 231 | 143.7 | 0.79933 | 230.2007 | 142.9007 | 1.6589 | 0.69866 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR32 | 20181113 | L11 | 7 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 126.7 | 81.87 | 0.79933 | 125.9007 | 81.0707 | 0.85189 | 0.48562 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (42% range)
AR32 | 20181113 | L11 | 7 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 189.1 | 120.3 | 0.79933 | 188.3007 | 119.5007 | 1.3074 | 0.66415 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | poor replication and this value almost as high as WSW
AR32 | 20181113 | L9 | 9 | 1 |  |  | a | 280 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.6 | 11.4 | 0.79933 | 12.8007 | 10.6007 | 0.043299 | 0.13784 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (49%0
AR32 | 20181113 | L9 | 9 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 15.6 | 11.82 | 0.79933 | 14.8007 | 11.0207 | 0.07183 | 0.10999 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | logsheet says chl mislabeled as C8 but truly C9, poor chl replication (49% range) student cruise-cruise logsheet notes filter leaked ~5ml people often underestimate volume so say 10ml, 9-1-2 in lab notebook,
AR32 | 20181113 | L9 | 9 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 55.81 | 41.47 | 0.79933 | 55.0107 | 40.6707 | 0.2725 | 0.39849 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L9 | 9 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 52.87 | 38.1 | 0.79933 | 52.0707 | 37.3007 | 0.28067 | 0.33472 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L9 | 9 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 320.1 | 195.3 | 0.79933 | 319.3007 | 194.5007 | 2.3715 | 0.83736 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L9 | 9 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 322.4 | 198.2 | 0.79933 | 321.6007 | 197.4007 | 2.3601 | 0.89661 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L9 | 9 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 161.2 | 101.8 | 0.79933 | 160.4007 | 101.0007 | 1.1288 | 0.53756 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L9 | 9 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 180.3 | 113.6 | 0.79933 | 179.5007 | 112.8007 | 1.2675 | 0.59352 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L4 | 11 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 320.3 | 183.7 | 0.79933 | 319.5007 | 182.9007 | 2.5958 | 0.42175 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor phaeo replication (87% range) all phaeo values of water column this station and surround station have phaeo >1.0 this seems odd, but L5 bottom water seems to be potential diff water mass
AR32 | 20181113 | L4 | 11 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 317.6 | 205.3 | 0.79933 | 316.8007 | 204.5007 | 2.134 | 1.2399 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor phaeo replication (87%)'
AR32 | 20181113 | L4 | 11 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 374.8 | 241.8 | 0.79933 | 374.0007 | 241.0007 | 2.5274 | 1.4487 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR32 | 20181113 | L4 | 11 | 3 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 345.8 | 222.7 | 0.79933 | 345.0007 | 221.9007 | 2.3803 | 1.3449 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | "~2ml leaked out" probably more than that so make it 5ml?
AR32 | 20181113 | L4 | 11 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 325.6 | 209.4 | 0.79933 | 324.8007 | 208.6007 | 2.2081 | 1.2334 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR32 | 20181113 | L4 | 11 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 393.9 | 259.6 | 0.79933 | 393.1007 | 258.8007 | 2.5521 | 1.7177 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR32 | 20181113 | L4 | 11 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 212.1 | 134.6 | 0.79933 | 211.3007 | 133.8007 | 1.4727 | 0.73475 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | very poor replication (99% range) could volume have been poored unevenly between the 2 replicates?
AR32 | 20181113 | L4 | 11 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 76.84 | 50.79 | 0.79933 | 76.0407 | 49.9907 | 0.49502 | 0.32973 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | very poor replication (99% range) could volume have been poored unevenly between the 2 replicates?
AR32 | 20181113 | L2 | 12 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 349.7 | 220.5 | 0.79933 | 348.9007 | 219.7007 | 7.3654 | 3.5085 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L2 | 12 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 337.4 | 211.3 | 0.79933 | 336.6007 | 210.5007 | 7.1887 | 3.2299 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L2 | 12 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 292.8 | 180.7 | 0.79933 | 292.0007 | 179.9007 | 6.3906 | 2.5135 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L2 | 12 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 299.5 | 186.6 | 0.79933 | 298.7007 | 185.8007 | 6.4362 | 2.7599 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR32 | 20181113 | L2 | 12 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 297.1 | 176.9 | 0.79933 | 296.3007 | 176.1007 | 6.8524 | 1.8636 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | cruise logsheet "both had leaks"
AR32 | 20181113 | L2 | 12 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 288.4 | 176.6 | 0.79933 | 287.6007 | 175.8007 | 6.3735 | 2.3276 | 20181101 | 9-53 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | cruise logsheet "both had leaks"
AR32 | 20181113 | L2 | 12 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 122.7 | 80.53 | 0.79933 | 121.9007 | 79.7307 | 0.80134 | 0.51406 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1
AR32 | 20181113 | L2 | 12 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 140.2 | 88.37 | 0.79933 | 139.4007 | 87.5707 | 0.98491 | 0.45984 | 20181101 | 9-54 |  | 3 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1
AR31A | 20181021 | L6 | 1 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 34.42 | 28.66 | 0.95867 | 33.4613 | 27.7013 | 0.10946 | 0.34756 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L6 | 1 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 33.66 | 27.77 | 0.95867 | 32.7013 | 26.8113 | 0.11193 | 0.33041 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L6 | 1 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 39.23 | 28.46 | 0.95867 | 38.2713 | 27.5013 | 0.20466 | 0.24906 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR31A | 20181021 | L6 | 1 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 34.65 | 25.87 | 0.95867 | 33.6913 | 24.9113 | 0.16684 | 0.24415 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR31A | 20181021 | L6 | 1 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 71.82 | 49.57 | 0.95867 | 70.8613 | 48.6113 | 0.42281 | 0.37919 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L6 | 1 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 69.63 | 48.26 | 0.95867 | 68.6713 | 47.3013 | 0.40609 | 0.3743 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L6 | 1 | 7 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 304.3 | 194.2 | 0.95867 | 303.3413 | 193.2413 | 2.1289 | 1.1152 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L6 | 1 | 7 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 324.5 | 204.3 | 0.95867 | 323.5413 | 203.3413 | 2.3242 | 1.0894 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L6 | 1 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 296.8 | 190.7 | 0.95867 | 295.8413 | 189.7413 | 2.0162 | 1.1142 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L6 | 1 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 308.2 | 195.2 | 0.95867 | 307.2413 | 194.2413 | 2.1473 | 1.0573 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L6 | 1 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 304.7 | 193 | 0.95867 | 303.7413 | 192.0413 | 2.1226 | 1.0457 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L6 | 1 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 305.8 | 197.1 | 0.95867 | 304.8413 | 196.1413 | 2.0656 | 1.1704 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L6 | 1 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 290.9 | 185.1 | 0.95867 | 289.9413 | 184.1413 | 2.0105 | 1.0275 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | next 4 stations sampled had somethingwrong with <10 filters, was there a problem with this cast?
AR31A | 20181021 | L6 | 1 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 264 | 166.7 | 0.95867 | 263.0413 | 165.7413 | 1.849 | 0.88546 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | next 4 stations sampled had somethingwrong with <10 filters, was there a problem with this cast?
AR31A | 20181021 | L5 | 2 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 65.83 | 50.47 | 0.79933 | 65.0307 | 49.6707 | 0.29188 | 0.52759 | 20181101 | 9-54 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L5 | 2 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 62.03 | 47.56 | 0.95867 | 61.0713 | 46.6013 | 0.27497 | 0.49387 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L5 | 2 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 82.29 | 56.89 | 0.95867 | 81.3313 | 55.9313 | 0.48267 | 0.44009 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L5 | 2 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 81.77 | 57.05 | 0.95867 | 80.8113 | 56.0913 | 0.46975 | 0.45566 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L5 | 2 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 164.2 | 108.1 | 0.95867 | 163.2413 | 107.1413 | 1.0661 | 0.70158 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L5 | 2 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 176.1 | 113.8 | 0.95867 | 175.1413 | 112.8413 | 1.1839 | 0.6778 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L5 | 2 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 283.2 | 181.2 | 0.95867 | 282.2413 | 180.2413 | 1.9383 | 1.0354 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L5 | 2 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 252.3 | 161 | 0.95867 | 251.3413 | 160.0413 | 1.7349 | 0.90544 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L5 | 2 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 272 | 173.3 | 0.95867 | 271.0413 | 172.3413 | 1.8756 | 0.96775 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L5 | 2 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 290.3 | 182.5 | 0.95867 | 289.3413 | 181.5413 | 2.0485 | 0.9466 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L5 | 2 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 109.7 | 73.75 | 0.95867 | 108.7413 | 72.7913 | 0.68315 | 0.51777 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication, something clearly wrong with <10 filters
AR31A | 20181021 | L5 | 2 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 234.6 | 149.6 | 0.95867 | 233.6413 | 148.6413 | 1.6152 | 0.83708 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | very poor replication, something clearly wrong with <10 filters and higher than WSW so toss this one
AR31A | 20181021 | L4 | 3 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 39.31 | 29.76 | 0.95867 | 38.3513 | 28.8013 | 0.18148 | 0.29369 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication
AR31A | 20181021 | L4 | 3 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 41.91 | 31.34 | 0.95867 | 40.9513 | 30.3813 | 0.60258 | 0.90113 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication
AR31A | 20181021 | L4 | 3 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 148.1 | 91.61 | 0.95867 | 147.1413 | 90.6513 | 3.2204 | 1.2663 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L4 | 3 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 159 | 99.05 | 0.95867 | 158.0413 | 98.0913 | 3.4176 | 1.4373 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L4 | 3 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 381.4 | 237.8 | 0.95867 | 380.4413 | 236.8413 | 2.7288 | 1.1787 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (21% range) but dilutions prone to more error
AR31A | 20181021 | L4 | 3 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 149.3 | 90.15 | 0.95867 | 148.3413 | 89.1913 | 3.372 | 1.0424 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | somewhat poor replication (21% range) but dilutions prone to more error
AR31A | 20181021 | L4 | 3 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 160.4 | 99.56 | 0.95867 | 159.4413 | 98.6013 | 3.4684 | 1.4118 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR31A | 20181021 | L4 | 3 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 149.9 | 89.91 | 0.95867 | 148.9413 | 88.9513 | 3.4199 | 0.98268 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR31A | 20181021 | L4 | 3 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 72.85 | 50.11 | 0.95867 | 71.8913 | 49.1513 | 0.43212 | 0.37878 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication, something clearly wrong with one of the <10 filters; could this one be accurate? Similar to C4N14 10a value
AR31A | 20181021 | L4 | 3 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 361.6 | 226.8 | 0.95867 | 360.6413 | 225.8413 | 2.5616 | 1.1644 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | very poor replication, something clearly wrong with this <10 filter; toss
AR31A | 20181021 | L3 | 4 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 169.9 | 113.5 | 0.95867 | 168.9413 | 112.5413 | 1.0718 | 0.78497 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L3 | 4 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 145.6 | 98.58 | 0.95867 | 144.6413 | 97.6213 | 0.89351 | 0.71706 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L3 | 4 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 235.4 | 150.4 | 0.95867 | 234.4413 | 149.4413 | 4.8457 | 2.5508 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L3 | 4 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 237.7 | 152.5 | 0.95867 | 236.7413 | 151.5413 | 4.8571 | 2.6434 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L3 | 4 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 160.4 | 100.6 | 0.95867 | 159.4413 | 99.6413 | 3.4091 | 1.5226 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L3 | 4 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 156.7 | 100 | 0.95867 | 155.7413 | 99.0413 | 3.2324 | 1.6696 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L3 | 4 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 152.3 | 99.56 | 0.95867 | 151.3413 | 98.6013 | 3.0066 | 1.8736 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L3 | 4 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 154.2 | 97.79 | 0.95867 | 153.2413 | 96.8313 | 3.2158 | 1.5768 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L3 | 4 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 67.77 | 45.12 | 0.95867 | 66.8113 | 44.1613 | 0.43041 | 0.29817 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication, something clearly wrong with <10 filters
AR31A | 20181021 | L3 | 4 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 386.4 | 240.7 | 0.95867 | 385.4413 | 239.7413 | 2.7687 | 1.1866 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | very poor replication, something clearly wrong with <10 filters and higher than WSW so toss this one
AR31A | 20181021 | L2 | 5 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 245.5 | 162.3 | 0.95867 | 244.5413 | 161.3413 | 1.581 | 1.0808 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L2 | 5 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 265.7 | 171.2 | 0.95867 | 264.7413 | 170.2413 | 1.7958 | 1.0129 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L2 | 5 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 255.5 | 156.4 | 0.79933 | 254.7007 | 155.6007 | 5.6495 | 2.0519 | 20181101 | 9-54 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR31A | 20181021 | L2 | 5 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 257.7 | 158.5 | 0.95867 | 256.7413 | 157.5413 | 5.6552 | 2.1422 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR31A | 20181021 | L2 | 5 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 315.4 | 188.5 | 0.95867 | 314.4413 | 187.5413 | 7.2343 | 2.0479 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L2 | 5 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 296.8 | 180.8 | 0.95867 | 295.8413 | 179.8413 | 6.6129 | 2.2882 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L2 | 5 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 296.2 | 183.9 | 0.95867 | 295.2413 | 182.9413 | 6.402 | 2.6526 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L2 | 5 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 315.4 | 192.8 | 0.95867 | 314.4413 | 191.8413 | 6.9892 | 2.5059 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L2 | 5 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 201.3 | 130.4 | 0.95867 | 200.3413 | 129.4413 | 1.3473 | 0.78825 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | pvery poor replication, something clearly wrong with <10 filters on this night of sampling
AR31A | 20181021 | L2 | 5 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 271 | 165.7 | 0.95867 | 270.0413 | 164.7413 | 6.003 | 2.1508 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | very poor replication and says it was diluted in labNB. Was it really? Would be much more reasonbale it wasn't, though maybe it wasn't pre-filtered? Looks like WSW
AR31A | 20181021 | L1 | 6 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 202 | 125.3 | 0.95867 | 201.0413 | 124.3413 | 4.3725 | 1.7817 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor phaeo replication (34% range)
AR31A | 20181021 | L1 | 6 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 212.9 | 138.9 | 0.95867 | 211.9413 | 137.9413 | 4.2186 | 2.6087 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor phaeo replication (34% range)
AR31A | 20181021 | L1 | 6 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 220.4 | 140.2 | 0.95867 | 219.4413 | 139.2413 | 4.572 | 2.3196 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L1 | 6 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 210.1 | 134.9 | 0.95867 | 209.1413 | 133.9413 | 4.287 | 2.3423 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR31A | 20181021 | L1 | 6 | 3 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 65.68 | 44.21 | 0.95867 | 64.7213 | 43.2513 | 0.40799 | 0.30558 | 20181101 | 9-57 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | why do these <10 look ok when previous 4 casts are terrible?
AR31A | 20181021 | L1 | 6 | 3 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 79.8 | 53.7 | 0.95867 | 78.8413 | 52.7413 | 0.49597 | 0.37416 | 20181101 | 9-56 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | why do these <10 look ok when previous 4 casts are terrible?
EN627 | 20190201 | L1 | 3 | 1 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 183.9 | 111.5 | 0.628 | 183.272 | 110.872 | 2.7516 | 0.90677 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | this entire cast, tissue capsules were pre-labeled as C1 but actually ended up being C3 so corrected in this document
EN627 | 20190201 | L1 | 3 | 1 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 179.9 | 110.5 | 0.628 | 179.272 | 109.872 | 2.6376 | 0.98779 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | entered as 1-1b, entry in lab notebook kind of looks like 1-11a but theres no niskin 11
EN627 | 20190201 | L1 | 3 | 3 |  |  | a | 355 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 172.4 | 108.8 | 0.543 | 171.857 | 108.257 | 0.98728 | 0.47173 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 2
EN627 | 20190201 | L1 | 3 | 3 |  |  | b | 355 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 148 | 94.82 | 0.543 | 147.457 | 94.277 | 0.82553 | 0.44507 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 5
EN627 | 20190201 | L1 | 3 | 3 |  |  | a | 355 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 4.713 | 3.095 | 0.543 | 4.17 | 2.552 | 0.025117 | 0.0092775 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 10, qc perhaps all seawater dumped on one filter? DF thinks this is unlikely since even if the larger value of this rep was divided in 2, the value would still be larger than 5 replicates. Toss these reps | Taylor previously QC'ed
EN627 | 20190201 | L1 | 3 | 3 |  |  | b | 355 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 357.9 | 204.2 | 0.543 | 357.357 | 203.657 | 2.3859 | 0.35882 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 8, qc perhaps all seawater dumped on to one filter? ; DF thinks this is unlikely since even if the larger value of this rep was divided in 2, the value would still be larger than 5 replicates. Toss these reps | Taylor previously QC'ed
EN627 | 20190201 | L1 | 3 | 3 |  |  | a | 285 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 185 | 113.5 | 0.628 | 184.372 | 112.872 | 2.7651 | 1.0246 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190201 | L1 | 3 | 3 |  |  | b | 285 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 193.5 | 117.4 | 0.54233 | 192.9577 | 116.8577 | 2.943 | 0.98056 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190201 | L1 | 3 | 3 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 73.58 | 52.45 | 0.628 | 72.952 | 51.822 | 0.40857 | 0.46139 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190201 | L1 | 3 | 3 |  |  | b | 286 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 59.09 | 41.42 | 0.54233 | 58.5477 | 40.8777 | 0.34047 | 0.34336 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190201 | L1 | 3 | 5 |  |  | a | 355 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 279.8 | 165.8 | 0.543 | 279.257 | 165.257 | 1.7697 | 0.45757 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 3
EN627 | 20190201 | L1 | 3 | 5 |  |  | b | 355 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 240.1 | 146.6 | 0.543 | 239.557 | 146.057 | 1.4514 | 0.51703 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 6
EN627 | 20190201 | L1 | 3 | 5 |  |  | a | 355 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 173.8 | 102.2 | 0.543 | 173.257 | 101.657 | 1.1115 | 0.2586 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 9 | poor chl replication (50% range); Taylor previously QC'ed; QF was 2 but DF changed to 1 because within variability of size fractionated filters
EN627 | 20190201 | L1 | 3 | 5 |  |  | b | 355 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 232.7 | 132 | 0.543 | 232.157 | 131.457 | 1.5632 | 0.20849 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 12 | poor chl replication (50% range); Taylor previously QC'ed; QF was 2 but DF changed to 1 because within variability of size fractionated filters
EN627 | 20190201 | L1 | 3 | 5 |  |  | a | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 356.7 | 218.7 | 0.628 | 356.072 | 218.072 | 2.6498 | 0.98561 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | entered as 1-5a but recorded as 1-5b2 in lab notebook
EN627 | 20190201 | L1 | 3 | 5 |  |  | b | 287 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 178.3 | 109 | 0.628 | 177.672 | 108.372 | 2.6613 | 0.95195 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190201 | L1 | 3 | 5 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 70.26 | 49.11 | 0.628 | 69.632 | 48.482 | 0.40896 | 0.40494 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication
EN627 | 20190201 | L1 | 3 | 5 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 43.21 | 30.55 | 0.628 | 42.582 | 29.922 | 0.24479 | 0.25752 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | entered as 1-5<10b, recored as 1-5<10a2 in lab notebook, poor replication
EN627 | 20190201 | L1 | 3 | 7 |  |  | a | 355 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 165.6 | 104.4 | 0.543 | 165.057 | 103.857 | 0.95003 | 0.44969 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR 1 | >5 & >20 size fraction make up compared to WSW and size frac at other depths of this case make all these highly suspect; were small bottles filled from 1 larger bottle? could the bottle have not been mixed before allocating water? good recpliation but size frac makeup too suspect. Looking @SMD 5-200 and 20-200 values, >5 is about 0.4ug/l greater than >20s. SMD <200s same or greater value thanSosik WSW
EN627 | 20190201 | L1 | 3 | 7 |  |  | b | 355 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 151.5 | 93.55 | 0.543 | 150.957 | 93.007 | 0.89958 | 0.35391 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR 4 | >5 & >20 size fraction make up compared to WSW and size frac at other depths of this case make all these highly suspect; were small bottles filled from 1 larger bottle? could the bottle have not been mixed before allocating water? good recpliation but size frac makeup too suspect. Looking @SMD 5-200 and 20-200 values, >5 is about 0.4ug/l greater than >20s. SMD <200s same or greater value thanSosik WSW
EN627 | 20190201 | L1 | 3 | 7 |  |  | a | 355 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 389 | 228.1 | 0.543 | 388.457 | 227.557 | 2.4977 | 0.56916 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR 7 | >5 & >20 size fraction make up compared to WSW and size frac at other depths of this case make all these highly suspect; were small bottles filled from 1 larger bottle? could the bottle have not been mixed before allocating water? good recpliation but size frac makeup too suspect. 5's and 20's filtering coupled together and 5's are suspicious enough to question these values
EN627 | 20190201 | L1 | 3 | 7 |  |  | b | 355 | 20 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 141 | 82.68 | 0.543 | 140.457 | 82.137 | 2.716 | 0.60501 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR 11 | >5 & >20 size fraction make up compared to WSW and size frac at other depths of this case make all these highly suspect; were small bottles filled from 1 larger bottle? could the bottle have not been mixed before allocating water? good recpliation but size frac makeup too suspect
EN627 | 20190201 | L1 | 3 | 7 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 181.9 | 111.4 | 0.628 | 181.272 | 110.772 | 2.6794 | 0.97568 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190201 | L1 | 3 | 7 |  |  | b | 285 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 175.4 | 108.1 | 0.628 | 174.772 | 107.472 | 2.6026 | 1.0058 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190201 | L1 | 3 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 21.61 | 15.7 | 0.628 | 20.982 | 15.072 | 0.11231 | 0.13635 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication
EN627 | 20190201 | L1 | 3 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 58.95 | 42.27 | 0.628 | 58.322 | 41.642 | 0.31696 | 0.37005 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication
EN627 | 20190202 | L2 | 7 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 221.4 | 130.8 | 0.628 | 220.772 | 130.172 | 5.1649 | 1.2779 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L2 | 7 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 317 | 188.8 | 0.54233 | 316.4577 | 188.2577 | 4.8723 | 1.3395 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L2 | 7 | 4 |  |  | a | 355 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 280.7 | 162.5 | 0.543 | 280.157 | 161.957 | 1.8349 | 0.34789 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR 14 | very poor replication (52% range) could volume filtered be incorrect?; Taylor previously QC'ed
EN627 | 20190202 | L2 | 7 | 4 |  |  | b | 355 | 5 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 205 | 117.7 | 0.543 | 204.457 | 117.157 | 4.0656 | 0.67134 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR 17 | very poor replication (52% range) could volume filtered be incorrect? Could labNB be typo and it wasn't diluted?; Taylor previously QC'ed
EN627 | 20190202 | L2 | 7 | 4 |  |  | a | 355 | 20 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 194.7 | 113.5 | 0.543 | 194.157 | 112.957 | 3.7815 | 0.7856 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR 20 | very poor replication; Taylor previously QC'ed
EN627 | 20190202 | L2 | 7 | 4 |  |  | b | 355 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 333.9 | 191.7 | 0.543 | 333.357 | 191.157 | 2.2074 | 0.36887 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR 23 | very poor replication; Taylor previously QC'ed
EN627 | 20190202 | L2 | 7 | 4 |  |  | a | 285 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 149 | 90.81 | 0.628 | 148.372 | 90.182 | 3.3755 | 1.1663 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L2 | 7 | 4 |  |  | b | 286 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 226.4 | 135.2 | 0.628 | 225.772 | 134.572 | 3.5146 | 0.98791 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L2 | 7 | 4 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 80.22 | 54.73 | 0.628 | 79.592 | 54.102 | 0.49288 | 0.41537 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L2 | 7 | 4 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 80.67 | 54.22 | 0.628 | 80.042 | 53.592 | 0.51144 | 0.38824 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L2 | 7 | 7 |  |  | a | 355 | 5 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 140.6 | 84.72 | 0.543 | 140.057 | 84.177 | 2.6023 | 0.80112 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 15 | poor replication 81% chl range (dilution is more suspect) but could volume filtered be incorrect? Or could >5 and >20 be switched?; Taylor previously QC'edl volume is correct, DF thinks bottle just wasn't properly mixed prior to pouring
EN627 | 20190202 | L2 | 7 | 7 |  |  | b | 355 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 320 | 186.6 | 0.543 | 319.457 | 186.057 | 2.0708 | 0.43674 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 18 | poor replication 81% chl range, could vol filt be incorrect? Or maybe dilution is source of error? Or sample bottle not well mixed when pouring vol to filt, or could >5 and >20 be switched?; Taylor previously QC'ed; DF thinks bottle just wasn't properly mixed
EN627 | 20190202 | L2 | 7 | 7 |  |  | a | 355 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 231.3 | 135.9 | 0.543 | 230.757 | 135.357 | 1.4809 | 0.34333 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 21 | poor replication (52% chl range) could >5 and >20 be switched? Could vol filt be inaccurate?; Taylor previously QC'ed; DF thinks bottle just wasn't properly mixed
EN627 | 20190202 | L2 | 7 | 7 |  |  | b | 355 | 20 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 138.6 | 84.09 | 0.543 | 138.057 | 83.547 | 2.5385 | 0.83945 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 24 | poor replication (52% chl range) could >5 and >20 be switched? Could vol filt be inaccurate?; Taylor previously QC'ed; DF thinks bottle just wasn't properly mixed
EN627 | 20190202 | L2 | 7 | 7 |  |  | a | 287 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 156.6 | 93.08 | 0.628 | 155.972 | 92.452 | 3.659 | 0.96468 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L2 | 7 | 7 |  |  | b | 287 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 155.9 | 94.06 | 0.628 | 155.272 | 93.432 | 3.5622 | 1.1105 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L2 | 7 | 7 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 60.04 | 41.39 | 0.54233 | 59.4977 | 40.8477 | 0.36062 | 0.32512 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L2 | 7 | 7 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 65.76 | 44.43 | 0.628 | 65.132 | 43.802 | 0.41244 | 0.32289 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L2 | 7 | 10 |  |  | a | 355 | 5 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 143 | 84.19 | 0.543 | 142.457 | 83.647 | 2.7388 | 0.64324 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 13 | Taylor previously QC'ed
EN627 | 20190202 | L2 | 7 | 10 |  |  | b | 355 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 397.2 | 229.7 | 0.543 | 396.657 | 229.157 | 2.6002 | 0.48827 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 16 | Taylor previously QC'ed
EN627 | 20190202 | L2 | 7 | 10 |  |  | a | 355 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 268.6 | 156.7 | 0.543 | 268.057 | 156.157 | 1.7371 | 0.36752 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 19 | somewhat poor chl replication (33% range); Taylor previously QC'ed; DF thinks this is just within biological variability of size fractions
EN627 | 20190202 | L2 | 7 | 10 |  |  | b | 355 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 187.9 | 108.3 | 0.543 | 187.357 | 107.757 | 1.2357 | 0.21662 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 22 | somewhat poor chl replication (33% range); Taylor previously QC'ed; DF thinks this is just within biological variability of size fractions
EN627 | 20190202 | L2 | 7 | 10 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 240.7 | 141.8 | 0.628 | 240.072 | 141.172 | 3.7587 | 0.89942 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L2 | 7 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 161.5 | 97.03 | 0.628 | 160.872 | 96.402 | 3.6753 | 1.096 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L2 | 7 | 10 |  |  | a | 290 | 10 | 5 | 1 |  | 1 | Cal_110119 | 1.8682 | 0.5122 | 49.99 | 33.9 | 0.628 | 49.362 | 33.272 | 0.30575 | 0.24317 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | cruise logsheet "surf <10a problem funnel" so toss; process of elimination assume tissue capsule was mislabeled with C7N4<10a, this sample was missing and there were 2 of C7N4<10a and the other one matched the other 10b | flag 3 log sheet says problem with <10a nitex funnel, probably filt slow
EN627 | 20190202 | L2 | 7 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 76.02 | 52.43 | 0.628 | 75.392 | 51.802 | 0.44827 | 0.40636 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | replicate was bad but had logsheet note of problem with <10 filter so leave this as flag 1, this value similar to other <10's in this cast
EN627 | 20190202 | L3 | 8 | 1 |  |  | a | 400 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 189.4 | 107.3 | 0.543 | 188.857 | 106.757 | 1.1311 | 0.14585 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 27 | somehwat poor replication (25% range) could >5's and >20's tubes been mixed up? "sampled 200ml x2 from amber bottles (no small bottle transfer)"; Taylor previously QC'ed; DF thinks this is a 20 rep
EN627 | 20190202 | L3 | 8 | 1 |  |  | b | 400 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 237.6 | 131.3 | 0.543 | 237.057 | 130.757 | 1.4645 | 0.099514 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 30 | somehwat poor replication (25% range) could >5's and >20's tubes been mixed up? "sampled 200ml x2 from amber bottles (no small bottle transfer)"; Taylor previously QC'ed; DF thinks this is a 20 rep
EN627 | 20190202 | L3 | 8 | 1 |  |  | a | 400 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 259.1 | 150.6 | 0.543 | 258.557 | 150.057 | 1.4948 | 0.30005 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 33 | very poor replication (61% range) poor dilution? Inccorect vol filt recorded? "sampled 200ml x2 from amber bottles (no small bottle transfer)"; Taylor previously QC'ed; DF thinks this is a 5 rep
EN627 | 20190202 | L3 | 8 | 1 |  |  | b | 400 | 20 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 162.3 | 94.38 | 0.543 | 161.757 | 93.837 | 2.8072 | 0.56 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 36 | very poor replication (61% range) poor dilution? Inccorect vol filt recorded? This sample almost same as WSW and greater than >5's so toss this one; "sampled 200ml x2 from amber bottles (no small bottle transfer)"; Taylor previously QC'ed
EN627 | 20190202 | L3 | 8 | 1 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 388.7 | 234.3 | 0.628 | 388.072 | 233.672 | 2.9855 | 0.9373 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L3 | 8 | 1 |  |  | b | 285 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 136.9 | 82.03 | 0.54233 | 136.3577 | 81.4877 | 3.1829 | 0.92103 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L3 | 8 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 375.1 | 226.7 | 0.628 | 374.472 | 226.072 | 2.82 | 0.90977 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L3 | 8 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 344.2 | 209.8 | 0.54233 | 343.6577 | 209.2577 | 2.554 | 0.8984 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L3 | 8 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 354.1 | 214.3 | 0.628 | 353.472 | 213.672 | 2.6566 | 0.86861 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L3 | 8 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 323.8 | 195.7 | 0.54233 | 323.2577 | 195.1577 | 2.4342 | 0.78549 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L3 | 8 | 6 |  |  | a | 400 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 97.2 | 54.56 | 0.543 | 96.657 | 54.017 | 0.58745 | 0.058656 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 25 | poor replication (30% range), but DF thinks within biological variation-- concern is that >5's are lower than >20s….. Could these have been completely switched?; Taylor previously QC'ed, DF thinks these were likely switched; also no small bottle transfer, sampled 200mL x2 from larger amber bottle
EN627 | 20190202 | L3 | 8 | 6 |  |  | b | 400 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 131.3 | 73.33 | 0.543 | 130.757 | 72.787 | 0.79865 | 0.071966 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 29 | poor replication (30% range) but DF thinks within biological variation-- concern is that >5's are lower than >20s ….. Could these have been completely switched?; Taylor previously QC'ed, DF thinks these were likely switched; also no small bottle transfer, sampled 200mL x2 from larger amber bottle
EN627 | 20190202 | L3 | 8 | 6 |  |  | a | 400 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 286.3 | 164 | 0.543 | 285.757 | 163.457 | 1.6849 | 0.27021 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 32 | poor replication (25% range) but DF thinks within biological variation-- concern is that >5's are lower than >20s which doesn't make sense, could they have been switched?; Taylor previously QC'ed, DF thinks these were likely switched; also no small bottle transfer, sampled 200mL x2 from larger amber bottle
EN627 | 20190202 | L3 | 8 | 6 |  |  | b | 400 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 364.9 | 206.1 | 0.543 | 364.357 | 205.557 | 2.1878 | 0.27092 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 35 | poor replication (25% range) but DF thinks within biological variation-- concern is that >5's are lower than >20s which doesn't make sense, could they have been switched?; Taylor previously QC'ed, DF thinks these were likely switched; also no small bottle transfer, sampled 200mL x2 from larger amber bottle
EN627 | 20190202 | L3 | 8 | 6 |  |  | a | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 344.6 | 208.4 | 0.628 | 343.972 | 207.772 | 2.6152 | 0.84846 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L3 | 8 | 6 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 324 | 197.5 | 0.628 | 323.372 | 196.872 | 2.446 | 0.85899 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L3 | 8 | 8 |  |  | a | 400 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 223.5 | 126.4 | 0.543 | 222.957 | 125.857 | 1.3377 | 0.16765 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 26 | very poor replication (55% range) could vol filt be wrong? Could this be >20 and accidentally switched?; Taylor previously QC'ed, DF thinks this is likely a >20 rep but since replication is so poor with other >20 rep, toss. QF was a 2 but DF changed to 3
EN627 | 20190202 | L3 | 8 | 8 |  |  | b | 400 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 389.1 | 217.5 | 0.543 | 388.557 | 216.957 | 2.3641 | 0.23093 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 28 | very poor replication (55% range) could vol filt be wrong? Or the other replication is not the replicate ; Taylor previously QC'ed
EN627 | 20190202 | L3 | 8 | 8 |  |  | a | 400 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 348.6 | 200.8 | 0.543 | 348.057 | 200.257 | 2.0362 | 0.35907 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 31 | poor replication (23%); Taylor previously QC'ed
EN627 | 20190202 | L3 | 8 | 8 |  |  | b | 400 | 20 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 149.3 | 86.87 | 0.543 | 148.757 | 86.327 | 2.5803 | 0.51743 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 34 | poor replication (23% range) and value higher than the >5's could this be the >5a and the other is the >20b? Dilution could be source of error?; Taylor previously QC'ed, DF thinks this is the other >5 rep but hard to say because replication is so poor, so toss
EN627 | 20190202 | L3 | 8 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 391.4 | 235.8 | 0.628 | 390.772 | 235.172 | 2.9568 | 0.92308 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L3 | 8 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 383.5 | 229.8 | 0.628 | 382.872 | 229.172 | 2.9207 | 0.8602 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L3 | 8 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 58.97 | 41.68 | 0.628 | 58.342 | 41.052 | 0.32856 | 0.34873 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replicaiton (53% range) and weird phaeo is nearly 1:1 at surf but similar at L2
EN627 | 20190202 | L3 | 8 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 36.7 | 26.74 | 0.628 | 36.072 | 26.112 | 0.18927 | 0.24153 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replicaiton (53% range) could this be the "problem funnel" from the previous cast that caused a lower value? and weird phaeo is nearly 1:1 at surf but similar at L2
EN627 | 20190202 | L4 | 12 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 283.6 | 171.9 | 0.628 | 282.972 | 171.272 | 2.1226 | 0.70307 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L4 | 12 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 150.9 | 93.37 | 0.628 | 150.272 | 92.742 | 2.1864 | 0.87369 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L4 | 12 | 4 |  |  | a | 356 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 317.7 | 182.4 | 0.543 | 317.157 | 181.857 | 2.0944 | 0.34966 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 41 | very poor phaeo replication (75% range) but chl match?; Taylor previously QC'ed
EN627 | 20190202 | L4 | 12 | 4 |  |  | b | 356 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 295.8 | 162.6 | 0.543 | 295.257 | 162.057 | 2.0619 | 0.11607 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 42 | very poor phaeo replication (75% range) but chl match? This phaeo much lower relative to other surrounding values; Taylor previously QC'ed
EN627 | 20190202 | L4 | 12 | 4 |  |  | a | 358 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 307.7 | 178 | 0.543 | 307.157 | 177.457 | 1.9965 | 0.37511 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 47
EN627 | 20190202 | L4 | 12 | 4 |  |  | b | 358 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 285.7 | 167.4 | 0.543 | 285.157 | 166.857 | 1.821 | 0.40893 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 48
EN627 | 20190202 | L4 | 12 | 4 |  |  | a | 287 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 161.8 | 97.06 | 0.628 | 161.172 | 96.432 | 2.4862 | 0.72897 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L4 | 12 | 4 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 271.6 | 164.7 | 0.628 | 270.972 | 164.072 | 2.067 | 0.68735 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L4 | 12 | 4 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 48.16 | 33.39 | 0.628 | 47.532 | 32.762 | 0.28559 | 0.2644 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L4 | 12 | 4 |  |  | b | 287 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 53.2 | 37.32 | 0.628 | 52.572 | 36.692 | 0.30492 | 0.30676 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L4 | 12 | 6 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 282.4 | 164.2 | 0.543 | 281.857 | 163.657 | 1.84 | 0.37185 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 39 | value is larger than WSW is this rep 20b?; DF thinks all chl signal is just all >5 and this is why the 5s/20s are more than the WSW -- just biological variation between samples.
EN627 | 20190202 | L4 | 12 | 6 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 288.4 | 163.1 | 0.543 | 287.857 | 162.557 | 1.9506 | 0.24646 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 40 | value is larger than WSW; DF thinks all chl signal is just all >5 and this is why the 5s/20s are more than the WSW -- just biological variation between samples.
EN627 | 20190202 | L4 | 12 | 6 |  |  | a | 356 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 300.1 | 177.1 | 0.543 | 299.557 | 176.557 | 1.904 | 0.46883 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 45 | value is larger than WSW; DF thinks all chl signal is just all >5 and this is why the 5s/20s are more than the WSW -- just biological variation between samples.
EN627 | 20190202 | L4 | 12 | 6 |  |  | b | 356 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 296.4 | 169.3 | 0.543 | 295.857 | 168.757 | 1.9675 | 0.30054 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 46 | value is larger than WSW could this be rep 5a?; DF thinks all chl signal is just all >5 and this is why the 5s/20s are more than the WSW -- just biological variation between samples.
EN627 | 20190202 | L4 | 12 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 216.5 | 131.9 | 0.628 | 215.872 | 131.272 | 1.6076 | 0.55812 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L4 | 12 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 194.4 | 119.8 | 0.628 | 193.772 | 119.172 | 1.4176 | 0.54852 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L4 | 12 | 8 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 258.7 | 157.6 | 0.628 | 258.072 | 156.972 | 1.9549 | 0.6803 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L4 | 12 | 8 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 296 | 179.6 | 0.628 | 295.372 | 178.972 | 2.2507 | 0.75379 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L4 | 12 | 8 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 227.8 | 137.8 | 0.54233 | 227.2577 | 137.2577 | 1.7402 | 0.56398 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | major problems with <10um nitex filters on this cruise, nitex was missing from filter funnel for a while; this value close to WSW value so toss and only use rep b
EN627 | 20190202 | L4 | 12 | 8 |  |  | b | 286 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 61.84 | 42.28 | 0.628 | 61.212 | 41.652 | 0.37689 | 0.3199 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | QC 10a & 10b very different but this value is consistent with size fraction pattern of cruise
EN627 | 20190202 | L4 | 12 | 10 |  |  | a | 356 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 304.1 | 171 | 0.543 | 303.557 | 170.457 | 2.0604 | 0.2305 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 37 | very poor phaeo replication (102% range); Taylor previously QC'ed
EN627 | 20190202 | L4 | 12 | 10 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 342.7 | 183.9 | 0.543 | 342.157 | 183.357 | 2.4721 | 0.0060797 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 38 | very poor phaeo replication (102% range) this phaeo seems too low compared to others, chl also happens to be > WSW; Taylor previously QC'ed
EN627 | 20190202 | L4 | 12 | 10 |  |  | a | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 317 | 184.3 | 0.543 | 316.457 | 183.757 | 2.0658 | 0.41779 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 43 | Taylor previously QC'ed
EN627 | 20190202 | L4 | 12 | 10 |  |  | b | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 306.9 | 178 | 0.543 | 306.357 | 177.457 | 2.0066 | 0.3918 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 44 | Taylor previously QC'ed
EN627 | 20190202 | L4 | 12 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 279.2 | 167.6 | 0.54233 | 278.6577 | 167.0577 | 2.1207 | 0.63544 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L4 | 12 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 323.8 | 195.9 | 0.54233 | 323.2577 | 195.3577 | 2.4304 | 0.79259 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L4 | 12 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 57.34 | 39.5 | 0.628 | 56.712 | 38.872 | 0.33901 | 0.30231 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | QC 10a & 10b very different but this value is consistent with size fraction pattern of cruise
EN627 | 20190202 | L4 | 12 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 266.8 | 162.5 | 0.54233 | 266.2577 | 161.9577 | 1.982 | 0.69002 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | major problems with <10um nitex filters on this cruise, nitex was missing from filter funnel for a while; this value close to WSW value so toss and only use rep a
EN627 | 20190202 | L5 | 13 | 2 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 92.84 | 60.81 | 0.628 | 92.212 | 60.182 | 1.2173 | 0.76847 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L5 | 13 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 99.54 | 64.63 | 0.628 | 98.912 | 64.002 | 1.3268 | 0.78506 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L5 | 13 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 115.2 | 71.8 | 0.628 | 114.572 | 71.172 | 0.82472 | 0.34949 | 20181101 | 9-78 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L5 | 13 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 125 | 77.53 | 0.54233 | 124.4577 | 76.9877 | 0.90206 | 0.3681 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L5 | 13 | 6 |  |  | a | 400 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 111.3 | 66.96 | 0.543 | 110.757 | 66.417 | 0.61087 | 0.18355 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 53
EN627 | 20190202 | L5 | 13 | 6 |  |  | b | 400 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 110.8 | 66.01 | 0.543 | 110.257 | 65.467 | 0.61707 | 0.16599 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 54
EN627 | 20190202 | L5 | 13 | 6 |  |  | a | 400 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 83.95 | 48.81 | 0.543 | 83.407 | 48.267 | 0.48412 | 0.093207 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 59
EN627 | 20190202 | L5 | 13 | 6 |  |  | b | 400 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 90.3 | 52.91 | 0.543 | 89.757 | 52.367 | 0.51512 | 0.11125 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 60
EN627 | 20190202 | L5 | 13 | 6 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 118.8 | 73.79 | 0.54233 | 118.2577 | 73.2477 | 0.87032 | 0.35933 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | water sample volume is guessed, not written on log sheet
EN627 | 20190202 | L5 | 13 | 6 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 119.6 | 73.85 | 0.54233 | 119.0577 | 73.3077 | 0.88463 | 0.34603 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L5 | 13 | 7 |  |  | a | 400 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 106.3 | 64.3 | 0.543 | 105.757 | 63.757 | 0.57863 | 0.18398 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 51
EN627 | 20190202 | L5 | 13 | 7 |  |  | b | 400 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 111.9 | 66.76 | 0.543 | 111.357 | 66.217 | 0.62189 | 0.17014 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 52
EN627 | 20190202 | L5 | 13 | 7 |  |  | a | 400 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 85.71 | 50.57 | 0.543 | 85.167 | 50.027 | 0.48412 | 0.11426 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 57
EN627 | 20190202 | L5 | 13 | 7 |  |  | b | 400 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 79.32 | 46.78 | 0.543 | 78.777 | 46.237 | 0.4483 | 0.10475 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 58
EN627 | 20190202 | L5 | 13 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 122.1 | 76 | 0.54233 | 121.5577 | 75.4577 | 0.87602 | 0.36889 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L5 | 13 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 120.6 | 75.2 | 0.628 | 119.972 | 74.572 | 0.86272 | 0.36758 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190202 | L5 | 13 | 10 |  |  | a | 400 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 104.7 | 63.33 | 0.543 | 104.157 | 62.787 | 0.56995 | 0.18105 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 49
EN627 | 20190202 | L5 | 13 | 10 |  |  | b | 400 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 103.7 | 62 | 0.543 | 103.157 | 61.457 | 0.5745 | 0.1606 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 50
EN627 | 20190202 | L5 | 13 | 10 |  |  | a | 400 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 77.36 | 45.55 | 0.543 | 76.817 | 45.007 | 0.43824 | 0.10009 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 55
EN627 | 20190202 | L5 | 13 | 10 |  |  | b | 400 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 75.06 | 44.38 | 0.543 | 74.517 | 43.837 | 0.42268 | 0.10166 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 56
EN627 | 20190202 | L5 | 13 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 142.6 | 87.56 | 0.54233 | 142.0577 | 87.0177 | 1.0459 | 0.38972 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | somewhat poor replication (26% range)
EN627 | 20190202 | L5 | 13 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 115.5 | 73.14 | 0.628 | 114.872 | 72.512 | 0.80495 | 0.39136 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | somewhat poor replication (26% range)
EN627 | 20190202 | L5 | 13 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 132 | 80.52 | 0.54233 | 131.4577 | 79.9777 | 0.97826 | 0.34122 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | major problems with <10um on this cruise, sometimes funnel missing nitex, this matches WSW
EN627 | 20190202 | L5 | 13 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 53.7 | 35.39 | 0.54233 | 53.1577 | 34.8477 | 0.34794 | 0.22698 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190203 | L6 | 16 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 155.7 | 100.2 | 0.54233 | 155.1577 | 99.6577 | 1.0546 | 0.58952 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 154.9 | 98.69 | 0.54233 | 154.3577 | 98.1477 | 1.0681 | 0.55111 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 42.07 | 27.93 | 0.54233 | 41.5277 | 27.3877 | 0.2687 | 0.18315 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 42.86 | 28.47 | 0.54233 | 42.3177 | 27.9277 | 0.27345 | 0.18731 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 120.3 | 72.68 | 0.54233 | 119.7577 | 72.1377 | 0.90491 | 0.28523 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 142.7 | 85.5 | 0.54233 | 142.1577 | 84.9577 | 1.087 | 0.31469 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 6 |  |  | a | 360 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 180 | 105.5 | 0.628 | 179.372 | 104.872 | 1.1404 | 0.25334 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 63
EN627 | 20190203 | L6 | 16 | 6 |  |  | b | 360 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 191.7 | 113 | 0.628 | 191.072 | 112.372 | 1.2047 | 0.28872 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 64
EN627 | 20190203 | L6 | 16 | 6 |  |  | a | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 117.2 | 68.55 | 0.628 | 116.572 | 67.922 | 0.75734 | 0.16065 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 69
EN627 | 20190203 | L6 | 16 | 6 |  |  | b | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 106.9 | 62.08 | 0.628 | 106.272 | 61.452 | 0.69772 | 0.13283 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 70
EN627 | 20190203 | L6 | 16 | 6 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 195.8 | 117.5 | 0.54233 | 195.2577 | 116.9577 | 1.514 | 0.44942 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 6 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 188.3 | 114.1 | 0.54233 | 187.7577 | 113.5577 | 1.4347 | 0.47162 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 6 |  |  | a | 286 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 42.6 | 27.65 | 0.54233 | 42.0577 | 27.1077 | 0.28806 | 0.16542 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 6 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 50.54 | 32.35 | 0.54233 | 49.9977 | 31.8077 | 0.35172 | 0.18225 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 9 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 208.8 | 121.5 | 0.628 | 208.172 | 120.872 | 1.359 | 0.27462 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 65
EN627 | 20190203 | L6 | 16 | 9 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 223.1 | 131.4 | 0.628 | 222.472 | 130.772 | 1.4275 | 0.33993 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 66
EN627 | 20190203 | L6 | 16 | 9 |  |  | a | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 194.8 | 113.5 | 0.628 | 194.172 | 112.872 | 1.2656 | 0.2599 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 71
EN627 | 20190203 | L6 | 16 | 9 |  |  | b | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 189 | 108.9 | 0.628 | 188.372 | 108.272 | 1.2469 | 0.21641 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 72
EN627 | 20190203 | L6 | 16 | 9 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 228 | 131.7 | 0.54233 | 227.4577 | 131.1577 | 1.8621 | 0.33976 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 9 |  |  | b | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 232.8 | 138.6 | 0.54233 | 232.2577 | 138.0577 | 1.8088 | 0.49274 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 9 |  |  | a | 287 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 49.79 | 31.9 | 0.54233 | 49.2477 | 31.3577 | 0.34351 | 0.17924 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | QC 10a & 10b very different but this value is consistent with size fraction pattern of cruise
EN627 | 20190203 | L6 | 16 | 9 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 208.2 | 124.2 | 0.54233 | 207.6577 | 123.6577 | 1.6242 | 0.45168 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | major problems with <10um nitex filters for part of this cruise nitex wasn't in filter funnel; this value close to WSW value so toss and only use rep a
EN627 | 20190203 | L6 | 16 | 11 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 217.6 | 128 | 0.628 | 216.972 | 127.372 | 1.3948 | 0.32667 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 61 | DF thinks this is >20a; QF was 1, DF changed to 2
EN627 | 20190203 | L6 | 16 | 11 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 239.9 | 140.9 | 0.628 | 239.272 | 140.272 | 1.5412 | 0.35468 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 62
EN627 | 20190203 | L6 | 16 | 11 |  |  | a | 230 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 172.2 | 100.9 | 0.628 | 171.572 | 100.272 | 1.7083 | 0.37752 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 67 | switched vol filt between TAR 67 & 68, "were sample tubes switched?" noted in logsheet, This makes replicates more reasonable; Taylor previously QC'ed, value larger than either 5um rep; DF thinks this is >5a
EN627 | 20190203 | L6 | 16 | 11 |  |  | b | 177 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 98.56 | 58.23 | 0.628 | 97.932 | 57.602 | 1.2556 | 0.30138 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 68 | switched vol filt between TAR 67 & 68, "were sample tubes switched?" noted in logsheet, This makes replicates more reasonable; Taylor previously QC'ed; QF was 2, changed to 1
EN627 | 20190203 | L6 | 16 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 225 | 133 | 0.54233 | 224.4577 | 132.4577 | 1.7482 | 0.43706 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 11 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 133 | 78.62 | 0.54233 | 132.4577 | 78.0777 | 2.0667 | 0.50954 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 37.65 | 25.08 | 0.54233 | 37.1077 | 24.5377 | 0.23886 | 0.16596 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L6 | 16 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 41.18 | 26.94 | 0.54233 | 40.6377 | 26.3977 | 0.2706 | 0.16491 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L7 | 17 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 63.88 | 42.98 | 0.54233 | 63.3377 | 42.4377 | 0.39716 | 0.30299 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L7 | 17 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 62.79 | 40.8 | 0.58083 | 62.2092 | 40.2192 | 0.41787 | 0.24567 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L7 | 17 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 117.5 | 72.79 | 0.54233 | 116.9577 | 72.2477 | 0.84961 | 0.34234 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (28%), visually confirm fewer large cells on this GFF
EN627 | 20190203 | L7 | 17 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 154.5 | 94.93 | 0.54233 | 153.9577 | 94.3877 | 1.132 | 0.42523 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | somewhat poor replication (28% range), but could this be C17N7b?? Would make more sense with CTD fluor trace and match other sample better; visually confirm more large cells on this GFF
EN627 | 20190203 | L7 | 17 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 124.7 | 76.76 | 0.54233 | 124.1577 | 76.2177 | 0.91099 | 0.34646 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | somehwat poor replication (26% range) but could this be C17N5b?? Would match that replicate better and the lower value is more consistent with CTD fluor, visually confirm less large cells
EN627 | 20190203 | L7 | 17 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 158.5 | 95.78 | 0.54233 | 157.9577 | 95.2377 | 1.1918 | 0.3794 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (26% range) visually confirm more large cells on this filter than rep a
EN627 | 20190203 | L7 | 17 | 8 |  |  | a | 360 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 191.6 | 111 | 0.628 | 190.972 | 110.372 | 1.2338 | 0.23306 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 75
EN627 | 20190203 | L7 | 17 | 8 |  |  | b | 360 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 220.8 | 127.6 | 0.628 | 220.172 | 126.972 | 1.4267 | 0.2608 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 76
EN627 | 20190203 | L7 | 17 | 8 |  |  | a | 360 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 264.8 | 152.2 | 0.628 | 264.172 | 151.572 | 1.7236 | 0.29077 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 0 | TAR 81. >20 shouldn't be larger than >5 values
EN627 | 20190203 | L7 | 17 | 8 |  |  | b |  | 20 |  |  | 0 | [7 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 1 |  |  |  | 9 | 0 | TAR 82 no sample out of water
EN627 | 20190203 | L7 | 17 | 8 |  |  | a | 285 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 146.9 | 86.86 | 0.54233 | 146.3577 | 86.3177 | 2.3219 | 0.57625 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L7 | 17 | 8 |  |  | b | 286 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 151.7 | 90.48 | 0.54233 | 151.1577 | 89.9377 | 2.3592 | 0.64988 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L7 | 17 | 9 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 174.1 | 101.5 | 0.628 | 173.472 | 100.872 | 1.1302 | 0.23315 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 77 | poor replication (55% range) could this be swapped with the >20b?; Taylor previously QC'ed; DF thinks tthis is >20b rep
EN627 | 20190203 | L7 | 17 | 9 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 303.8 | 175.9 | 0.628 | 303.172 | 175.272 | 1.991 | 0.37783 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 78 | poor replication (55% range) incorrect vol filt? Poor mixing of larger sample bottle before pooring? Switched replicates of 5's and 20's?; Taylor previously QC'ed
EN627 | 20190203 | L7 | 17 | 9 |  |  | a | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 165.6 | 96.08 | 0.628 | 164.972 | 95.452 | 1.0822 | 0.20784 | 20181101 | 9-81 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 83 | poor replication (33% range) incorrect vol filt? Poor mixing of larger sample bottle before pooring? Switched replicates of 5's and 20's?; Taylor previously QC'ed
EN627 | 20190203 | L7 | 17 | 9 |  |  | b | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 226.3 | 129.2 | 0.628 | 225.672 | 128.572 | 1.5116 | 0.22613 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 84 | poor replication (33%) could this be the >5a replicate?; Taylor previously QC'ed; DF thinks this is >5a sample
EN627 | 20190203 | L7 | 17 | 9 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 310.9 | 186.1 | 0.54233 | 310.3577 | 185.5577 | 2.4131 | 0.70192 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor chl replication (32% range) could C17N9 and N10 replicates be switched? Visually confirm substantially more large cells on this filter | Taylor previously QC'ed
EN627 | 20190203 | L7 | 17 | 9 |  |  | b | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 225.9 | 135.3 | 0.54233 | 225.3577 | 134.7577 | 1.7396 | 0.50685 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor chl replication (32% range) could C17N9 and N10 replicates be switched? Would be better match if this was surf N10b visually confirm large cells on GFF | Taylor previously QC'ed
EN627 | 20190203 | L7 | 17 | 10 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 265.1 | 154.5 | 0.628 | 264.472 | 153.872 | 1.7217 | 0.35792 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 73 | very poor replication (45% range) incorrect vol filt? Uneven mixing of larger sample bottle before pouring? ; Taylor previously QC'ed
EN627 | 20190203 | L7 | 17 | 10 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 164.6 | 94.87 | 0.628 | 163.972 | 94.242 | 1.0855 | 0.18822 | 20181101 | 9-81 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 74 | very poor replication (45% range) incorrect vol filt? Could this be N10 >20b? Uneven mixing of larger sample bottle before pouring? ; Taylor previously QC'ed; DF thinks this is >20b rep
EN627 | 20190203 | L7 | 17 | 10 |  |  | a | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 200 | 115.3 | 0.628 | 199.372 | 114.672 | 1.3185 | 0.2313 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 79 | poor replication (27% range) vol filt got mixed up? Uneven mixing of bottle?; Taylor previously QC'ed
EN627 | 20190203 | L7 | 17 | 10 |  |  | b | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 265.1 | 153.7 | 0.628 | 264.472 | 153.072 | 1.7342 | 0.33465 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 80 | could this be N>5b? Poor replication, or wrong vol filt? ; Taylor previously QC'ed; DF thinks this is >5b rep
EN627 | 20190203 | L7 | 17 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 245 | 145.4 | 0.58083 | 244.4192 | 144.8192 | 1.8927 | 0.49658 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | somewhat poor replication (27% range) could replicates of N9 and N10 been switched? Visually confirm large cells | Taylor previously QC'ed
EN627 | 20190203 | L7 | 17 | 10 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 162.2 | 96.27 | 0.54233 | 161.6577 | 95.7277 | 2.5057 | 0.65296 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | visually confirm this GFF has more large cells on it. could this be C17N9b? Reps would match better and this 20m sample is sitting closer to 40m CM but well mixed surf layer. somewhat poor replication (27% range) could replicates of N9 and N10 been switched? | Taylor previously QC'ed
EN627 | 20190203 | L7 | 17 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 66.43 | 41.88 | 0.54233 | 65.8877 | 41.3377 | 0.46652 | 0.21548 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (39% range)
EN627 | 20190203 | L7 | 17 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 98.29 | 61.6 | 0.54233 | 97.7477 | 61.0577 | 0.69721 | 0.31013 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (39% range)
EN627 | 20190203 | L8 | 19 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 29.88 | 20.48 | 0.58083 | 29.2992 | 19.8992 | 0.17863 | 0.14967 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L8 | 19 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 26.15 | 18.23 | 0.54233 | 25.6077 | 17.6877 | 0.1505 | 0.14131 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L8 | 19 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 38.87 | 25.39 | 0.54233 | 38.3277 | 24.8477 | 0.25616 | 0.15378 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L8 | 19 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 33.08 | 21.43 | 0.58083 | 32.4992 | 20.8492 | 0.22138 | 0.12259 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L8 | 19 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 83.8 | 52.91 | 0.58083 | 83.2192 | 52.3292 | 0.58699 | 0.27634 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | very poor replication (63% range) but visually confirmed from photo of filter uneven large cells captured on the replicates and in dynamic trace of the water column looking at CTD data
EN627 | 20190203 | L8 | 19 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 47.04 | 30.97 | 0.58083 | 46.4592 | 30.3892 | 0.30537 | 0.19599 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | very poor replication (63% range) but visually confirmed from photo of filter uneven large cells captured on the replicates but in dynamic trace of the water column looking at CTD data
EN627 | 20190203 | L8 | 19 | 5 |  |  | a | 360 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 100.1 | 59.88 | 0.628 | 99.472 | 59.252 | 0.61568 | 0.17179 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 89 | somewhat poor chl replication (23% range) but would match very well if replicates were swapped); variation within WSW on this cast makes this sample's variation possible; DF thinks that chl signal for >5 and >20 is just the same, hence variation within 5s and 20s and that they're bascially the same
EN627 | 20190203 | L8 | 19 | 5 |  |  | b | 360 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 81.21 | 49.48 | 0.628 | 80.582 | 48.852 | 0.48571 | 0.16354 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 90 | somewhat poor replication (23% range) could this be the 20a?; variation within WSW on this cast makes this sample's variation possible; DF thinks that chl signal for >5 and >20 is just the same, hence variation within 5s and 20s and that they're bascially the same
EN627 | 20190203 | L8 | 19 | 5 |  |  | a | 360 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 96.36 | 56.03 | 0.628 | 95.732 | 55.402 | 0.61736 | 0.11894 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 95 | somehwat poor replication (21% range) could this be 5b?; variation within WSW on this cast makes this sample's variation possible; DF thinks that chl signal for >5 and >20 is just the same, hence variation within 5s and 20s and that they're bascially the same
EN627 | 20190203 | L8 | 19 | 5 |  |  | b | 360 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 80.45 | 47.97 | 0.628 | 79.822 | 47.342 | 0.49719 | 0.13199 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 96 | somewhat poor chl replication (21% range);variation within WSW on this cast makes this sample's variation possible; DF thinks that chl signal for >5 and >20 is just the same, hence variation within 5s and 20s and that they're bascially the same
EN627 | 20190203 | L8 | 19 | 5 |  |  | a | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 112.5 | 69.67 | 0.58083 | 111.9192 | 69.0892 | 0.82239 | 0.32936 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190203 | L8 | 19 | 5 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 133.7 | 80.87 | 0.58083 | 133.1192 | 80.2892 | 1.0215 | 0.32634 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190203 | L8 | 19 | 5 |  |  | a | 287 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 59.47 | 37.77 | 0.54233 | 58.9277 | 37.2277 | 0.41667 | 0.20394 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication (53% range)
EN627 | 20190203 | L8 | 19 | 5 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 36.53 | 24.08 | 0.54233 | 35.9877 | 23.5377 | 0.24073 | 0.15441 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication (53% range)
EN627 | 20190203 | L8 | 19 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 176.7 | 107.2 | 0.54233 | 176.1577 | 106.6577 | 1.3207 | 0.43897 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L8 | 19 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 169.3 | 101.8 | 0.54233 | 168.7577 | 101.2577 | 1.2827 | 0.38788 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L8 | 19 | 7 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 205.8 | 120.1 | 0.628 | 205.172 | 119.472 | 1.3341 | 0.28061 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR 87 | poor replication (37% rnage) but WSW also range so maybe it's if a big cell is present or not; this sample within size fraction pattern
EN627 | 20190203 | L8 | 19 | 7 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 297.7 | 172.3 | 0.628 | 297.072 | 171.672 | 1.9521 | 0.3681 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR 88 | poor replication (37% rnage) but WSW also range so maybe it's if a big cell is present or not; value greater than WSW so flag as suspicious this replicate variation flagged
EN627 | 20190203 | L8 | 19 | 7 |  |  | a | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 162.8 | 92.38 | 0.628 | 162.172 | 91.752 | 1.0962 | 0.14383 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 93 | poor phaeo replication (49% range); Taylor previously QC'ed
EN627 | 20190203 | L8 | 19 | 7 |  |  | b | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 176.9 | 104.1 | 0.628 | 176.272 | 103.472 | 1.1333 | 0.26518 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 94 | poor phaeo replication (49% range); Taylor previously QC'ed
EN627 | 20190203 | L8 | 19 | 7 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 227.1 | 136.9 | 0.628 | 226.472 | 136.272 | 1.7441 | 0.54356 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (21% range) but visually confirmed from filter photos that this had more larger cells
EN627 | 20190203 | L8 | 19 | 7 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 181.7 | 108.7 | 0.58083 | 181.1192 | 108.1192 | 1.4115 | 0.40353 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (21% range) but visually confirmed from photo of filter that this one is less
EN627 | 20190203 | L8 | 19 | 7 |  |  | a | 286 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 62.95 | 40.14 | 0.54233 | 62.4077 | 39.5977 | 0.43951 | 0.22291 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L8 | 19 | 7 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 62.76 | 39.99 | 0.54233 | 62.2177 | 39.4477 | 0.44028 | 0.22195 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L8 | 19 | 9 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 117.9 | 68.83 | 0.628 | 117.272 | 68.202 | 0.76388 | 0.1579 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 85 | Taylor previously QC'ed
EN627 | 20190203 | L8 | 19 | 9 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 99.52 | 58.67 | 0.628 | 98.892 | 58.042 | 0.63592 | 0.14854 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 86 | Taylor previously QC'ed
EN627 | 20190203 | L8 | 19 | 9 |  |  | a | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 82.97 | 48.53 | 0.628 | 82.342 | 47.902 | 0.53613 | 0.11128 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 91 | poor phaeo replication (48% range); Taylor previously QC'ed
EN627 | 20190203 | L8 | 19 | 9 |  |  | b | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 80.48 | 45.42 | 0.628 | 79.852 | 44.792 | 0.54579 | 0.059598 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 92 | poor phaeo replication (48% range) rep 20b a bit low?; Taylor previously QC'ed
EN627 | 20190203 | L8 | 19 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 148.1 | 89.72 | 0.54233 | 147.5577 | 89.1777 | 1.1094 | 0.36189 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L8 | 19 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 139.7 | 84.84 | 0.54233 | 139.1577 | 84.2977 | 1.0425 | 0.34827 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L8 | 19 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 54.59 | 33.38 | 0.54233 | 54.0477 | 32.8377 | 0.40305 | 0.13871 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190203 | L8 | 19 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 55.18 | 34.51 | 0.54233 | 54.6377 | 33.9677 | 0.39279 | 0.16762 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L9 | 22 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 70.41 | 48.73 | 0.58083 | 69.8292 | 48.1492 | 0.41198 | 0.38239 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L9 | 22 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 64.93 | 44.75 | 0.54233 | 64.3877 | 44.2077 | 0.38347 | 0.34587 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L9 | 22 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 38.36 | 25.11 | 0.58083 | 37.7792 | 24.5292 | 0.25179 | 0.1529 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication (41% range) visually confirm significantly more large cells on this GFF
EN627 | 20190204 | L9 | 22 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 27.81 | 19.11 | 0.58083 | 27.2292 | 18.5292 | 0.16532 | 0.14037 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication (41% range) visually confirm significantly fewer large cells on this filter
EN627 | 20190204 | L9 | 22 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 137 | 83.7 | 0.58083 | 136.4192 | 83.1192 | 1.0128 | 0.35847 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | visually confirm both have large cells, this less general color besides cells
EN627 | 20190204 | L9 | 22 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 168.3 | 102.6 | 0.74183 | 167.5582 | 101.8582 | 1.2485 | 0.43199 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | visually confirm both reps have large cells this one more color tinted from small cells
EN627 | 20190204 | L9 | 22 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 280.1 | 166.5 | 0.74183 | 279.3582 | 165.7582 | 2.1587 | 0.576 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (29% range) but most of cast dealing with patchiness of large cells
EN627 | 20190204 | L9 | 22 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 210.9 | 126.7 | 0.74183 | 210.1582 | 125.9582 | 1.6 | 0.47805 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (29% range) but most of cast dealing with patchiness of large cells
EN627 | 20190204 | L9 | 22 | 6 |  |  | a | 356 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 320.2 | 187.6 | 0.628 | 319.572 | 186.972 | 2.0526 | 0.4602 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR 101 | value larger than either WSW value, QF was 1
EN627 | 20190204 | L9 | 22 | 6 |  |  | b | 356 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 270.5 | 158.3 | 0.628 | 269.872 | 157.672 | 1.7368 | 0.38221 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR 102 | value larger than WSW repa, QF was 1
EN627 | 20190204 | L9 | 22 | 6 |  |  | a | 356 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 307.4 | 177.4 | 0.628 | 306.772 | 176.772 | 2.0124 | 0.36336 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR 107 | value larger than >5 repb, QF was 1
EN627 | 20190204 | L9 | 22 | 6 |  |  | b | 356 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 344.3 | 196.8 | 0.628 | 343.672 | 196.172 | 2.2833 | 0.35319 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR 108 | value larger than both >5 reps and WSW reps, QF was 1
EN627 | 20190204 | L9 | 22 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 221.9 | 135 | 0.58083 | 221.3192 | 134.4192 | 1.6513 | 0.56633 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L9 | 22 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 255.1 | 152.1 | 0.58083 | 254.5192 | 151.5192 | 1.9573 | 0.54251 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L9 | 22 | 7 |  |  | a | 356 | 5 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 147.9 | 86.32 | 0.628 | 147.272 | 85.692 | 2.8597 | 0.59525 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 99 | value larger than WSWand dilution could cause increased variation
EN627 | 20190204 | L9 | 22 | 7 |  |  | b | 356 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 380.9 | 219.7 | 0.628 | 380.272 | 219.072 | 2.4953 | 0.44889 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 100
EN627 | 20190204 | L9 | 22 | 7 |  |  | a | 356 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 294.7 | 173.9 | 0.628 | 294.072 | 173.272 | 1.87 | 0.45874 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 105
EN627 | 20190204 | L9 | 22 | 7 |  |  | b | 356 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 322.2 | 188.5 | 0.628 | 321.572 | 187.872 | 2.0696 | 0.45527 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 106
EN627 | 20190204 | L9 | 22 | 7 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 341.1 | 199.1 | 0.74183 | 340.3582 | 198.3582 | 2.7457 | 0.58423 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication (69% range) high value visually confirm lots of cosconodiscus compared to other N7's |  | pic of filter available
EN627 | 20190204 | L9 | 22 | 7 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 173.2 | 104.8 | 0.58083 | 172.6192 | 104.2192 | 1.3226 | 0.427 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | 2 tissue capsules labeled C22N7b, this sample seems similar to C22N11b; very poor replication (69% range) |  | pic of filter available
EN627 | 20190204 | L9 | 22 | 7 |  |  | a | 286 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 112.6 | 70.42 | 0.54233 | 112.0577 | 69.8777 | 0.81274 | 0.35623 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | 2 tissue capsules labeled C22N7<10a see C22N11<10a for more notes, large cells found on both filters…. Def larger than 10um, toss both samples?
EN627 | 20190204 | L9 | 22 | 7 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 125.8 | 77.65 | 0.54233 | 125.2577 | 77.1077 | 0.93103 | 0.36342 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | 2 tissue capsules labeled C22N7<10b see C22N11<10b for more notes, large cells found on both filters…. Def larger than 10um, toss both samples?
EN627 | 20190204 | L9 | 22 | 9 |  |  | a | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 311 | 185.3 | 0.58083 | 310.4192 | 184.7192 | 2.4136 | 0.66577 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L9 | 22 | 9 |  |  | b | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 286 | 168.6 | 0.74183 | 285.2582 | 167.8582 | 2.2542 | 0.54406 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L9 | 22 | 11 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 165.4 | 97.93 | 0.628 | 164.772 | 97.302 | 1.0503 | 0.26476 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 97 | smaller than either >20 reps; could this be a >20 rep instead? ; DF thinks this is a >20 rep
EN627 | 20190204 | L9 | 22 | 11 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 177.1 | 104.2 | 0.628 | 176.472 | 103.572 | 1.1348 | 0.26497 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 98 | smaller than either >20 reps; could this be a >20 rep instead? Yes, DF thinks this is a >20 rep
EN627 | 20190204 | L9 | 22 | 11 |  |  | a | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 186.7 | 107.8 | 0.628 | 186.072 | 107.172 | 1.2283 | 0.22022 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 103 | larger than either >5 rep, could this be a >5 rep instead?; yes, DF thinks this is a >5 rep
EN627 | 20190204 | L9 | 22 | 11 |  |  | b | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 185.8 | 107.1 | 0.628 | 185.172 | 106.472 | 1.2251 | 0.21388 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 104 | larger than either >5 rep, could this be a >5 rep instead?; yes, DF thinks this is a >5 rep
EN627 | 20190204 | L9 | 22 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 225 | 136.7 | 0.54233 | 224.4577 | 136.1577 | 1.6779 | 0.56841 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 |  |  | pic of filter available
EN627 | 20190204 | L9 | 22 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 214.4 | 131.2 | 0.58083 | 213.8192 | 130.6192 | 1.581 | 0.57395 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | 2 tissue capsules labeled C22N7b, this sample seems similar to C22N11a in value AND visually similar looking at photos of filters |  | pic of filter available
EN627 | 20190204 | L9 | 22 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 91.4 | 58.28 | 0.54233 | 90.8577 | 57.7377 | 0.62937 | 0.32319 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | tissue capsule was C22N7<10a but this sample was missing so comparing with other chl's from this N moved to N11
EN627 | 20190204 | L9 | 22 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 91.95 | 58.41 | 0.54233 | 91.4077 | 57.8677 | 0.63735 | 0.31736 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | tissue capsule was C22N7<10b but this sample was missing so comparing with other chl's from this N moved to N11
EN627 | 20190204 | L10 | 24 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 49.9 | 31.56 | 0.58083 | 49.3192 | 30.9792 | 0.34851 | 0.16259 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L10 | 24 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 57.83 | 35.69 | 0.69017 | 57.1398 | 34.9998 | 0.42072 | 0.15671 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | filled in niskin number with process of elimination
EN627 | 20190204 | L10 | 24 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 47.33 | 28.78 | 0.58083 | 46.7492 | 28.1992 | 0.3525 | 0.11273 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (23% range chl)
EN627 | 20190204 | L10 | 24 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 39.26 | 24.66 | 0.74183 | 38.5182 | 23.9182 | 0.27744 | 0.11717 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (23% range chl)
EN627 | 20190204 | L10 | 24 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 71.87 | 44.25 | 0.74183 | 71.1282 | 43.5082 | 0.52485 | 0.19295 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (39%) but a lot of random large cells throughout water column
EN627 | 20190204 | L10 | 24 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 48.95 | 30.39 | 0.74183 | 48.2082 | 29.6482 | 0.35269 | 0.13645 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication (39%) but a lot of random large cells throughout water column
EN627 | 20190204 | L10 | 24 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 59.27 | 37.08 | 0.58083 | 58.6892 | 36.4992 | 0.42167 | 0.1805 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication (47% range)
EN627 | 20190204 | L10 | 24 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 89.6 | 53.58 | 0.74183 | 88.8582 | 52.8382 | 0.68448 | 0.18725 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication (47% range)
EN627 | 20190204 | L10 | 24 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 151.5 | 91.5 | 0.74183 | 150.7582 | 90.7582 | 1.1402 | 0.35718 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication (42% range) no photo but confirmed many large cells in water can make high rnage of reps
EN627 | 20190204 | L10 | 24 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 102.4 | 63.32 | 0.58083 | 101.8192 | 62.7392 | 0.74263 | 0.29245 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication (42% range) visually confirm this filter only has large cells
EN627 | 20190204 | L10 | 24 | 6 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 266.3 | 154 | 0.628 | 265.672 | 153.372 | 1.7482 | 0.32469 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 111 | very poor replication (47% range) but large cells throughout water column; Taylor previously QC'ed; DF thinks just biological variation -- winter cruise so lots of large cells
EN627 | 20190204 | L10 | 24 | 6 |  |  | b | 354 | 5 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 144.1 | 83.22 | 0.628 | 143.472 | 82.592 | 2.8432 | 0.50561 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 112 | very poor replication (47% range) but large cells throughout water column, and this sample diluted which can cause more error; Taylor previously QC'ed; DF thinks this is just bio variation -- winter cruise so lots of large cells
EN627 | 20190204 | L10 | 24 | 6 |  |  | a | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 271.1 | 156.3 | 0.628 | 270.472 | 155.672 | 1.7871 | 0.31686 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 117
EN627 | 20190204 | L10 | 24 | 6 |  |  | b | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 306.5 | 178.1 | 0.628 | 305.872 | 177.472 | 1.9988 | 0.39978 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 118
EN627 | 20190204 | L10 | 24 | 6 |  |  | a | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 397.4 | 236.5 | 0.74183 | 396.6582 | 235.7582 | 3.1003 | 0.84367 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | somewhat poor replication (29% range) but most of cast dealing with patchiness of large cells
EN627 | 20190204 | L10 | 24 | 6 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 294.6 | 175 | 0.74183 | 293.8582 | 174.2582 | 2.3126 | 0.61278 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | somewhat poor replication (29% range) but most of cast dealing with patchiness of large cells
EN627 | 20190204 | L10 | 24 | 6 |  |  | a | 285 | 10 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 179.2 | 105.7 | 0.54233 | 178.6577 | 105.1577 | 2.8424 | 0.68828 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | ??<10um nitex not in filter funnel, approx match to WSW?? Def happened on cruise, no notes of which casts
EN627 | 20190204 | L10 | 24 | 6 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 98.77 | 62.67 | 0.74183 | 98.0282 | 61.9282 | 0.69803 | 0.34159 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190204 | L10 | 24 | 8 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 135.9 | 80.33 | 0.628 | 135.272 | 79.702 | 0.86507 | 0.21214 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 113 | very poor replication (64% range) but large cells in water column, could this have been switched with >20a? Toss this since value less than both >20's; Taylor previously QC'ed
EN627 | 20190204 | L10 | 24 | 8 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 263.7 | 155.8 | 0.628 | 263.072 | 155.172 | 1.6797 | 0.41752 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 114 | very poor replication (64% range) but large cells in water column; Taylor previously QC'ed; QF was 2 but DF changed to 1
EN627 | 20190204 | L10 | 24 | 8 |  |  | a | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 207.5 | 120 | 0.628 | 206.872 | 119.372 | 1.3621 | 0.25124 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 119 | could this have been swapped with >5a?
EN627 | 20190204 | L10 | 24 | 8 |  |  | b | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 171.9 | 99.39 | 0.628 | 171.272 | 98.762 | 1.1288 | 0.20603 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 120
EN627 | 20190204 | L10 | 24 | 8 |  |  | a | 287 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 134.3 | 80.19 | 0.58083 | 133.7192 | 79.6092 | 2.078 | 0.5763 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L10 | 24 | 8 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 318.1 | 189.1 | 0.74183 | 317.3582 | 188.3582 | 2.4944 | 0.66772 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L10 | 24 | 8 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 283.2 | 170.3 | 0.54233 | 282.6577 | 169.7577 | 2.183 | 0.66678 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | ??<10um nitex not in filter funnel, approx match to WSW?? Def happened on cruise, no notes of which casts
EN627 | 20190204 | L10 | 24 | 8 |  |  | b | 287 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 99.85 | 63.18 | 0.54233 | 99.3077 | 62.6377 | 0.70411 | 0.34009 | 20181101 | 9-77 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190204 | L10 | 24 | 9 |  |  | a | 360 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 168.9 | 100.2 | 0.628 | 168.272 | 99.572 | 1.0516 | 0.27169 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 109
EN627 | 20190204 | L10 | 24 | 9 |  |  | b | 360 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 180.3 | 107.1 | 0.628 | 179.672 | 106.472 | 1.1205 | 0.29451 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 110
EN627 | 20190204 | L10 | 24 | 9 |  |  | a | 360 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 93.09 | 54.89 | 0.628 | 92.462 | 54.262 | 0.58476 | 0.1364 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 115
EN627 | 20190204 | L10 | 24 | 9 |  |  | b | 360 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 107 | 62.01 | 0.628 | 106.372 | 61.382 | 0.68869 | 0.12708 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 116
EN627 | 20190204 | L10 | 24 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 257 | 156 | 0.74183 | 256.2582 | 155.2582 | 1.9193 | 0.6422 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L10 | 24 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 254.6 | 152.7 | 0.58083 | 254.0192 | 152.1192 | 1.9364 | 0.57331 | 20181101 | 9-75 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L10 | 24 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 95.19 | 60.41 | 0.54233 | 94.6477 | 59.8677 | 0.66091 | 0.32679 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | odd that N9 <10 ok but other <10 on C24 bad, we sometimes have 4 funnels out
EN627 | 20190204 | L10 | 24 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 81.93 | 51.84 | 0.54233 | 81.3877 | 51.2977 | 0.57179 | 0.27452 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | odd that N9 <10 ok but other <10 on C24 bad, we sometimes have 4 funnels out
EN627 | 20190204 | L11 | 26 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 18.77 | 11.94 | 0.69017 | 18.0798 | 11.2498 | 0.12979 | 0.055813 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | poor replication, odd that more of a chl signal and phaeo not higher than chl at 500m so toss this one
EN627 | 20190204 | L11 | 26 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.865 | 5.809 | 0.69017 | 7.1748 | 5.1188 | 0.03907 | 0.045382 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication but flag as 2 because these values more believable for 500m depth
EN627 | 20190204 | L11 | 26 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 95.35 | 58.28 | 0.69017 | 94.6598 | 57.5898 | 0.70443 | 0.2457 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L11 | 26 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 78.75 | 48.06 | 0.69017 | 78.0598 | 47.3698 | 0.58319 | 0.19832 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L11 | 26 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 54.77 | 34.09 | 0.69017 | 54.0798 | 33.3998 | 0.39298 | 0.15806 | 20181101 | 9-71 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication
EN627 | 20190204 | L11 | 26 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 33.03 | 21.6 | 0.69017 | 32.3398 | 20.9098 | 0.2172 | 0.12777 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication
EN627 | 20190204 | L11 | 26 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 102.2 | 62.05 | 0.69017 | 101.5098 | 61.3598 | 0.76296 | 0.24936 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication. Which would yuou toss?
EN627 | 20190204 | L11 | 26 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 51.23 | 32.72 | 0.69017 | 50.5398 | 32.0298 | 0.35174 | 0.17669 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication. Which would yuou toss?
EN627 | 20190204 | L11 | 26 | 6 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.853 | 1.048 | 0.628 | 1.225 | 0.42 | 0.012532 | -0.0068551 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR 123 | how are these values basically zero when >20's are almost 1 for both reps, is this the >20s?; Taylor previously QC'ed, DF thinks >5 reps and >20 reps got switched or maybe not water got poured over? Seems unlikely this is a 20 rep given WSW values and 20 values in the correct spot
EN627 | 20190204 | L11 | 26 | 6 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 0.991 | 0.836 | 0.628 | 0.363 | 0.208 | 0.0024129 | 0.0003983 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR 125 | how are these values basically zero when >20's are almost 1 for both reps, is this the >20s?; Taylor previously QC'ed, DF thinks >5 reps and >20 reps got switched; actually on second thought, DF doesn't think water got poured through these filters
EN627 | 20190204 | L11 | 26 | 6 |  |  | a | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 200.7 | 114.9 | 0.628 | 200.072 | 114.272 | 1.3357 | 0.20877 | 20181101 | 9-81 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR 130 | how are these values of 1 when >5's virtually zero, could these be the >5's; Taylor previously QC'ed, DF thinks >5 reps and >20 reps got switched
EN627 | 20190204 | L11 | 26 | 6 |  |  | b | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 133.7 | 78 | 0.628 | 133.072 | 77.372 | 0.86709 | 0.17862 | 20181101 | 9-81 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR 132 | how are these values of 1 when >5's virtually zero, could these be the >5's; Taylor previously QC'ed, DF thinks >5 reps and >20 reps got switched
EN627 | 20190204 | L11 | 26 | 6 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 266.8 | 158.9 | 0.69017 | 266.1098 | 158.2098 | 2.0864 | 0.5696 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L11 | 26 | 6 |  |  | b | 287 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 254.1 | 152.5 | 0.69017 | 253.4098 | 151.8098 | 1.9509 | 0.57991 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L11 | 26 | 6 |  |  | a | 287 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 106.6 | 63.86 | 0.69017 | 105.9098 | 63.1698 | 0.82066 | 0.23241 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | C26N6<10a logsheet note "10a chl lost volume" but was it actually <10b
EN627 | 20190204 | L11 | 26 | 6 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 64.75 | 41.88 | 0.69017 | 64.0598 | 41.1898 | 0.44222 | 0.24926 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | don't trust either of them, 10a is at least similar values to other <10 depths on this cast
EN627 | 20190204 | L11 | 26 | 7 |  |  | a | 360 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 229.7 | 135.3 | 0.628 | 229.072 | 134.672 | 1.445 | 0.34476 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 122 | Taylor previously QC'ed
EN627 | 20190204 | L11 | 26 | 7 |  |  | b | 360 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 167 | 97.59 | 0.628 | 166.372 | 96.962 | 1.0625 | 0.22613 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 124 | could this be >20a?; 20s look like all water poured into 1 filter so no basis for switching this to a 20; DF thinks this is within bio variation of size fractionated filters. QF was 2 but changed to 1
EN627 | 20190204 | L11 | 26 | 7 |  |  | a | 360 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 333 | 192.8 | 0.628 | 332.372 | 192.172 | 2.1461 | 0.40785 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 129 | >20 larger than >5? Could this be the >5b?; 20b didn't have any water poured onto filter so is this double concentration on 20 filter? DF thinks this is definitely a double pour and 20 value should be half this
EN627 | 20190204 | L11 | 26 | 7 |  |  | b | 360 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.064 | 1.056 | 0.628 | 0.436 | 0.428 | 0.00012246 | 0.0055657 | 20181101 | 9-80 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 131 | clearly something went wrong with this sample, Rb = 1.064 Ra=1.056; Taylor previously QC'ed
EN627 | 20190204 | L11 | 26 | 7 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 377.8 | 224 | 0.69017 | 377.1098 | 223.3098 | 2.9739 | 0.77494 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication (43% range)
EN627 | 20190204 | L11 | 26 | 7 |  |  | b | 286 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 191.8 | 111.6 | 0.74183 | 191.0582 | 110.8582 | 4.636 | 0.92761 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication but one was diluted
EN627 | 20190204 | L11 | 26 | 7 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 334.3 | 199.8 | 0.69017 | 333.6098 | 199.1098 | 2.6007 | 0.74187 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | ??<10um nitex not in filter funnel, approx match to WSW?? Def happened on cruise, no notes of which casts.
EN627 | 20190204 | L11 | 26 | 7 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 123.6 | 78.61 | 0.69017 | 122.9098 | 77.9198 | 0.86993 | 0.43816 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190204 | L11 | 26 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 317.4 | 188.7 | 0.74183 | 316.6582 | 187.9582 | 2.4456 | 0.65531 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L11 | 26 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 314.3 | 188.7 | 0.74183 | 313.5582 | 187.9582 | 2.3867 | 0.71422 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L11 | 26 | 11 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 292.9 | 174.1 | 0.628 | 292.272 | 173.472 | 1.8494 | 0.49517 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 121
EN627 | 20190204 | L11 | 26 | 11 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 312.5 | 181.5 | 0.628 | 311.872 | 180.872 | 2.0393 | 0.40526 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 127
EN627 | 20190204 | L11 | 26 | 11 |  |  | a | 354 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 283 | 167.6 | 0.628 | 282.372 | 166.972 | 1.7965 | 0.46025 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 126 | very poore replication (72% range) vol filt incorrect?; Taylor previously QC'ed
EN627 | 20190204 | L11 | 26 | 11 |  |  | b | 150 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 249 | 144.2 | 0.628 | 248.372 | 143.572 | 3.8502 | 0.72923 | 20181101 | 9-79 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 128 | very poore replication (72% range) vol filt incorrect? This value much higher than both WSW reps so toss; Taylor previously QC'ed
EN627 | 20190204 | L11 | 26 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 394.5 | 233.2 | 0.74183 | 393.7582 | 232.4582 | 3.0651 | 0.76999 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L11 | 26 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 392.7 | 236.8 | 0.69017 | 392.0098 | 236.1098 | 2.9625 | 0.93285 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190204 | L11 | 26 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 334.6 | 199.7 | 0.54233 | 334.0577 | 199.1577 | 2.5635 | 0.72227 | 20181101 | 9-76 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | ??<10um nitex not in filter funnel part of this cruise nitext was missing from a filter funnel, approx match to WSW?? Def happened on cruise, no notes of which casts
EN627 | 20190204 | L11 | 26 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 130 | 82.32 | 0.69017 | 129.3098 | 81.6298 | 0.90605 | 0.44069 | 20181101 | 9-70 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190205 | u11c | 29 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.29 | 2.815 | 0.69017 | 2.5998 | 2.1248 | 0.0090263 | 0.02603 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | in noise of fluorometer
EN627 | 20190205 | u11c | 29 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.508 | 3.117 | 0.69017 | 2.8178 | 2.4268 | 0.0074301 | 0.032608 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | in noise of fluorometer
EN627 | 20190205 | u11c | 29 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.331 | 2.757 | 0.69017 | 2.6408 | 2.0668 | 0.010908 | 0.023191 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor repilcation but values of zero. In noise of fluorometer
EN627 | 20190205 | u11c | 29 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.629 | 3.433 | 0.74183 | 2.8872 | 2.6912 | 0.0037245 | 0.040675 | 20181101 | 9-72 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor repilcation but values of zero. In noise of fluorometer
EN627 | 20190205 | u11c | 29 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 8.137 | 6.486 | 0.69017 | 7.4468 | 5.7958 | 0.031373 | 0.064247 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190205 | u11c | 29 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 8.199 | 6.89 | 0.69017 | 7.5088 | 6.1998 | 0.024875 | 0.077411 | 20181101 | 9-71 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190205 | u11c | 29 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 81.08 | 50.72 | 0.69017 | 80.3898 | 50.0298 | 0.57692 | 0.24848 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190205 | u11c | 29 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 82.65 | 52.23 | 0.69017 | 81.9598 | 51.5398 | 0.57806 | 0.27225 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190205 | u11c | 29 | 6 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 46.52 | 28.74 | 0.543 | 45.977 | 28.197 | 0.27678 | 0.10431 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 147 | Taylor previously QC'ed
EN627 | 20190205 | u11c | 29 | 6 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 39.05 | 24.96 | 0.543 | 38.507 | 24.417 | 0.21934 | 0.11066 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 148 | Taylor previously QC'ed
EN627 | 20190205 | u11c | 29 | 6 |  |  | a | 354 | 20 | 2 | 1 | 0 | 0.66667 | Cal_110119 | 1.8682 | 0.5122 | 158.7 | 94.75 | 0.543 | 158.157 | 94.207 | 0.59731 | 0.16664 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 153 | extracted in 2ml cryo vials because ran out of 5ml vials on cruise, pipetted 1.8ml sample with 0.9ml acetone to have enough volume for fluorometer reading; Taylor previously QC'ed
EN627 | 20190205 | u11c | 29 | 6 |  |  | b | 354 | 20 | 2 | 1 | 0 | 0.66667 | Cal_110119 | 1.8682 | 0.5122 | 88.65 | 51.57 | 0.543 | 88.107 | 51.027 | 0.34634 | 0.067452 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 154 | extracted in 2ml cryo vials because ran out of 5ml vials on cruise, pipetted 1.8ml sample with 0.9ml acetone to have enough volume for fluorometer reading; Taylor previously QC'ed
EN627 | 20190205 | u11c | 29 | 6 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 106.8 | 67.3 | 0.69017 | 106.1098 | 66.6098 | 0.76377 | 0.35444 | 20181101 | 9-71 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190205 | u11c | 29 | 6 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 107.5 | 65.45 | 0.69017 | 106.8098 | 64.7598 | 0.81308 | 0.27408 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190205 | u11c | 29 | 6 |  |  | a | 286 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 27.82 | 18.43 | 0.69017 | 27.1298 | 17.7398 | 0.18093 | 0.11584 | 20181101 | 9-71 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190205 | u11c | 29 | 6 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 29.04 | 19.83 | 0.69017 | 28.3498 | 19.1398 | 0.17809 | 0.14323 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190205 | u11c | 29 | 8 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 85.67 | 50.49 | 0.543 | 85.127 | 49.947 | 0.54765 | 0.1274 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 149 | value very high relative to WSW compared to pattern on rest of cast, flag as 2; DF thinks that all chl signal is just >5, so this is why these values are close to WSW values
EN627 | 20190205 | u11c | 29 | 8 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 107.9 | 65.38 | 0.543 | 107.357 | 64.837 | 0.66192 | 0.21438 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 150 | value higher than WSW; DF thinks that all chl signal is just >5, so this is why these values are close to WSW values
EN627 | 20190205 | u11c | 29 | 8 |  |  | a | 354 | 20 | 2 | 1 | 0 | 0.66667 | Cal_110119 | 1.8682 | 0.5122 | 38.13 | 25.61 | 0.543 | 37.587 | 25.067 | 0.11694 | 0.086334 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR 155 | poor replication, which to believe?extracted in 2ml cryo vials because ran out of 5ml vials on cruise, pipetted 1.8ml sample with 0.9ml acetone to have enough volume for fluorometer reading
EN627 | 20190205 | u11c | 29 | 8 |  |  | b | 354 | 20 | 2 | 1 | 0 | 0.66667 | Cal_110119 | 1.8682 | 0.5122 | 55.39 | 34.95 | 0.543 | 54.847 | 34.407 | 0.19092 | 0.088099 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR 156 | poor replication, which to believe?extracted in 2ml cryo vials because ran out of 5ml vials on cruise, pipetted 1.8ml sample with 0.9ml acetone to have enough volume for fluorometer reading
EN627 | 20190205 | u11c | 29 | 8 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 81.85 | 52.77 | 0.69017 | 81.1598 | 52.0798 | 0.56229 | 0.312 | 20181101 | 9-71 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190205 | u11c | 29 | 8 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 67.09 | 43.41 | 0.69017 | 66.3998 | 42.7198 | 0.45788 | 0.25928 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190205 | u11c | 29 | 8 |  |  | a | 287 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 29.94 | 20.36 | 0.69017 | 29.2498 | 19.6698 | 0.18395 | 0.14396 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190205 | u11c | 29 | 8 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 30.97 | 20.72 | 0.54233 | 30.4277 | 20.1777 | 0.19819 | 0.14054 | 20181101 | 9-76 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190205 | u11c | 29 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 51.71 | 34.75 | 0.69017 | 51.0198 | 34.0598 | 0.32229 | 0.23964 | 20181101 | 9-71 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190205 | u11c | 29 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 62.54 | 41.73 | 0.69017 | 61.8498 | 41.0398 | 0.39545 | 0.28163 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190205 | u11c | 29 | 10 |  |  | a | 360 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 20.7 | 13.79 | 0.543 | 20.157 | 13.247 | 0.10578 | 0.070278 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 145 | poor replication (30% range) ; Taylor previously QC'ed; DF thinks within bio variation; QF was 2, DF changed to 1
EN627 | 20190205 | u11c | 29 | 10 |  |  | b | 360 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 27.85 | 18.46 | 0.543 | 27.307 | 17.917 | 0.14374 | 0.09438 | 20181101 | 9-74 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 146 | poor replication (30% range) ; Taylor previously QC'ed; DF thinks within bio variation; QF was 2, DF changed to 1
EN627 | 20190205 | u11c | 29 | 10 |  |  | a | 360 | 20 | 2 | 1 | 0 | 0.66667 | Cal_110119 | 1.8682 | 0.5122 | 27 | 17.95 | 0.543 | 26.457 | 17.407 | 0.083121 | 0.055684 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR 151 | extracted in 2ml cryo vials because ran out of 5ml vials on cruise, pipetted 1.8ml sample with 0.9ml acetone to have enough volume for fluorometer reading; Taylor previously QC'ed, replication isn't bad, QF was 2
EN627 | 20190205 | u11c | 29 | 10 |  |  | b | 360 | 20 | 2 | 1 | 0 | 0.66667 | Cal_110119 | 1.8682 | 0.5122 | 18.53 | 12.12 | 0.543 | 17.987 | 11.577 | 0.058874 | 0.033443 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR 152 | extracted in 2ml cryo vials because ran out of 5ml vials on cruise, pipetted 1.8ml sample with 0.9ml acetone to have enough volume for fluorometer reading; Taylor previously QC'ed, replication isn't bad, QF was 2
EN627 | 20190205 | u11c | 29 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 55.46 | 36.43 | 0.69017 | 54.7698 | 35.7398 | 0.36162 | 0.22802 | 20181101 | 9-71 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190205 | u11c | 29 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 50.59 | 33.81 | 0.69017 | 49.8998 | 33.1198 | 0.31887 | 0.22755 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190205 | u11c | 29 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 29.65 | 19.89 | 0.69017 | 28.9598 | 19.1998 | 0.18547 | 0.1313 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190205 | u11c | 29 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 28.53 | 19.24 | 0.69017 | 27.8398 | 18.5498 | 0.17654 | 0.1295 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN627 | 20190206 | MVCO | 45 | 1 |  |  | a | 354 | 5 | 5 | 1 | 0 | 0.33 | Cal_110119 | 1.8682 | 0.5122 | 148.1 | 88.37 | 0.7025 | 147.3975 | 87.6675 | 2.8177 | 0.77284 | 20181101 | 9-82 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR 135 | sample was originally labeled as >20a but values indicate mistakenly swapped and should be >5a
EN627 | 20190206 | MVCO | 45 | 1 |  |  | b | 354 | 5 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 158.9 | 102.3 | 0.7025 | 158.1975 | 101.5975 | 2.6433 | 1.4761 | 20181101 | 9-82 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR 138 | very slow filtering
EN627 | 20190206 | MVCO | 45 | 1 |  |  | a | 374 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 357.4 | 231 | 0.7025 | 356.6975 | 230.2975 | 1.8625 | 1.0837 | 20181101 | 9-82 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR 143 | sample was originally labeled as >5a but values indicate mistakenly swapped and should be >20a
EN627 | 20190206 | MVCO | 45 | 1 |  |  | b | 374 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 251.4 | 150.3 | 0.7025 | 250.6975 | 149.5975 | 1.4897 | 0.42407 | 20181101 | 9-82 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR 144
EN627 | 20190206 | MVCO | 45 | 1 |  |  | a | 287 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 235.4 | 150.2 | 0.69017 | 234.7098 | 149.5098 | 3.2719 | 1.7129 | 20181101 | 9-71 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication (56% range)
EN627 | 20190206 | MVCO | 45 | 1 |  |  | b | 287 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 131.7 | 84.16 | 0.69017 | 131.0098 | 83.4698 | 1.8257 | 0.95732 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication (56% range)
EN627 | 20190206 | MVCO | 45 | 1 |  |  | a | 102 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 23.35 | 15.9 | 0.69017 | 22.6598 | 15.2098 | 0.4025 | 0.31094 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | <10 taking long time to filter, used less volume
EN627 | 20190206 | MVCO | 45 | 1 |  |  | b | 255 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 45.48 | 31.24 | 0.69017 | 44.7898 | 30.5498 | 0.30774 | 0.26545 | 20181101 | 9-71 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | <10 taking long time to filter, used less volume
EN627 | 20190206 | MVCO | 45 | 5 |  |  | a | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 292.6 | 182.8 | 0.543 | 292.057 | 182.257 | 1.7093 | 0.754 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 134 | very slow filtering; Taylor previously QC'ed
EN627 | 20190206 | MVCO | 45 | 5 |  |  | b | 354 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 327.5 | 205.3 | 0.7025 | 326.7975 | 204.5975 | 1.9023 | 0.86291 | 20181101 | 9-82 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 137 | very slow filtering; Taylor previously QC'ed
EN627 | 20190206 | MVCO | 45 | 5 |  |  | a | 334 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 340.1 | 201.2 | 0.7025 | 339.3975 | 200.4975 | 2.2918 | 0.58031 | 20181101 | 9-82 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 141 | very poor replication (69% range) and this basically higher than both WSW values so toss; Taylor previously QC'ed
EN627 | 20190206 | MVCO | 45 | 5 |  |  | b | 374 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 179.8 | 104.5 | 0.543 | 179.257 | 103.957 | 1.1095 | 0.22036 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 142 | very poor replication (69%); Taylor previously QC'ed
EN627 | 20190206 | MVCO | 45 | 5 |  |  | a | 286 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 158 | 102.7 | 0.69017 | 157.3098 | 102.0098 | 2.1311 | 1.2819 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190206 | MVCO | 45 | 5 |  |  | b | 285 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 164.4 | 105.9 | 0.69017 | 163.7098 | 105.2098 | 2.2623 | 1.2701 | 20181101 | 9-71 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190206 | MVCO | 45 | 5 |  |  | a | 263 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 23.62 | 16.73 | 0.69017 | 22.9298 | 16.0398 | 0.14437 | 0.14742 | 20181101 | 9-71 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | <10 taking long time to filter, used less volume
EN627 | 20190206 | MVCO | 45 | 5 |  |  | b | 216 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 31.36 | 21.33 | 0.69017 | 30.6698 | 20.6398 | 0.25589 | 0.20128 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | <10 taking long time to filter, used less volume
EN627 | 20190206 | MVCO | 45 | 9 |  |  | a | 360 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 224.3 | 141.2 | 0.543 | 223.757 | 140.657 | 1.2721 | 0.59728 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 133 | very slow filtering; Taylor previously QC'ed
EN627 | 20190206 | MVCO | 45 | 9 |  |  | b | 360 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 248.7 | 157.1 | 0.543 | 248.157 | 156.557 | 1.4022 | 0.67848 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 136 | very slow filtering; Taylor previously QC'ed
EN627 | 20190206 | MVCO | 45 | 9 |  |  | a | 360 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 203.8 | 119.6 | 0.543 | 203.257 | 119.057 | 1.2889 | 0.29338 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR 139 | poor replication (31% range) nut isn't too bad for the >20 size fraction
EN627 | 20190206 | MVCO | 45 | 9 |  |  | b | 360 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 147.8 | 86.51 | 0.543 | 147.257 | 85.967 | 0.93821 | 0.20431 | 20181101 | 9-73 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR 140 | poor replication (31% range) nut isn't too bad for the >20 size fraction
EN627 | 20190206 | MVCO | 45 | 9 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 133.7 | 85.27 | 0.69017 | 133.0098 | 84.5798 | 1.8406 | 0.95022 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190206 | MVCO | 45 | 9 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 140.2 | 90.93 | 0.69017 | 139.5098 | 90.2398 | 1.8725 | 1.1051 | 20181101 | 9-71 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN627 | 20190206 | MVCO | 45 | 9 |  |  | a | 211 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 36.45 | 25.14 | 0.69017 | 35.7598 | 24.4498 | 0.29539 | 0.25902 | 20181101 | 9-71 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | <10 taking long time to filter, used less volume
EN627 | 20190206 | MVCO | 45 | 9 |  |  | b | 214 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 44.54 | 29.36 | 0.69017 | 43.8498 | 28.6698 | 0.3909 | 0.25007 | 20181101 | 9-70 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | <10 taking long time to filter, used less volume
AR28A | 20180324 | L10 | 1 | 5 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 231.6 | 145.4 | 0.89267 | 230.7073 | 144.5073 | 0.89883 | 0.40939 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 01/01
AR28A | 20180324 | L10 | 1 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 238.7 | 152 | 0.89267 | 237.8073 | 151.1073 | 0.9049 | 0.46436 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 01/02
AR28A | 20180324 | L10 | 1 | 7 |  |  | a | 1056 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! |  |  |  | 4 |  |  |  |  | 0 | 01/03
AR28A | 20180324 | L10 | 1 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 181 | 113.9 | 0.89267 | 180.1073 | 113.0073 | 0.70033 | 0.32368 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 01/04
AR28A | 20180324 | L8 | 2 | 3 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 279.6 | 175.7 | 0.89267 | 278.7073 | 174.8073 | 1.0834 | 0.49913 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/01
AR28A | 20180324 | L8 | 2 | 4 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 193.3 | 125.5 | 0.89267 | 192.4073 | 124.6073 | 1.0615 | 0.63224 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/02
AR28A | 20180324 | L8 | 2 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 195.1 | 121.8 | 0.89267 | 194.2073 | 120.9073 | 0.76504 | 0.33056 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/03
AR28A | 20180324 | L8 | 2 | 6 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 147.3 | 92.48 | 0.89267 | 146.4073 | 91.5873 | 0.85824 | 0.38663 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/04
AR28A | 20180324 | L8 | 2 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 253.8 | 161.3 | 0.89267 | 252.9073 | 160.4073 | 0.48226 | 0.24382 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/05
AR28A | 20180324 | L8 | 2 | 8 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 250.4 | 160.2 | 0.89267 | 249.5073 | 159.3073 | 0.47071 | 0.25107 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/06
AR28A | 20180325 | L6 | 3 | 1 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 224.1 | 146.6 | 0.89267 | 223.2073 | 145.7073 | 0.80811 | 0.51097 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 03/01
AR28A | 20180325 | L6 | 3 | 3 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 323.7 | 197.3 | 0.89267 | 322.8073 | 196.4073 | 0.65962 | 0.23025 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 03/02
AR28A | 20180325 | L6 | 3 | 5 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 342.8 | 210.7 | 0.89267 | 341.9073 | 209.8073 | 0.68937 | 0.26121 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 03/03
AR28A | 20180325 | L6 | 3 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 281 | 172.6 | 0.89267 | 280.1073 | 171.7073 | 0.56569 | 0.21227 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 03/04
AR28B | 20180403 | L1 | 1 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 178.9 | 113.3 | 0.81033 | 178.0897 | 112.4897 | 1.2466 | 0.60929 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180403 | L1 | 1 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 163.7 | 101.8 | 0.81033 | 162.8897 | 100.9897 | 1.1763 | 0.48987 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180403 | L1 | 1 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 125.9 | 78.4 | 0.81033 | 125.0897 | 77.5897 | 0.90263 | 0.37746 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180403 | L1 | 1 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 117 | 72.21 | 0.81033 | 116.1897 | 71.3997 | 0.85113 | 0.32683 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180403 | L1 | 1 | 8 |  |  | a | 288 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 114 | 72.29 | 0.81033 | 113.1897 | 71.4797 | 0.79811 | 0.38936 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180403 | L1 | 1 | 8 |  |  | b | 289 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 123.1 | 76.37 | 0.81033 | 122.2897 | 75.5597 | 0.89107 | 0.35984 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180403 | L1 | 1 | 8 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 73.1 | 45.66 | 0.81033 | 72.2897 | 44.8497 | 0.53058 | 0.22234 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication (47% range)
AR28B | 20180403 | L1 | 1 | 8 |  |  | b | 286 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 113.9 | 69.76 | 0.81033 | 113.0897 | 68.9497 | 0.85051 | 0.30294 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | very poor replication (47% range) should i toss this one since basiaclly same as WSW? Pattern of entire 5 stations of sampling one <10 sample off, so toss the bad rep of higher value for each of the first 5 stations of cruise
AR28B | 20180403 | L2 | 2 | 1 |  |  | a | 290 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 143.2 | 93.72 | 0.81033 | 142.3897 | 92.9097 | 5.6415 | 3.5555 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180403 | L2 | 2 | 1 |  |  | b | 290 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 165.4 | 105.9 | 0.81033 | 164.5897 | 105.0897 | 6.784 | 3.6187 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180403 | L2 | 2 | 5 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 362.9 | 237.1 | 0.81033 | 362.0897 | 236.2897 | 4.7811 | 3.0156 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180403 | L2 | 2 | 5 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 366.3 | 236.6 | 0.81033 | 365.4897 | 235.7897 | 4.9293 | 2.8509 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180403 | L2 | 2 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 72.19 | 46.04 | 0.81033 | 71.3797 | 45.2297 | 0.49692 | 0.24928 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180403 | L2 | 2 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 66.92 | 42.64 | 0.81033 | 66.1097 | 41.8297 | 0.46139 | 0.22873 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180403 | L2 | 2 | 11 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 65.23 | 40.81 | 0.81033 | 64.4197 | 39.9997 | 0.47219 | 0.19931 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180403 | L2 | 2 | 11 |  |  | b | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 71.06 | 43.73 | 0.81033 | 70.2497 | 42.9197 | 0.52661 | 0.19139 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180403 | L2 | 2 | 11 |  |  | a | 289 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 53.41 | 33.45 | 0.81033 | 52.5997 | 32.6397 | 0.38061 | 0.15975 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (25% range) is this a pattern with one of the <10 funnels?
AR28B | 20180403 | L2 | 2 | 11 |  |  | b | 288 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 66.7 | 40.98 | 0.81033 | 65.8897 | 40.1697 | 0.49214 | 0.17518 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | poor replication (25% range) is this a pattern with one of the <10 funnels? This is basically same as WSW, should I toss?
AR28B | 20180404 | L3 | 3 | 2 |  |  | a | 200 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 222.8 | 134.3 | 0.81033 | 221.9897 | 133.4897 | 14.6311 | 4.5292 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L3 | 3 | 2 |  |  | b | 187 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 192.9 | 119.4 | 0.81033 | 192.0897 | 118.5897 | 12.996 | 5.2089 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L3 | 3 | 4 |  |  | a | 290 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 233.1 | 143.1 | 0.81033 | 232.2897 | 142.2897 | 10.2615 | 3.8236 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L3 | 3 | 4 |  |  | b | 290 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 224.7 | 136 | 0.81033 | 223.8897 | 135.1897 | 10.1132 | 3.269 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L3 | 3 | 6 |  |  | a | 390 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 270.3 | 159.9 | 0.81033 | 269.4897 | 159.0897 | 9.3598 | 2.3503 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor phaeo ratio agreement (48% phaeo range) both reps are dilutions
AR28B | 20180404 | L3 | 3 | 6 |  |  | b | 290 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 248.4 | 152.1 | 0.81033 | 247.5897 | 151.2897 | 10.9798 | 3.9962 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor phaeo ratio agreement (48% phaeo range) both reps are dilutions
AR28B | 20180404 | L3 | 3 | 7 |  |  | a | 282 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 112.5 | 70.53 | 0.81033 | 111.6897 | 69.7197 | 0.82017 | 0.36271 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (23% range)
AR28B | 20180404 | L3 | 3 | 7 |  |  | b | 283 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 138.2 | 84.97 | 0.81033 | 137.3897 | 84.1597 | 1.0365 | 0.38629 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (23% range)
AR28B | 20180404 | L3 | 3 | 9 |  |  | a | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 107.1 | 65.33 | 0.81033 | 106.2897 | 64.5197 | 0.80484 | 0.2745 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L3 | 3 | 9 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 123.4 | 75.1 | 0.81033 | 122.5897 | 74.2897 | 0.93393 | 0.31321 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L3 | 3 | 9 |  |  | a | 289 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 79.77 | 49.27 | 0.81033 | 78.9597 | 48.4597 | 0.58159 | 0.22067 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | poor replication (32% range) is this a pattern with one of the <10 filters?
AR28B | 20180404 | L3 | 3 | 9 |  |  | b | 288 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 57.95 | 36.11 | 0.81033 | 57.1397 | 35.2997 | 0.4179 | 0.16852 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication (32% range) is this a pattern with one of the <10 filters?
AR28B | 20180404 | L4 | 4 | 2 |  |  | a | 218 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 342.5 | 213.9 | 0.81033 | 341.6897 | 213.0897 | 6.5017 | 2.8517 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L4 | 4 | 2 |  |  | b | 224 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 334.1 | 208.7 | 0.81033 | 333.2897 | 207.8897 | 6.1701 | 2.7106 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L4 | 4 | 3 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 278.4 | 171.2 | 0.81033 | 277.5897 | 170.3897 | 4.0742 | 1.548 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L4 | 4 | 3 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 264.4 | 165.2 | 0.81033 | 263.5897 | 164.3897 | 3.7701 | 1.6541 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L4 | 4 | 6 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 288.2 | 176.7 | 0.81033 | 287.3897 | 175.8897 | 4.2376 | 1.5661 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L4 | 4 | 6 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 274 | 169.1 | 0.81033 | 273.1897 | 168.2897 | 3.9868 | 1.5662 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L4 | 4 | 7 |  |  | a | 282 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 238.8 | 144.8 | 0.81033 | 237.9897 | 143.9897 | 1.8369 | 0.60603 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L4 | 4 | 7 |  |  | b | 283 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 129.8 | 77.81 | 0.81033 | 128.9897 | 76.9997 | 2.0248 | 0.57877 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L4 | 4 | 10 |  |  | a | 289 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 214.7 | 130.6 | 0.81033 | 213.8897 | 129.7897 | 1.6037 | 0.54504 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L4 | 4 | 10 |  |  | b | 288 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 219.1 | 132.3 | 0.81033 | 218.2897 | 131.4897 | 1.6609 | 0.52351 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L4 | 4 | 10 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 106 | 64.16 | 0.81033 | 105.1897 | 63.3497 | 0.80902 | 0.25447 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication (53% range) and pattern of 1 bad <10 for each of these stations the first night sampling, this seems more reasonable
AR28B | 20180404 | L4 | 4 | 10 |  |  | b | 286 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 184 | 111.2 | 0.81033 | 183.1897 | 110.3897 | 1.4027 | 0.44395 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | very poor replication (53% range) and pattern of 1 bad <10 for each of these stations. Toss, value basically same as WSW
AR28B | 20180404 | L5 | 5 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 77.89 | 50.21 | 0.81033 | 77.0797 | 49.3997 | 0.52599 | 0.28901 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L5 | 5 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 73.51 | 47.72 | 0.81033 | 72.6997 | 46.9097 | 0.49008 | 0.28384 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L5 | 5 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 78.72 | 51.44 | 0.81033 | 77.9097 | 50.6297 | 0.51839 | 0.3169 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L5 | 5 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 82.26 | 53.22 | 0.81033 | 81.4497 | 52.4097 | 0.55184 | 0.31282 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L5 | 5 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 115.9 | 72.94 | 0.81033 | 115.0897 | 72.1297 | 0.81636 | 0.37365 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L5 | 5 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 98.88 | 62.8 | 0.81033 | 98.0697 | 61.9897 | 0.68562 | 0.3371 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L5 | 5 | 7 |  |  | a | 283 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 141.8 | 87.1 | 0.81033 | 140.9897 | 86.2897 | 1.0652 | 0.39367 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L5 | 5 | 7 |  |  | b | 282 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 149.1 | 90.24 | 0.81033 | 148.2897 | 89.4297 | 1.1502 | 0.36705 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L5 | 5 | 10 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 124 | 74.98 | 0.81033 | 123.1897 | 74.1697 | 0.94785 | 0.29727 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L5 | 5 | 10 |  |  | b | 286 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 148.8 | 89.12 | 0.81033 | 147.9897 | 88.3097 | 1.1499 | 0.32738 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180404 | L5 | 5 | 10 |  |  | a | 288 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 89.84 | 54.84 | 0.81033 | 89.0297 | 54.0297 | 0.66971 | 0.22787 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (36% range)
AR28B | 20180404 | L5 | 5 | 10 |  |  | b | 289 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 127.8 | 77.21 | 0.81033 | 126.9897 | 76.3997 | 0.96467 | 0.30014 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | poor replication (36% range) and pattern of one not great higher <10 for each of the first 5 stations during a crazy night of sampling, maybe nitex wasn't fullly on for one of the filter funnels, toss each higher value from this night
AR28B | 20180405 | L13 | 6 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 117.2 | 72.69 | 0.89267 | 116.3073 | 71.7973 | 0.45848 | 0.1836 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/01
AR28B | 20180405 | L13 | 6 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 110.4 | 68.16 | 0.89267 | 109.5073 | 67.2673 | 0.43509 | 0.16647 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/02
AR28B | 20180405 | L13 | 6 | 5 |  |  | a | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 117 | 74.08 | 0.89267 | 116.1073 | 73.1873 | 0.4324 | 0.20775 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/03
AR28B | 20180405 | L13 | 6 | 5 |  |  | b | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 115.7 | 72.62 | 0.89267 | 114.8073 | 71.7273 | 0.43401 | 0.19337 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/04
AR28B | 20180405 | L13 | 6 | 7 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 110.1 | 68.92 | 0.89267 | 109.2073 | 68.0273 | 0.42025 | 0.18248 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/05
AR28B | 20180405 | L13 | 6 | 7 |  |  | b | 540 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 116.9 | 73.48 | 0.89267 | 116.0073 | 72.5873 | 0.44311 | 0.20002 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/06
AR28B | 20180409 | L10 | 10 | 3 |  |  | a | 485 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.738 | 6.451 | 0.81033 | 6.9277 | 5.6407 | 0.014623 | 0.041021 | 20181101 | 9-66 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor phaeo:chl ratio agreement
AR28B | 20180409 | L10 | 10 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 9.867 | 8.339 | 0.81033 | 9.0567 | 7.5287 | 0.015739 | 0.051589 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor phaeo:chl ratio agreement
AR28B | 20180409 | L10 | 10 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 8.534 | 7.362 | 0.81033 | 7.7237 | 6.5517 | 0.012072 | 0.046519 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180409 | L10 | 10 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 8.646 | 7.421 | 0.81033 | 7.8357 | 6.6107 | 0.012618 | 0.0465 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180409 | L10 | 10 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.51 | 11.8 | 0.81033 | 12.6997 | 10.9897 | 0.017614 | 0.080666 | 20181101 | 9-67 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180409 | L10 | 10 | 6 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 10.24 | 8.622 | 0.81033 | 9.4297 | 7.8117 | 0.016666 | 0.053193 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180409 | L10 | 10 | 7 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 14.6 | 11.48 | 0.81033 | 13.7897 | 10.6697 | 0.032138 | 0.06328 | 20181101 | 9-67 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180409 | L10 | 10 | 7 |  |  | b | 532 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.41 | 10.7 | 0.81033 | 12.5997 | 9.8897 | 0.027914 | 0.060528 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180409 | L10 | 10 | 8 |  |  | a | 520 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 14.71 | 11.35 | 0.81033 | 13.8997 | 10.5397 | 0.035608 | 0.061366 | 20181101 | 9-67 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180409 | L10 | 10 | 8 |  |  | b | 520 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 16.98 | 13.02 | 0.778 | 16.202 | 12.242 | 0.041967 | 0.070671 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180409 | L10 | 10 | 9 |  |  | a | 526 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 133.1 | 87.33 | 0.81033 | 132.2897 | 86.5197 | 0.47952 | 0.30746 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180409 | L10 | 10 | 9 |  |  | b | 530 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 125 | 83.12 | 0.778 | 124.222 | 82.342 | 0.43546 | 0.30787 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180409 | L10 | 10 | 10 |  |  | a | 549 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 166.9 | 109.8 | 0.81033 | 166.0897 | 108.9897 | 0.57316 | 0.37667 | 20181101 | 9-67 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180409 | L10 | 10 | 10 |  |  | b | 549 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 196.5 | 132.2 | 0.778 | 195.722 | 131.422 | 0.64543 | 0.49989 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180409 | L10 | 10 | 11 |  |  | a | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 187.9 | 122.6 | 0.89267 | 187.0073 | 121.7073 | 0.65787 | 0.40667 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 10/01
AR28B | 20180409 | L10 | 10 | 12 |  |  | a | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 192.7 | 125.4 | 0.89267 | 191.8073 | 124.5073 | 0.67802 | 0.41101 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 10/02
AR28B | 20180409 | L10 | 10 | 13 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 195.9 | 124.8 | 0.89267 | 195.0073 | 123.9073 | 0.72559 | 0.37225 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 10/03
AR28B | 20180409 | L10 | 10 | 14 |  |  | a | 540 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 192.9 | 124.2 | 0.89267 | 192.0073 | 123.3073 | 0.70109 | 0.39142 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 10/04
AR28B | 20180409 | L10 | 10 | 16 |  |  | a | 535 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 135.7 | 88.31 | 0.81033 | 134.8897 | 87.4997 | 0.48814 | 0.29436 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (30% range) is filter funnel nitex still messed up?
AR28B | 20180409 | L10 | 10 | 16 |  |  | b | 547 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 182 | 116.3 | 0.81033 | 181.1897 | 115.4897 | 0.6619 | 0.34826 | 20181101 | 9-67 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | poor replication (30% range) is filter funnel nitex still messed up? This is basicall same value as WSW so toss
AR28B | 20180410 | L12 | 11 | 6 |  |  | a | 535 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 346.6 | 211 | 0.89267 | 345.7073 | 210.1073 | 2.7935 | 0.96444 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 11/01
AR28B | 20180410 | L12 | 11 | 7 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 225.2 | 137.4 | 0.89267 | 224.3073 | 136.5073 | 0.90439 | 0.31639 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 11/02
AR28B | 20180410 | L12 | 11 | 8 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 388.9 | 234.5 | 0.89267 | 388.0073 | 233.6073 | 1.5904 | 0.49873 | 20181101 | 9-68 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 11/03
AR28B | 20180410 | L12 | 11 | 9 |  |  | a | 535 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 380.7 | 228.2 | 0.89267 | 379.8073 | 227.3073 | 3.1417 | 0.92392 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 11/04
AR28B | 20180410 | L11 | 12 | 1 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 9.936 | 8.744 | 0.81033 | 9.1257 | 7.9337 | 0.012278 | 0.058672 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication (40% chl range) but very low values
AR28B | 20180410 | L11 | 12 | 1 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 8.583 | 7.794 | 0.81033 | 7.7727 | 6.9837 | 0.0081271 | 0.054327 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication (40% chl range) but very low values
AR28B | 20180410 | L11 | 12 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 10.84 | 9.623 | 0.81033 | 10.0297 | 8.8127 | 0.012536 | 0.066275 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR28B | 20180410 | L11 | 12 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 12.25 | 10.7 | 0.778 | 11.472 | 9.922 | 0.015966 | 0.072766 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR28B | 20180410 | L11 | 12 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.976 | 6.2 | 0.81033 | 6.1657 | 5.3897 | 0.0079932 | 0.040206 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (27% range) but low value
AR28B | 20180410 | L11 | 12 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.152 | 6.133 | 0.81033 | 6.3417 | 5.3227 | 0.010496 | 0.037104 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (27% range) but low value
AR28B | 20180410 | L11 | 12 | 6 |  |  | a | 532 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.87 | 10.7 | 0.778 | 13.092 | 9.922 | 0.032837 | 0.056395 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L11 | 12 | 6 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 12.98 | 10.06 | 0.81033 | 12.1697 | 9.2497 | 0.030078 | 0.052641 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L11 | 12 | 7 |  |  | a | 525 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 125.6 | 80.14 | 0.81033 | 124.7897 | 79.3297 | 0.47718 | 0.24577 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L11 | 12 | 7 |  |  | b | 525 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 122.8 | 78.28 | 0.778 | 122.022 | 77.502 | 0.46731 | 0.23898 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L11 | 12 | 8 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 182.2 | 115.9 | 0.81033 | 181.3897 | 115.0897 | 0.68292 | 0.34631 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L11 | 12 | 8 |  |  | b | 531 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 186.2 | 118.7 | 0.81033 | 185.3897 | 117.8897 | 0.70052 | 0.3617 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L11 | 12 | 10 |  |  | a | 549 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 171.8 | 104.8 | 0.81033 | 170.9897 | 103.9897 | 0.67254 | 0.23372 | 20181101 | 9-67 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR28B | 20180410 | L11 | 12 | 10 |  |  | b | 549 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 182.8 | 116.3 | 0.778 | 182.022 | 115.522 | 0.66752 | 0.33924 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR28B | 20180410 | L11 | 12 | 11 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 128.8 | 82.18 | 0.81033 | 127.9897 | 81.3697 | 0.48021 | 0.24747 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L11 | 12 | 11 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 141.7 | 91.3 | 0.778 | 140.922 | 90.522 | 0.51915 | 0.29038 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L11 | 12 | 14 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 182 | 113.8 | 0.81033 | 181.1897 | 112.9897 | 0.7025 | 0.30796 | 20181101 | 9-67 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L11 | 12 | 14 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 203.8 | 128.9 | 0.81033 | 202.9897 | 128.0897 | 0.77151 | 0.37399 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L11 | 12 | 15 |  |  | a | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 264.3 | 164.7 | 0.778 | 263.522 | 163.922 | 1.0034 | 0.43036 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L11 | 12 | 15 |  |  | b | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 254.1 | 159.2 | 0.778 | 253.322 | 158.422 | 0.95608 | 0.4296 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L11 | 12 | 19 |  |  | a | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 262.9 | 164.8 | 0.81033 | 262.0897 | 163.9897 | 0.98831 | 0.44606 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L11 | 12 | 19 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 222.3 | 139.7 | 0.81033 | 221.4897 | 138.8897 | 0.85082 | 0.39126 | 20181101 | 9-67 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L11 | 12 | 19 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 94.11 | 63.38 | 0.778 | 93.332 | 62.602 | 0.3136 | 0.24106 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication but rep 10b clearly wrong
AR28B | 20180410 | L11 | 12 | 19 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 231.6 | 145 | 0.778 | 230.822 | 144.222 | 0.88377 | 0.39406 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | poor replication, nearly the same as WSW values, was the nitex in place? Was this poured through the nitex?
AR28B | 20180410 | L9 | 17 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 37.27 | 28.15 | 0.81033 | 36.4597 | 27.3397 | 0.093941 | 0.15056 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L9 | 17 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 35.02 | 27.24 | 0.81033 | 34.2097 | 26.4297 | 0.080138 | 0.15622 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L9 | 17 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 39.35 | 30.66 | 0.778 | 38.572 | 29.882 | 0.089512 | 0.17772 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L9 | 17 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 44.12 | 33.98 | 0.778 | 43.342 | 33.202 | 0.10445 | 0.19248 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L9 | 17 | 7 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 53.19 | 39.1 | 0.778 | 52.412 | 38.322 | 0.14513 | 0.19758 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L9 | 17 | 7 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 55.17 | 40.63 | 0.778 | 54.392 | 39.852 | 0.14977 | 0.20662 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L9 | 17 | 8 |  |  | a | 531 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 99.89 | 65.31 | 0.778 | 99.112 | 64.532 | 0.35888 | 0.22258 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L9 | 17 | 8 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 101.5 | 68.3 | 0.778 | 100.722 | 67.522 | 0.34198 | 0.26187 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L9 | 17 | 9 |  |  | a | 525 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 90.58 | 57.73 | 0.89267 | 89.6873 | 56.8373 | 0.34482 | 0.17316 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 17/01
AR28B | 20180410 | L9 | 17 | 10 |  |  | a | 525 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 92.28 | 59.48 | 0.89267 | 91.3873 | 58.5873 | 0.34429 | 0.18963 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 17/02
AR28B | 20180410 | L9 | 17 | 12 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 120.1 | 73.9 | 0.778 | 119.322 | 73.122 | 0.47588 | 0.17804 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L9 | 17 | 12 |  |  | b | 532 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 126.6 | 78.76 | 0.81033 | 125.7897 | 77.9497 | 0.49556 | 0.20547 | 20181101 | 9-65 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L9 | 17 | 13 |  |  | a | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 170.2 | 102.2 | 0.778 | 169.422 | 101.422 | 0.68507 | 0.20204 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L9 | 17 | 13 |  |  | b | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 160.7 | 96.29 | 0.778 | 159.922 | 95.512 | 0.6489 | 0.18652 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L9 | 17 | 14 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 177.5 | 106.7 | 0.89267 | 176.6073 | 105.8073 | 0.72928 | 0.21695 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 17/03
AR28B | 20180410 | L9 | 17 | 15 |  |  | a | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 154 | 93.02 | 0.89267 | 153.1073 | 92.1273 | 0.61435 | 0.19147 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 17/04
AR28B | 20180410 | L9 | 17 | 17 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 180.9 | 107.5 | 0.778 | 180.122 | 106.722 | 0.74906 | 0.19651 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180410 | L9 | 17 | 17 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 180.3 | 107.5 | 0.778 | 179.522 | 106.722 | 0.74293 | 0.20263 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L7 | 27 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 68.76 | 52.16 | 0.778 | 67.982 | 51.382 | 0.17099 | 0.28852 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L7 | 27 | 3 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 73.25 | 56.35 | 0.778 | 72.472 | 55.572 | 0.17408 | 0.3229 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L7 | 27 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 50.31 | 37.09 | 0.778 | 49.532 | 36.312 | 0.13617 | 0.18856 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L7 | 27 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 48.89 | 36.42 | 0.778 | 48.112 | 35.642 | 0.12845 | 0.1903 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L7 | 27 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 82.39 | 58.96 | 0.778 | 81.612 | 58.182 | 0.24134 | 0.27898 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L7 | 27 | 5 |  |  | b | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 81.55 | 57.37 | 0.778 | 80.772 | 56.592 | 0.2436 | 0.25139 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L7 | 27 | 9 |  |  | a | 532 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 73 | 49.21 | 0.778 | 72.222 | 48.432 | 0.24643 | 0.18913 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L7 | 27 | 9 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 81.5 | 56 | 0.778 | 80.722 | 55.222 | 0.26266 | 0.23118 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L7 | 27 | 10 |  |  | a | 525 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 154.3 | 95.89 | 0.89267 | 153.4073 | 94.9973 | 0.61311 | 0.25262 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 27/01
AR28B | 20180411 | L7 | 27 | 11 |  |  | a | 525 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 157.7 | 98.07 | 0.89267 | 156.8073 | 97.1773 | 0.62592 | 0.25968 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 27/02
AR28B | 20180411 | L7 | 27 | 12 |  |  | a | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 157.2 | 96.22 | 0.78917 | 156.4108 | 95.4308 | 0.61435 | 0.22036 | 20181101 | 9-58 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L7 | 27 | 12 |  |  | b | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 163 | 100.9 | 0.78917 | 162.2108 | 100.1108 | 0.62563 | 0.25001 | 20181101 | 9-58 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L7 | 27 | 13 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 168.9 | 105.1 | 0.778 | 168.122 | 104.322 | 0.65717 | 0.27577 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L7 | 27 | 13 |  |  | b | 531 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 161.5 | 101.3 | 0.778 | 160.722 | 100.522 | 0.62476 | 0.28097 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L7 | 27 | 14 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 164.1 | 101 | 0.89267 | 163.2073 | 100.1073 | 0.64996 | 0.24529 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 27/03
AR28B | 20180411 | L7 | 27 | 15 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 168.7 | 101.5 | 0.89267 | 167.8073 | 100.6073 | 0.6922 | 0.20753 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 27/04
AR28B | 20180411 | L7 | 27 | 17 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 169 | 103.7 | 0.78917 | 168.2108 | 102.9108 | 0.6664 | 0.24541 | 20181101 | 9-58 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L7 | 27 | 17 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 163.1 | 98.41 | 0.778 | 162.322 | 97.632 | 0.66017 | 0.20486 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L6 | 28 | 3 |  |  | a | 535 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 81.6 | 61.44 | 0.778 | 80.822 | 60.662 | 0.41532 | 0.66967 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L6 | 28 | 3 |  |  | b | 535 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 84.22 | 63.08 | 0.778 | 83.442 | 62.302 | 0.43551 | 0.67882 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L6 | 28 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 89.81 | 69.7 | 0.778 | 89.032 | 68.922 | 0.20714 | 0.40922 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L6 | 28 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 92.52 | 70.68 | 0.778 | 91.742 | 69.902 | 0.22496 | 0.40016 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L6 | 28 | 6 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 41.95 | 27.76 | 0.778 | 41.172 | 26.982 | 0.14616 | 0.095133 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L6 | 28 | 6 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 39.66 | 26.68 | 0.778 | 38.882 | 25.902 | 0.1337 | 0.097939 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L6 | 28 | 7 |  |  | a | 525 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 123.2 | 73.76 | 0.89267 | 122.3073 | 72.8673 | 0.51896 | 0.1451 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 28/01
AR28B | 20180411 | L6 | 28 | 8 |  |  | a | 525 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 129.3 | 79.31 | 0.89267 | 128.4073 | 78.4173 | 0.52473 | 0.18991 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 28/02
AR28B | 20180411 | L6 | 28 | 10 |  |  | a | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 141.6 | 86.21 | 0.778 | 140.822 | 85.432 | 0.55803 | 0.18922 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L6 | 28 | 10 |  |  | b | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 141.7 | 86.22 | 0.78917 | 140.9108 | 85.4308 | 0.55894 | 0.1883 | 20181101 | 9-58 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L6 | 28 | 11 |  |  | a | 532 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 137.8 | 83.34 | 0.89267 | 136.9073 | 82.4473 | 0.56413 | 0.17735 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 28/03
AR28B | 20180411 | L6 | 28 | 12 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 143.4 | 86.04 | 0.89267 | 142.5073 | 85.1473 | 0.59084 | 0.17063 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 28/04
AR28B | 20180411 | L6 | 28 | 14 |  |  | a | 545 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 124 | 73.57 | 0.89267 | 123.1073 | 72.6773 | 0.50992 | 0.1281 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 28/05
AR28B | 20180411 | L6 | 28 | 14 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 110.5 | 66.99 | 0.778 | 109.722 | 66.212 | 0.44403 | 0.14262 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L6 | 28 | 14 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 100 | 60.84 | 0.778 | 99.222 | 60.062 | 0.39963 | 0.13252 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180411 | L6 | 28 | 15 |  |  | a | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 116.5 | 71.15 | 0.89267 | 115.6073 | 70.2573 | 0.45688 | 0.15764 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 28/06
AR28B | 20180412 | L8 | 37 | 3 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 98.91 | 76.09 | 0.78917 | 98.1208 | 75.3008 | 0.23506 | 0.43835 | 20181101 | 9-58 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180412 | L8 | 37 | 3 |  |  | b | 535 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 50.6 | 39.07 | 0.778 | 49.822 | 38.292 | 0.23753 | 0.44735 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180412 | L8 | 37 | 4 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 68.38 | 52.15 | 0.78917 | 67.5908 | 51.3608 | 0.16718 | 0.29214 | 20181101 | 9-58 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180412 | L8 | 37 | 4 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 63.95 | 48.69 | 0.778 | 63.172 | 47.912 | 0.15719 | 0.27129 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180412 | L8 | 37 | 5 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 67.66 | 47.23 | 0.778 | 66.882 | 46.452 | 0.21044 | 0.20498 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180412 | L8 | 37 | 5 |  |  | b | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 67.66 | 47.23 | 0.778 | 66.882 | 46.452 | 0.21044 | 0.20498 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180412 | L8 | 37 | 6 |  |  | a | 525 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 86.01 | 54.67 | 0.89267 | 185.9073 | 112.3073 | 0.77256 | 0.25093 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 37/01
AR28B | 20180412 | L8 | 37 | 7 |  |  | a | 525 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 85.86 | 53.87 | 0.89267 | 84.9673 | 52.9773 | 0.33579 | 0.14701 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 37/02
AR28B | 20180412 | L8 | 37 | 10 |  |  | a | 532 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 139 | 86.67 | 0.89267 | 138.1073 | 85.7773 | 0.54207 | 0.22936 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 37/03
AR28B | 20180412 | L8 | 37 | 11 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 135.8 | 84.41 | 0.89267 | 134.9073 | 83.5173 | 0.52934 | 0.21754 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 37/04
AR28B | 20180412 | L8 | 37 | 13 |  |  | a | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 156.4 | 96.05 | 0.778 | 155.622 | 95.272 | 0.608 | 0.22532 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180412 | L8 | 37 | 13 |  |  | b | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 157.6 | 95.63 | 0.778 | 156.822 | 94.852 | 0.62432 | 0.20532 | 20181101 | 9-64 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR28B | 20180412 | L8 | 37 | 14 |  |  | a | 535 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 198.5 | 119.9 | 0.89267 | 197.6073 | 119.0073 | 0.80962 | 0.25465 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 37/05
AR28B | 20180412 | L8 | 37 | 14 |  |  | a | 540 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 209.4 | 126.8 | 0.778 | 208.622 | 126.022 | 0.84295 | 0.27362 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | and are just WSWhighly suspect. These were missing values and there was a pair of <10 in the lab notebook that remained unentered. Looks like they weren't filtered through the nitex
AR28B | 20180412 | L8 | 37 | 14 |  |  | b | 540 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 193.3 | 114.9 | 0.778 | 192.522 | 114.122 | 0.80008 | 0.21105 | 20181101 | 9-63 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | highly suspect. These were missing values and there was a pair of <10 in the lab notebook that remained unentered. Looks like just WSW ad was forgotten to go through nitex
AR28B | 20180412 | L8 | 37 | 15 |  |  | a | 547 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 186.8 | 113.2 | 0.89267 | 185.9073 | 112.3073 | 0.74149 | 0.24083 | 20181101 | 9-69 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 37/06
AR31B | 20181025 |  | 1 | 5 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 21.74 | 18.83 | 0.7025 | 21.0375 | 18.1275 | 0.015172 | 0.066882 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 01/01
AR31B | 20181025 |  | 1 | 6 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 15.44 | 13.22 | 0.7025 | 14.7375 | 12.5175 | 0.011585 | 0.045128 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 01/02
AR31B | 20181025 |  | 1 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 226.7 | 140.7 | 0.7025 | 225.9975 | 139.9975 | 2.6928 | 1.113 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 01/03
AR31B | 20181025 |  | 1 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 244.3 | 152 | 0.7025 | 243.5975 | 151.2975 | 2.89 | 1.2229 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 01/04
AR31B | 20181025 |  | 2 | 6 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 23.77 | 19.29 | 0.7025 | 23.0675 | 18.5875 | 0.023357 | 0.060778 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/01
AR31B | 20181025 |  | 2 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 99.49 | 77.98 | 0.7025 | 98.7875 | 77.2775 | 0.11225 | 0.23787 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/02
AR31B | 20181025 |  | 2 | 7 |  |  | b | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 104.1 | 81.95 | 0.7025 | 103.3975 | 81.2475 | 0.11559 | 0.25252 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/03
AR31B | 20181025 |  | 2 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 237 | 146.9 | 0.7025 | 236.2975 | 146.1975 | 2.8211 | 1.1532 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/04
AR31B | 20181025 |  | 2 | 8 |  |  | b | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 283.1 | 175 | 0.7025 | 282.3975 | 174.2975 | 3.3815 | 1.3522 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/05
AR31B | 20181025 |  | 2 | 9 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 236.8 | 148.7 | 0.7025 | 236.0975 | 147.9975 | 2.7585 | 1.2647 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/06
AR31B | 20181026 |  | 3 | 3 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 43.19 | 33.43 | 0.7025 | 42.4875 | 32.7275 | 0.050885 | 0.097255 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 03/01
AR31B | 20181026 |  | 3 | 4 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 369.2 | 242.7 | 0.7025 | 368.4975 | 241.9975 | 0.66015 | 0.43628 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 03/02
AR31B | 20181026 |  | 3 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 208.1 | 129.9 | 0.7025 | 207.3975 | 129.1975 | 2.4485 | 1.0636 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 03/03
AR31B | 20181026 |  | 3 | 5 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 222.6 | 138.7 | 0.7025 | 221.8975 | 137.9975 | 2.627 | 1.1244 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 03/04
AR31B | 20181026 |  | 3 | 6 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 217 | 137.6 | 0.7025 | 216.2975 | 136.8975 | 2.4838 | 1.2342 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 03/05
AR31B | 20181026 |  | 4 | 4 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 142.6 | 89.39 | 0.7025 | 141.8975 | 88.6875 | 1.6645 | 0.74414 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 04/01
AR31B | 20181026 |  | 4 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 182 | 113.5 | 0.7025 | 181.2975 | 112.7975 | 2.1448 | 0.92152 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 04/02
AR31B | 20181026 |  | 4 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 194 | 121 | 0.7025 | 193.2975 | 120.2975 | 2.2857 | 0.9845 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 04/03
AR31B | 20181026 |  | 4 | 9 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 182.4 | 113.9 | 0.7025 | 181.6975 | 113.1975 | 2.1448 | 0.93239 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 04/04
AR31B | 20181029 |  | 5 | 2 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 77.52 | 64.29 | 0.7025 | 76.8175 | 63.5875 | 0.13795 | 0.4377 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 05/01
AR31B | 20181029 |  | 5 | 4 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 228.9 | 142.4 | 0.7025 | 228.1975 | 141.6975 | 2.7084 | 1.1435 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 05/02
AR31B | 20181029 |  | 5 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 265.9 | 164.6 | 0.7025 | 265.1975 | 163.8975 | 3.1718 | 1.2836 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 05/03
AR31B | 20181029 |  | 5 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 271.3 | 168.9 | 0.7025 | 270.5975 | 168.1975 | 3.2063 | 1.3661 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 05/04
AR31B | 20181029 |  | 5 | 7 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 267.7 | 162 | 0.7025 | 266.9975 | 161.2975 | 3.3065 | 1.0742 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 05/05
AR31B | 20181029 |  | 5 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 274.8 | 172.9 | 0.7025 | 274.0975 | 172.1975 | 3.1906 | 1.4905 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 05/06
AR31B | 20181029 |  | 6 | 3 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 46.31 | 36.12 | 0.7025 | 45.6075 | 35.4175 | 0.053127 | 0.10719 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/01
AR31B | 20181029 |  | 6 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 233.7 | 145.9 | 0.7025 | 232.9975 | 145.1975 | 2.7491 | 1.198 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/02
AR31B | 20181029 |  | 6 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 313.7 | 191.2 | 0.7025 | 312.9975 | 190.4975 | 3.8356 | 1.3429 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/03
AR31B | 20181029 |  | 6 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 260.1 | 161.2 | 0.7025 | 259.3975 | 160.4975 | 3.0967 | 1.2664 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/04
AR31B | 20181029 |  | 6 | 8 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 290.2 | 180.2 | 0.7025 | 289.4975 | 179.4975 | 3.441 | 1.4339 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/05
AR31B | 20181030 |  | 7 | 4 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 165.6 | 117.8 | 0.7025 | 164.8975 | 117.0975 | 0.24921 | 0.28083 | 20181101 | 9-82 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 07/01
AR31B | 20181030 |  | 7 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 343.8 | 211 | 0.7025 | 343.0975 | 210.2975 | 4.1581 | 1.5587 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | 07/02
AR31B | 20181030 |  | 7 | 5 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 280.2 | 173.2 | 0.7025 | 279.4975 | 172.4975 | 3.3503 | 1.3389 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | 07/03
AR31B | 20181030 |  | 7 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 358.5 | 218.1 | 0.7025 | 357.7975 | 217.3975 | 4.3961 | 1.5137 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 07/04
AR31B | 20181030 |  | 7 | 6 |  |  | b | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 339.4 | 206.6 | 0.7025 | 338.6975 | 205.8975 | 4.1542 | 1.4377 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 07/05
AR31B | 20181031 |  | 8 | 3 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 175.9 | 129.7 | 0.7025 | 175.1975 | 128.9975 | 0.24087 | 0.34303 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 08/01
AR31B | 20181031 |  | 8 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 331.6 | 203.7 | 0.7025 | 330.8975 | 202.9975 | 4.0047 | 1.5137 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 08/02
AR31B | 20181031 |  | 8 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 348 | 214.1 | 0.7025 | 347.2975 | 213.3975 | 4.1926 | 1.6085 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 08/03
AR31B | 20181031 |  | 8 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 341.6 | 209.6 | 0.7025 | 340.8975 | 208.8975 | 4.1331 | 1.5457 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 08/04
AR31B | 20181031 |  | 8 | 8 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 348.2 | 217.6 | 0.7025 | 347.4975 | 216.8975 | 4.0854 | 1.8053 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 08/05
AR31C | 20181102 |  | 1 | 3 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 60.99 | 48.89 | 0.7025 | 60.2875 | 48.1875 | 0.063085 | 0.15503 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 01/01
AR31C | 20181102 |  | 1 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 347.2 | 215.6 | 0.7025 | 346.4975 | 214.8975 | 4.1206 | 1.7213 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 01/02
AR31C | 20181102 |  | 1 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 314.6 | 198.5 | 0.7025 | 313.8975 | 197.7975 | 3.6352 | 1.7418 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 01/03
AR31C | 20181102 |  | 1 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 253 | 158.8 | 0.7025 | 252.2975 | 158.0975 | 2.9495 | 1.3483 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 01/04
AR31C | 20181102 |  | 1 | 8 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 321.3 | 196.3 | 0.7025 | 320.5975 | 195.5975 | 3.9102 | 1.402 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 01/05
AR31C | 20181102 |  | 2 | 4 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 59.08 | 48.11 | 0.7025 | 58.3775 | 47.4075 | 0.057193 | 0.15739 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/01
AR31C | 20181102 |  | 2 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 296.7 | 185.1 | 0.7025 | 295.9975 | 184.3975 | 3.4943 | 1.5184 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/02
AR31C | 20181102 |  | 2 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 280.9 | 172.8 | 0.7025 | 280.1975 | 172.0975 | 3.3847 | 1.2936 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/03
AR31C | 20181102 |  | 2 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 295.9 | 180.5 | 0.7025 | 295.1975 | 179.7975 | 3.6133 | 1.2744 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/04
AR31C | 20181102 |  | 2 | 8 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 268.5 | 165.9 | 0.7025 | 267.7975 | 165.1975 | 3.2095 | 1.2771 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 02/05
AR31C | 20181104 |  | 4 | 4 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 85.39 | 54.41 | 0.7025 | 84.6875 | 53.7075 | 0.96911 | 0.48952 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 04/01
AR31C | 20181104 |  | 4 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 374.9 | 225.7 | 0.7025 | 374.1975 | 224.9975 | 4.6716 | 1.4448 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 04/02
AR31C | 20181104 |  | 4 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 320.2 | 196.4 | 0.7025 | 319.4975 | 195.6975 | 3.8763 | 1.4436 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 04/03
AR31C | 20181104 |  | 4 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 337.3 | 208 | 0.7025 | 336.5975 | 207.2975 | 4.0485 | 1.5867 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 04/04
AR31C | 20181104 |  | 4 | 8 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 390.8 | 237.8 | 0.7025 | 390.0975 | 237.0975 | 4.7861 | 1.6532 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 04/05
AR31C | 20181105 |  | 11 | 4 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 111.8 | 88.71 | 0.7025 | 111.0975 | 88.0075 | 0.12038 | 0.27798 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 11/01
AR31C | 20181105 |  | 11 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 335.1 | 193.6 | 0.7025 | 334.3975 | 192.8975 | 4.4305 | 0.81327 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 11/02
AR31C | 20181105 |  | 11 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 302.4 | 183 | 0.7025 | 301.6975 | 182.2975 | 3.7386 | 1.2171 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 11/03
AR31C | 20181105 |  | 11 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 293 | 183 | 0.7025 | 292.2975 | 182.2975 | 3.4442 | 1.5114 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 11/04
AR31C | 20181105 |  | 11 | 8 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 351.5 | 216.2 | 0.7025 | 350.7975 | 215.4975 | 4.2324 | 1.6202 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 11/05
AR31C | 20181106 |  | 12 | 6 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 22.77 | 18.94 | 0.7025 | 22.0675 | 18.2375 | 0.019968 | 0.062583 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 12/01
AR31C | 20181106 |  | 12 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 293.4 | 180.7 | 0.7025 | 292.6975 | 179.9975 | 3.5288 | 1.3644 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 12/02
AR31C | 20181106 |  | 12 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 314.4 | 191.7 | 0.7025 | 313.6975 | 190.9975 | 3.8419 | 1.3503 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 12/03
AR31C | 20181107 |  | 20 | 7 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 219.9 | 138.2 | 0.7025 | 219.1975 | 137.4975 | 2.5557 | 1.1785 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 20/01
AR31C | 20181107 |  | 20 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 224.2 | 139.5 | 0.7025 | 223.4975 | 138.7975 | 2.6521 | 1.1211 | 20181101 | 9-83 |  | 4 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 20/02
AR34A | 20190405 |  | 5 | 7 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 147 | 90.26 | 0.81233 | 146.1877 | 89.4477 | 1.7749 | 0.65436 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 05/01
AR34A | 20190405 |  | 5 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 162.9 | 98.95 | 0.81233 | 162.0877 | 98.1377 | 2.0024 | 0.66546 | 20181101 | 9-94 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 05/02
AR34A | 20190405 |  | 6 | 5 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 289.6 | 189.7 | 0.81233 | 288.7877 | 188.8877 | 0.52084 | 0.33415 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/01
AR34A | 20190405 |  | 6 | 6 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 303 | 197.6 | 0.81233 | 302.1877 | 196.7877 | 0.55003 | 0.34156 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/02
AR34A | 20190405 |  | 6 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 161.4 | 101.8 | 0.81233 | 160.5877 | 100.9877 | 1.8662 | 0.87914 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/03
AR34A | 20190405 |  | 6 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 169.9 | 105.7 | 0.81233 | 169.0877 | 104.8877 | 2.0102 | 0.84113 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 06/04
AR34A | 20190406 |  | 9 | 5 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 245.6 | 155.2 | 0.81233 | 244.7877 | 154.3877 | 0.94262 | 0.45504 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 09/01
AR34A | 20190406 |  | 9 | 6 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 168 | 109.5 | 0.81233 | 167.1877 | 108.6877 | 0.91585 | 0.56145 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 09/02
AR34A | 20190406 |  | 9 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 202.6 | 123.8 | 0.81233 | 201.7877 | 122.9877 | 2.4673 | 0.87602 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 09/03
AR34A | 20190406 |  | 9 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 202.1 | 125.4 | 0.81233 | 201.2877 | 124.5877 | 2.4016 | 0.98527 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 09/04
AR34A | 20190407 |  | 10 | 3 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 364.3 | 238.2 | 0.81233 | 363.4877 | 237.3877 | 0.65744 | 0.41709 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 10/01
AR34A | 20190407 |  | 10 | 4 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 352.8 | 229.8 | 0.81233 | 351.9877 | 228.9877 | 0.64188 | 0.3956 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 10/02
AR34A | 20190407 |  | 10 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 337 | 206.1 | 0.81233 | 336.1877 | 205.2877 | 1.3662 | 0.494 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 10/03
AR34A | 20190407 |  | 10 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 338.4 | 207.6 | 0.81233 | 337.5877 | 206.7877 | 1.3652 | 0.50863 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 10/04
AR34A | 20190407 |  | 10 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 203.7 | 126.2 | 0.81233 | 202.8877 | 125.3877 | 2.4266 | 0.98197 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 10/05
AR34A | 20190407 |  | 10 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 211.5 | 129.3 | 0.81233 | 210.6877 | 128.4877 | 2.5738 | 0.91908 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 10/06
AR34A | 20190407 |  | 11 | 5 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 365.5 | 229.8 | 0.81233 | 364.6877 | 228.9877 | 1.415 | 0.65803 | 20181101 | 9-94 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 11/01
AR34A | 20190407 |  | 11 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 188.5 | 114.6 | 0.81233 | 187.6877 | 113.7877 | 2.3139 | 0.77935 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 11/02
AR34A | 20190407 |  | 12 | 3 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 272.9 | 180.1 | 0.81233 | 272.0877 | 179.2877 | 0.48382 | 0.32771 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 12/01
AR34A | 20190407 |  | 12 | 4 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 276.8 | 182.4 | 0.81233 | 275.9877 | 181.5877 | 0.49263 | 0.3301 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 12/02
AR34A | 20190407 |  | 12 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 144.1 | 89.63 | 0.81233 | 143.2877 | 88.8177 | 1.7055 | 0.70893 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 12/03
AR34A | 20190407 |  | 12 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 138.7 | 86.32 | 0.81233 | 137.8877 | 85.5077 | 1.6401 | 0.68439 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 12/04
AR34A | 20190407 |  | 12 | 7 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 162.1 | 100.1 | 0.81233 | 161.2877 | 99.2877 | 1.9395 | 0.75706 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 12/05
AR34A | 20190407 |  | 12 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 169.3 | 104.5 | 0.81233 | 168.4877 | 103.6877 | 2.029 | 0.78972 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 12/06
AR34A | 20190408 |  | 13 | 5 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 42.48 | 31.72 | 0.81233 | 41.6677 | 30.9077 | 0.056098 | 0.083804 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 13/01
AR34A | 20190408 |  | 13 | 6 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 308.1 | 201.3 | 0.81233 | 307.2877 | 200.4877 | 0.55734 | 0.35102 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 13/02
AR34A | 20190408 |  | 13 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 273.2 | 168.6 | 0.81233 | 272.3877 | 167.7877 | 3.2752 | 1.2861 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 13/03
AR34A | 20190408 |  | 13 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 196.3 | 118.5 | 0.81233 | 195.4877 | 117.6877 | 2.436 | 0.76326 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 13/04
AR34A | 20190409 |  | 17 | 4 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 164.3 | 111.5 | 0.81233 | 163.4877 | 110.6877 | 0.27528 | 0.22574 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 17/01
AR34A | 20190409 |  | 17 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 183.2 | 115.2 | 0.81233 | 182.3877 | 114.3877 | 2.1292 | 0.9804 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 17/02
AR34A | 20190409 |  | 17 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 170.8 | 105.5 | 0.81233 | 169.9877 | 104.6877 | 2.0446 | 0.80125 | 20181101 | 9-92 |  | 2 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | 17/03
AR34B | 20190415 | L13 | 1 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 47.56 | 33.72 | 0.8215 | 46.7385 | 32.8985 | 0.263 | 0.27977 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | LTER sampled/filtered chl for OOI
AR34B | 20190415 | L13 | 1 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 40.98 | 29.17 | 0.8215 | 40.1585 | 28.3485 | 0.22442 | 0.24328 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L13 | 1 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 89.09 | 60.88 | 0.8215 | 88.2685 | 60.0585 | 0.53607 | 0.45479 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L13 | 1 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 85.12 | 60.46 | 0.8215 | 84.2985 | 59.6385 | 0.46861 | 0.51532 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L13 | 1 | 7 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 189.7 | 115.9 | 0.8215 | 188.8785 | 115.0785 | 2.8048 | 0.99236 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L13 | 1 | 8 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 203.6 | 126.4 | 0.8215 | 202.7785 | 125.5785 | 2.934 | 1.2096 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L4 | 2 | 2 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 268 | 168.1 | 0.8215 | 267.1785 | 167.2785 | 3.7967 | 1.7228 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L4 | 2 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 264.1 | 160.7 | 0.8215 | 263.2785 | 159.8785 | 3.9298 | 1.3456 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L4 | 2 | 3 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 274.3 | 163.2 | 0.9095 | 273.3905 | 162.2905 | 4.2224 | 1.1326 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | N3chla might be bad but value seems ok, filter close to edge of funnel/frit = might have leaked some?
AR34B | 20190415 | L4 | 2 | 3 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 273.9 | 168 | 0.8215 | 273.0785 | 167.1785 | 4.0248 | 1.4915 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L4 | 2 | 4 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 311.2 | 191.5 | 0.8215 | 310.3785 | 190.6785 | 2.3145 | 0.8865 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L4 | 2 | 4 |  |  | b | 285 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 170.1 | 105.9 | 0.8215 | 169.2785 | 105.0785 | 2.4828 | 1.0453 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L4 | 2 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 98.39 | 67.46 | 0.8215 | 97.5685 | 66.6385 | 0.58775 | 0.51166 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L4 | 2 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 106.7 | 71.06 | 0.8215 | 105.8785 | 70.2385 | 0.67726 | 0.48155 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L4 | 2 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 55.5 | 34.82 | 0.8215 | 54.6785 | 33.9985 | 0.39298 | 0.16794 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L4 | 2 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 59.89 | 37.77 | 0.9095 | 58.9805 | 36.8605 | 0.42034 | 0.18779 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L4 | 2 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.17 | 12.54 | 0.8215 | 18.3485 | 11.7185 | 0.12599 | 0.067346 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (33% range)
AR34B | 20190415 | L4 | 2 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 26.48 | 17.15 | 0.8215 | 25.6585 | 16.3285 | 0.1773 | 0.092095 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication (33% range)
AR34B | 20190415 | L3 | 3 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 22.94 | 17.98 | 0.8215 | 22.1185 | 17.1585 | 0.094253 | 0.18883 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (24% range) but low values
AR34B | 20190415 | L3 | 3 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 20.72 | 16.84 | 0.8215 | 19.8985 | 16.0185 | 0.07373 | 0.19054 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor replication (24% range) but low values
AR34B | 20190415 | L3 | 3 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 51.11 | 32.36 | 0.8215 | 50.2885 | 31.5385 | 0.3563 | 0.16403 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L3 | 3 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 57.43 | 36.2 | 0.8215 | 56.6085 | 35.3785 | 0.40343 | 0.18025 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L3 | 3 | 5 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 74.3 | 44.96 | 0.8215 | 73.4785 | 44.1385 | 0.56732 | 0.17366 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L3 | 3 | 5 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 73.27 | 45.3 | 0.8215 | 72.4485 | 44.4785 | 0.54083 | 0.20586 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L3 | 3 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 74.3 | 45.34 | 0.8215 | 73.4785 | 44.5185 | 0.55032 | 0.18415 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L3 | 3 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 77.2 | 47.87 | 0.8215 | 76.3785 | 47.0485 | 0.55735 | 0.21886 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L3 | 3 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 26.68 | 16.9 | 0.8215 | 25.8585 | 16.0785 | 0.18585 | 0.079419 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L3 | 3 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 30.74 | 19.25 | 0.8215 | 29.9185 | 18.4285 | 0.21834 | 0.085695 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L2 | 4 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 56.66 | 38.3 | 0.8215 | 55.8385 | 37.4785 | 0.34889 | 0.26944 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L2 | 4 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 56.58 | 38.5 | 0.9095 | 55.6705 | 37.5905 | 0.34357 | 0.2766 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L2 | 4 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 79.95 | 51.01 | 0.8215 | 79.1285 | 50.1885 | 0.54994 | 0.27808 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L2 | 4 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 83.05 | 53.29 | 0.8215 | 82.2285 | 52.4685 | 0.56552 | 0.30011 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L2 | 4 | 5 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 89.84 | 56.81 | 0.8215 | 89.0185 | 55.9885 | 0.63867 | 0.30124 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | estimated small amount of volume missing from each replicate, slightly poor replication (21% range)
AR34B | 20190415 | L2 | 4 | 5 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 73.39 | 46.65 | 0.8215 | 72.5685 | 45.8285 | 0.51705 | 0.2523 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | estimated small amount of volume missing from each replicate, slightly poor replication (21% range)
AR34B | 20190415 | L2 | 4 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 84.54 | 51.73 | 0.81233 | 83.7277 | 50.9177 | 0.62348 | 0.21657 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L2 | 4 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 79.08 | 48.44 | 0.8215 | 78.2585 | 47.6185 | 0.58224 | 0.20337 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L2 | 4 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.64 | 12.83 | 0.8215 | 18.8185 | 12.0085 | 0.12941 | 0.06871 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L2 | 4 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 20.7 | 13.01 | 0.8215 | 19.8785 | 12.1885 | 0.14613 | 0.054957 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L1 | 5 | 2 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 67.83 | 44.04 | 0.8215 | 67.0085 | 43.2185 | 0.46001 | 0.26553 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L1 | 5 | 2 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 68.48 | 44.54 | 0.8215 | 67.6585 | 43.7185 | 0.46291 | 0.27102 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L1 | 5 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 68.61 | 44.69 | 0.8215 | 67.7885 | 43.8685 | 0.45454 | 0.2692 | 20181101 | 9-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor chl replication (26% range)
AR34B | 20190415 | L1 | 5 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 86.23 | 55.07 | 0.9095 | 85.3205 | 54.1605 | 0.59212 | 0.30142 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor chl replication (26% range)
AR34B | 20190415 | L1 | 5 | 4 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 33.06 | 21.56 | 0.81233 | 32.2477 | 20.7477 | 0.21853 | 0.12377 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190415 | L1 | 5 | 4 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 31.09 | 20.58 | 0.8215 | 30.2685 | 19.7585 | 0.19972 | 0.12626 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L7 | 6 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 15.8 | 13.06 | 0.8215 | 14.9785 | 12.2385 | 0.052067 | 0.14985 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | slightly poor chl replication (22% range)
AR34B | 20190417 | L7 | 6 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 16.32 | 14.13 | 0.8215 | 15.4985 | 13.3085 | 0.041616 | 0.17795 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | slightly poor chl replication (22% range)
AR34B | 20190417 | L7 | 6 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 18 | 14.76 | 0.9095 | 17.0905 | 13.8505 | 0.061569 | 0.16694 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | slightly poor chl replication (22% range)
AR34B | 20190417 | L7 | 6 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 14.62 | 12.03 | 0.9095 | 13.7105 | 11.1205 | 0.049217 | 0.13425 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | slightly poor chl replication (22% range)
AR34B | 20190417 | L7 | 6 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 55.31 | 37.85 | 0.9095 | 54.4005 | 36.9405 | 0.33179 | 0.27766 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L7 | 6 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 55.67 | 38.72 | 0.9095 | 54.7605 | 37.8105 | 0.3221 | 0.30171 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L7 | 6 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.21 | 15.43 | 0.81233 | 18.3977 | 14.6177 | 0.07183 | 0.16933 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L7 | 6 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 20.65 | 16.46 | 0.8215 | 19.8285 | 15.6385 | 0.079621 | 0.17838 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L7 | 6 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 22.92 | 17.37 | 0.8215 | 22.0985 | 16.5485 | 0.10546 | 0.16755 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L7 | 6 | 9 |  |  | b | 280 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 25.02 | 18.58 | 0.9095 | 24.1105 | 17.6705 | 0.12675 | 0.17519 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 0 | N9chlb spilt ~10ml
AR34B | 20190417 | L7 | 6 | 12 |  |  | a | 285 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 77.46 | 47.12 | 0.8215 | 76.6385 | 46.2985 | 1.1733 | 0.38117 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L7 | 6 | 12 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 151 | 94.86 | 0.8215 | 150.1785 | 94.0385 | 1.0855 | 0.49315 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L7 | 6 | 15 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 109.8 | 67.14 | 0.8215 | 108.9785 | 66.3185 | 1.6213 | 0.56695 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L7 | 6 | 15 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 103.2 | 63 | 0.8215 | 102.3785 | 62.1785 | 1.5278 | 0.52384 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L7 | 6 | 15 |  |  | a | 290 | 10 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 90.67 | 56.37 | 0.8215 | 89.8485 | 55.5485 | 1.3036 | 0.52931 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L7 | 6 | 15 |  |  | b | 290 | 10 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 92.79 | 57.21 | 0.8215 | 91.9685 | 56.3885 | 1.3522 | 0.50838 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L5 | 7 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 319.3 | 208.4 | 0.8215 | 318.4785 | 207.5785 | 2.1074 | 1.3173 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L5 | 7 | 1 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 171.5 | 109.6 | 0.9095 | 170.5905 | 108.6905 | 2.3525 | 1.2338 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L5 | 7 | 3 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 202.7 | 137.3 | 0.8215 | 201.8785 | 136.4785 | 2.4856 | 2.0177 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L5 | 7 | 3 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 212 | 144.2 | 0.9095 | 211.0905 | 143.2905 | 2.5768 | 2.1513 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L5 | 7 | 5 |  |  | a | 285 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 302.9 | 187.8 | 0.9095 | 301.9905 | 186.8905 | 4.4512 | 1.8237 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L5 | 7 | 5 |  |  | b | 285 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 293.7 | 190 | 0.9095 | 292.7905 | 189.0905 | 4.0103 | 2.3384 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L5 | 7 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 145.3 | 95.21 | 0.9095 | 144.3905 | 94.3005 | 0.95184 | 0.60394 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L5 | 7 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 147.5 | 95.85 | 0.9095 | 146.5905 | 94.9405 | 0.98149 | 0.58485 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L5 | 7 | 8 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 71.34 | 47.4 | 0.9095 | 70.4305 | 46.4905 | 0.90985 | 0.62416 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L5 | 7 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 135.5 | 88.87 | 0.9095 | 134.5905 | 87.9605 | 0.8861 | 0.56509 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L5 | 7 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 36.63 | 23.81 | 0.81233 | 35.8177 | 22.9977 | 0.24361 | 0.1358 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L5 | 7 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 33.16 | 21.59 | 0.81233 | 32.3477 | 20.7777 | 0.21986 | 0.12293 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L6 | 8 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 46.69 | 35.05 | 0.9095 | 45.7805 | 34.1405 | 0.22119 | 0.34206 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | somewhat poor chl replication (30% range)
AR34B | 20190417 | L6 | 8 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 33.21 | 24.69 | 0.8215 | 32.3885 | 23.8685 | 0.1619 | 0.23188 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | somewhat poor chl replication (30% range)
AR34B | 20190417 | L6 | 8 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 47.81 | 33.76 | 0.81233 | 46.9977 | 32.9477 | 0.26699 | 0.27659 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L6 | 8 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 43.37 | 29.97 | 0.9095 | 42.4605 | 29.0605 | 0.25464 | 0.22481 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L6 | 8 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 66.07 | 45.85 | 0.9095 | 65.1605 | 44.9405 | 0.38423 | 0.3572 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor chl replication (22%)
AR34B | 20190417 | L6 | 8 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 79.53 | 54.27 | 0.9095 | 78.6205 | 53.3605 | 0.48001 | 0.40034 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor chl replication (22%)
AR34B | 20190417 | L6 | 8 | 5 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 132.5 | 86.46 | 0.8215 | 131.6785 | 85.6385 | 0.89023 | 0.54743 | 20181101 | 9-97 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L6 | 8 | 5 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 140.5 | 90.77 | 0.81233 | 139.6877 | 89.9577 | 0.96158 | 0.54859 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L6 | 8 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 122.8 | 81.47 | 0.81233 | 121.9877 | 80.6577 | 0.78538 | 0.54532 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L6 | 8 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 117.1 | 79.8 | 0.9095 | 116.1905 | 78.8905 | 0.7088 | 0.59275 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L6 | 8 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 127.9 | 85.48 | 0.81233 | 127.0877 | 84.6677 | 0.80609 | 0.59076 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR34B | 20190417 | L6 | 8 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 104.6 | 68 | 0.9095 | 103.6905 | 67.0905 | 0.6955 | 0.41137 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR34B | 20190417 | L6 | 8 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 44.12 | 28.55 | 0.9095 | 43.2105 | 27.6405 | 0.29587 | 0.16014 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190417 | L6 | 8 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 45.2 | 29.96 | 0.9095 | 44.2905 | 29.0505 | 0.2896 | 0.18968 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 0 | N8<10b found copepod on filter, not a good sign, but signal low and great match to <10a??
AR34B | 20190417 | L8 | 9 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 9.415 | 7.854 | 0.81233 | 8.6027 | 7.0417 | 0.029663 | 0.086511 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication (27% range) but phaeo match and just very low values in deep water
AR34B | 20190418 | L8 | 9 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 10.37 | 8.317 | 0.81233 | 9.5577 | 7.5047 | 0.039013 | 0.084801 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication (27% range) but phaeo match and just very low values in deep water
AR34B | 20190418 | L8 | 9 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 10.77 | 8.537 | 0.81233 | 9.9577 | 7.7247 | 0.042433 | 0.08501 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR34B | 20190418 | L8 | 9 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 11.48 | 9.136 | 0.81233 | 10.6677 | 8.3237 | 0.044542 | 0.092783 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR34B | 20190418 | L8 | 9 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 15.54 | 12.53 | 0.81233 | 14.7277 | 11.7177 | 0.057198 | 0.13612 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor chl replication (42% range) but very low values
AR34B | 20190418 | L8 | 9 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 10.89 | 8.935 | 0.9095 | 9.9805 | 8.0255 | 0.03715 | 0.095256 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor chl replication (42% range) but very low values
AR34B | 20190418 | L8 | 9 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 57.28 | 41.54 | 0.9095 | 56.3705 | 40.6305 | 0.2991 | 0.37123 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L8 | 9 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 57.17 | 40.94 | 0.81233 | 56.3577 | 40.1277 | 0.30841 | 0.35362 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | originally was written as C9N4b in lab NB but deduced it's N4
AR34B | 20190418 | L8 | 9 | 6 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 79.56 | 53.98 | 0.9095 | 78.6505 | 53.0705 | 0.49462 | 0.39631 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L8 | 9 | 6 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 77.17 | 50.77 | 0.81233 | 76.3577 | 49.9577 | 0.51047 | 0.3282 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L8 | 9 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 43.29 | 27.6 | 0.81233 | 42.4777 | 26.7877 | 0.29815 | 0.14379 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L8 | 9 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 44.84 | 28.9 | 0.81233 | 44.0277 | 28.0877 | 0.3029 | 0.16049 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L8 | 9 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 33 | 20.23 | 0.81233 | 32.1877 | 19.4177 | 0.24266 | 0.077691 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR34B | 20190418 | L8 | 9 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 31.63 | 19.57 | 0.9095 | 30.7205 | 18.6605 | 0.22917 | 0.078691 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR34B | 20190418 | L8 | 9 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 21 | 12.74 | 0.9095 | 20.0905 | 11.8305 | 0.15696 | 0.038219 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication but other replicate highly suspect
AR34B | 20190418 | L8 | 9 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 33.22 | 20.37 | 0.9095 | 32.3105 | 19.4605 | 0.24418 | 0.076877 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | poor replicatio (43% range) and thi value = WSW values so toss
AR34B | 20190418 | L11 | 10 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 11.03 | 8.668 | 0.81233 | 10.2177 | 7.8557 | 0.044884 | 0.08472 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L11 | 10 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 11.17 | 8.754 | 0.81233 | 10.3577 | 7.9417 | 0.045911 | 0.085112 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L11 | 10 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 15.32 | 12.33 | 0.81233 | 14.5077 | 11.5177 | 0.056818 | 0.1332 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L11 | 10 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 17.18 | 13.93 | 0.81233 | 16.3677 | 13.1177 | 0.061759 | 0.15466 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L11 | 10 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 31.34 | 22.99 | 0.81233 | 30.5277 | 22.1777 | 0.15867 | 0.20722 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L11 | 10 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 30.3 | 22.94 | 0.81233 | 29.4877 | 22.1277 | 0.13986 | 0.22521 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L11 | 10 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 343.7 | 213 | 0.9095 | 342.7905 | 212.0905 | 2.4837 | 1.0154 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L11 | 10 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 387.7 | 239.2 | 0.81233 | 386.8877 | 238.3877 | 2.8219 | 1.1111 | 20181101 | 9-92 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L11 | 10 | 12 |  |  | a | 285 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 207.7 | 128.5 | 0.81233 | 206.8877 | 127.6877 | 3.0628 | 1.2243 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L11 | 10 | 12 |  |  | b | 280 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 392.7 | 239.4 | 0.81233 | 391.8877 | 238.5877 | 3.0172 | 1.0597 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 0 | N12chlb spill ~5ml
AR34B | 20190418 | L11 | 10 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 157.3 | 96.48 | 0.81233 | 156.4877 | 95.6677 | 3.4672 | 1.2678 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | originally written as C9 in lab NB but deduced it should be C10
AR34B | 20190418 | L11 | 10 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_110119 | 1.8682 | 0.5122 | 155.4 | 97.46 | 0.81233 | 154.5877 | 96.6477 | 3.303 | 1.4805 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | originally written as C9 in lab NB but deduced it should be C10
AR34B | 20190418 | L11 | 10 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 107.7 | 67.91 | 0.9095 | 106.7905 | 67.0005 | 0.75612 | 0.34927 | 20181101 | 9-95 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190418 | L11 | 10 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 111.4 | 71.35 | 0.81233 | 110.5877 | 70.5377 | 0.76106 | 0.40268 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L9 | 18 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.471 | 6.546 | 0.81233 | 6.6587 | 5.7337 | 0.017577 | 0.077017 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication (35% range) but very low values
AR34B | 20190422 | L9 | 18 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 7.702 | 6.375 | 0.81233 | 6.8897 | 5.5627 | 0.025217 | 0.066557 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication (35% range) but very low values
AR34B | 20190422 | L9 | 18 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 11.21 | 9.817 | 0.81233 | 10.3977 | 9.0047 | 0.026471 | 0.12209 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | phaeo:chl ratio very different between replicates
AR34B | 20190422 | L9 | 18 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 10.25 | 8.721 | 0.81233 | 9.4377 | 7.9087 | 0.029055 | 0.10142 | 20181101 | 9-92 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | phaeo:chl ratio very different between replicates
AR34B | 20190422 | L9 | 18 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.89 | 11.61 | 0.81233 | 13.0777 | 10.7977 | 0.043326 | 0.13482 | 20181101 | 9-94 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L9 | 18 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 15.56 | 12.86 | 0.81233 | 14.7477 | 12.0477 | 0.051307 | 0.14746 | 20181101 | 9-94 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L9 | 18 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 22.8 | 18.91 | 0.81233 | 21.9877 | 18.0977 | 0.07392 | 0.22466 | 20181101 | 9-92 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR34B | 20190422 | L9 | 18 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 23.41 | 19 | 0.81233 | 22.5977 | 18.1877 | 0.083802 | 0.21626 | 20181101 | 9-92 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR34B | 20190422 | L9 | 18 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 25.03 | 20.05 | 0.72767 | 24.3023 | 19.3223 | 0.10273 | 0.22983 | 20201113 | 10-97 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | dropped N12chla into dewar; note during extraction filter very ripped but all of it present
AR34B | 20190422 | L9 | 18 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 22.88 | 18.58 | 0.81233 | 22.0677 | 17.7677 | 0.081712 | 0.21142 | 20181101 | 9-94 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L9 | 18 | 14 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 212.6 | 133.2 | 0.81233 | 211.7877 | 132.3877 | 1.5353 | 0.68718 | 20181101 | 9-94 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L9 | 18 | 14 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 250.3 | 154.6 | 0.81233 | 249.4877 | 153.7877 | 1.8505 | 0.73126 | 20181101 | 9-94 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L9 | 18 | 18 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 204.3 | 121.8 | 0.81233 | 203.4877 | 120.9877 | 3.1354 | 0.8567 | 20181101 | 9-94 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L9 | 18 | 18 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 224.3 | 134.7 | 0.81233 | 223.4877 | 133.8877 | 3.4053 | 1.0125 | 20181101 | 9-94 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L9 | 18 | 18 |  |  | a | 290 | 10 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 204.1 | 121.1 | 0.81233 | 203.2877 | 120.2877 | 3.1544 | 0.8146 | 20181101 | 9-94 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 0 | poor replication, and this sample is closer to WSW values than the other ,10 could it have not gotten filtered correctly? Or leaky filter funnel nitex?
AR34B | 20190422 | L9 | 18 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 350.1 | 209.1 | 0.81233 | 349.2877 | 208.2877 | 2.6794 | 0.75698 | 20181101 | 9-94 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 0
AR34B | 20190422 | L10 | 21 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 8.064 | 7.05 | 0.81233 | 7.2517 | 6.2377 | 0.019269 | 0.083641 | 20181101 | 9-92 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication (30% range) but very low values
AR34B | 20190422 | L10 | 21 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 8.317 | 6.932 | 0.81233 | 7.5047 | 6.1197 | 0.026319 | 0.074644 | 20181101 | 9-92 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication (30% range) but very low values
AR34B | 20190422 | L10 | 21 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 25.2 | 18.75 | 0.81233 | 24.3877 | 17.9377 | 0.12257 | 0.17337 | 20181101 | 9-94 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L10 | 21 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 25.46 | 19.03 | 0.81233 | 24.6477 | 18.2177 | 0.12219 | 0.17837 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L10 | 21 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 59.65 | 40.94 | 0.81233 | 58.8377 | 40.1277 | 0.35554 | 0.30649 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L10 | 21 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 65.02 | 44.34 | 0.81233 | 64.2077 | 43.5277 | 0.39298 | 0.32515 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L10 | 21 | 16 |  |  | a | 285 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 185.6 | 112.9 | 0.81233 | 184.7877 | 112.0877 | 2.8115 | 0.9519 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L10 | 21 | 16 |  |  | b | 285 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 188.5 | 113.7 | 0.81233 | 187.6877 | 112.8877 | 2.8927 | 0.89754 | 20181101 | 9-92 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | pick 1 copepod off GFF chlb, another too small to pick off
AR34B | 20190422 | L10 | 21 | 20 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 223.1 | 136.8 | 0.81233 | 222.2877 | 135.9877 | 3.2799 | 1.2072 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | both C21N20 written in labNB as rep a so assume one is rep b
AR34B | 20190422 | L10 | 21 | 20 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 246 | 150.9 | 0.81233 | 245.1877 | 150.0877 | 3.6143 | 1.338 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | both C21N20 written in labNB as rep a so assume one is rep b
AR34B | 20190422 | L10 | 21 | 20 |  |  | a | 290 | 10 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 160.2 | 97.88 | 0.81233 | 159.3877 | 97.0677 | 2.3685 | 0.83438 | 20181101 | 9-92 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L10 | 21 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 283.4 | 178.7 | 0.81233 | 282.5877 | 177.8877 | 1.9896 | 0.94523 | 20181101 | 9-94 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR34B | 20190422 | L12 | 22 | 7 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 228.4 | 138.7 | 0.81233 | 227.5877 | 137.8877 | 3.4091 | 1.1407 | 20181101 | 9-93 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | Error in CTD btl file, N1 misfired and did not get recorded in data. All btl values are 1 Niskin off, these 2 samples are from 7m depth and diff niskins. Taylor manually edited btl file so the Niskin numbers would be correct.
AR34B | 20190422 | L12 | 22 | 8 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 242.8 | 144.4 | 0.81233 | 241.9877 | 143.5877 | 3.7397 | 0.99813 | 20181101 | 9-94 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
EN644 | 20190820 | L1 | 1 | 18 |  |  | a | 280.3 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 29.19 | 18.02 | 0.79233 | 28.3977 | 17.2277 | 0.21961 | 0.074455 | 20181101 | 10-24 |  | 1 | MJM/ELJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 1
EN644 | 20190820 | L1 | 1 | 18 |  |  | b | 280.5 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 32.89 | 20 | 0.79233 | 32.0977 | 19.2077 | 0.25324 | 0.074383 | 20181101 | 10-24 |  | 1 | MJM/ELJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 2
EN644 | 20190820 | L1 | 1 | 8 |  |  | a | 279.89 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 22.1 | 14.42 | 0.79233 | 21.3077 | 13.6277 | 0.15121 | 0.08174 | 20181101 | 10-24 |  | 1 | MJM/ELJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 5
EN644 | 20190820 | L1 | 1 | 8 |  |  | b | 280.59 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 21.86 | 14.12 | 0.79233 | 21.0677 | 13.3277 | 0.15201 | 0.075242 | 20181101 | 10-24 |  | 1 | MJM/ELJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 6
EN644 | 20190820 | L1 | 1 | 3 |  |  | a | 281.05 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 101.4 | 63.73 | 0.79233 | 100.6077 | 62.9377 | 0.73863 | 0.33279 | 20181101 | 10-24 |  | 1 | MJM/ELJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 3
EN644 | 20190820 | L1 | 1 | 3 |  |  | b | 279.91 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 104.7 | 63.09 | 0.79233 | 103.9077 | 62.2977 | 0.8192 | 0.24564 | 20181101 | 10-24 |  | 1 | MJM/ELJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 4
EN644 | 20190820 | L1 | 1 | 18 |  |  | a | 280.16 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 11.02 | 6.819 | 0.79233 | 10.2277 | 6.0267 | 0.082634 | 0.020287 | 20181101 | 10-24 |  | 1 | MJM/ELJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 7 | poor phaeo replication (60% range); Taylor previously QC'ed
EN644 | 20190820 | L1 | 1 | 18 |  |  | b | 280.66 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 14.07 | 8.983 | 0.79233 | 13.2777 | 8.1907 | 0.099884 | 0.039744 | 20181101 | 10-24 |  | 1 | MJM/ELJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 8 | poor phaeo replication (60% range); Taylor previously QC'ed
EN644 | 20190820 | L1 | 1 | 8 |  |  | a | 280.02 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 5.545 | 3.631 | 0.79233 | 4.7527 | 2.8387 | 0.037667 | 0.010834 | 20181101 | 10-24 |  | 1 | MJM/ELJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 11
EN644 | 20190820 | L1 | 1 | 8 |  |  | b | 279.94 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.295 | 4.14 | 0.79233 | 5.5027 | 3.3477 | 0.042422 | 0.014793 | 20181101 | 10-24 |  | 1 | MJM/ELJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 12
EN644 | 20190820 | L1 | 1 | 3 |  |  | a | 280.43 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 32.84 | 19.6 | 0.79233 | 32.0477 | 18.8077 | 0.26018 | 0.060699 | 20181101 | 10-24 |  | 1 | MJM/ELJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 9 | poor replication (48% chl range)
EN644 | 20190820 | L1 | 1 | 3 |  |  | b | 280.09 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 53.16 | 31.36 | 0.79233 | 52.3677 | 30.5677 | 0.42892 | 0.093237 | 20181101 | 10-24 |  | 1 | MJM/ELJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 10 | poor replication (48% chl range)
EN644 | 20190820 | L1 | 1 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 71.05 | 43.87 | 0.72767 | 70.3223 | 43.1423 | 0.56068 | 0.18184 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190820 | L1 | 1 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 66.57 | 40.2 | 0.72767 | 65.8423 | 39.4723 | 0.54397 | 0.13538 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190820 | L1 | 1 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 77.25 | 49.64 | 0.72767 | 76.5223 | 48.9123 | 0.56955 | 0.27227 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
EN644 | 20190820 | L1 | 1 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 79.28 | 50.93 | 0.80267 | 78.4773 | 50.1273 | 0.58482 | 0.27792 | 20201113 | 10-80 | 20220312 | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | labNB written C1N1chlb but process of elimination it's N11 and fits with rep a and easy typo to be N1 to N11
EN644 | 20190820 | L1 | 1 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 81.63 | 50.08 | 0.72767 | 80.9023 | 49.3523 | 0.65083 | 0.19857 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | very poor replication, looks like could this be C1N11b? That sample is missing as of Apr20
EN644 | 20190820 | L1 | 1 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 382.5 | 234.3 | 0.72767 | 381.7723 | 233.5723 | 3.0571 | 0.96285 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | how does this compare to HPLC?
EN644 | 20190820 | L1 | 1 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 188.9 | 115.9 | 0.72767 | 188.1723 | 115.1723 | 1.5059 | 0.47634 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | on lab NB 10-80 sample C1N1chl b is written but no C1N1 in meta data, could it be a typo and be C1N2? Or is there a meta data typo? Check log sheets
EN644 | 20190820 | L1 | 1 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 181.5 | 114.7 | 0.80267 | 180.6973 | 113.8973 | 1.378 | 0.5823 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190820 | L1 | 1 | 17 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 91.62 | 59.09 | 0.72767 | 90.8923 | 58.3623 | 0.67104 | 0.33342 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | this <10 is higher than either WSW value; poor chl replication; this is somewhat high given that it's <10 and is higher than both WSW values at this depth
EN644 | 20190820 | L1 | 1 | 17 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 55.25 | 34.53 | 0.72767 | 54.5223 | 33.8023 | 0.42742 | 0.15435 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication; could N17 and N7 have mixed up replicates?
EN644 | 20190820 | L1 | 1 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 64.76 | 39.28 | 0.80267 | 63.9573 | 38.4773 | 0.52561 | 0.13662 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication; could N17 and N7 have mixed up replicates?
EN644 | 20190820 | L1 | 1 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 42.24 | 27.24 | 0.72767 | 41.5123 | 26.5123 | 0.30943 | 0.14687 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication; could N17 and N7 have mixed up replicates?
EN644 | 20190820 | L1 | 1 | 2 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 98.37 | 62.72 | 0.80267 | 97.5673 | 61.9173 | 0.7354 | 0.33025 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190820 | L1 | 1 | 2 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 99.02 | 60.42 | 0.80267 | 98.2173 | 59.6173 | 0.79626 | 0.22981 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L2 | 2 | 21 |  |  | a | 280.47 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 30.33 | 18.54 | 0.79233 | 29.5377 | 17.7477 | 0.23165 | 0.071098 | 20181101 | 10-22 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 13
EN644 | 20190821 | L2 | 2 | 21 |  |  | b | 280.16 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 33.27 | 20.58 | 0.79233 | 32.4777 | 19.7877 | 0.24961 | 0.088312 | 20181101 | 10-22 |  | 2 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 14
EN644 | 20190821 | L2 | 2 | 7 |  |  | a | 281.05 | 5 | 5 | 1 | 0 | 0.33 | Cal_110119 | 1.8682 | 0.5122 | 284.5 | 169.2 | 0.79233 | 283.7077 | 168.4077 | 6.8509 | 1.8367 | 20181101 | 10-22 |  | 3 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 15
EN644 | 20190821 | L2 | 2 | 7 |  |  | b | 279.91 | 5 | 5 | 1 | 0 | 0.33 | Cal_110119 | 1.8682 | 0.5122 | 296.8 | 175 | 0.79233 | 296.0077 | 174.2077 | 7.2665 | 1.7568 | 20181101 | 10-22 |  | 4 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 16
EN644 | 20190821 | L2 | 2 | 4 |  |  | a | 280.02 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 340.3 | 200 | 0.79233 | 339.5077 | 199.2077 | 2.7611 | 0.64259 | 20181101 | 10-22 |  | 5 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 17 | somewhat poor replication (25% range); Taylor previously QC'ed; DF thinks this is within bio variation of size fractionated filtering
EN644 | 20190821 | L2 | 2 | 4 |  |  | b | 280.59 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 398.2 | 216.6 | 0.79233 | 397.4077 | 215.8077 | 3.5666 | 0.11321 | 20181101 | 10-22 |  | 6 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 18 | somewhat poor replication (25% range); Taylor previously QC'ed; DF thinks this is within bio variation of size fractionated filtering
EN644 | 20190821 | L2 | 2 | 21 |  |  | a | 280.66 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.11 | 11.74 | 0.79233 | 18.3177 | 10.9477 | 0.14471 | 0.041916 | 20181101 | 10-22 |  | 7 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 19
EN644 | 20190821 | L2 | 2 | 21 |  |  | b | 280.31 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 20.09 | 12.42 | 0.79233 | 19.2977 | 11.6277 | 0.15079 | 0.047677 | 20181101 | 10-22 |  | 8 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 20
EN644 | 20190821 | L2 | 2 | 7 |  |  | a | 280.43 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.824 | 2.405 | 0.79233 | 3.0317 | 1.6127 | 0.027885 | -0.00037107 | 20181101 | 10-22 |  | 9 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 21 | completely incorrect duplicates, did no water get poured in this funnel?; Taylor previously QC'ed
EN644 | 20190821 | L2 | 2 | 7 |  |  | b | 280.09 | 20 | 5 | 1 | 0 | 0.5 | Cal_110119 | 1.8682 | 0.5122 | 360.7 | 206.4 | 0.79233 | 359.9077 | 205.6077 | 6.0717 | 0.95261 | 20181101 | 10-22 |  | 10 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 22 | completely bad duplicates, is this all the water (280x2)? Or possibly real?; Taylor previously QC'ed, DF thinks this is both duplicates so then >20 value would be half of this
EN644 | 20190821 | L2 | 2 | 4 |  |  | a | 279.89 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 316.6 | 185 | 0.79233 | 315.8077 | 184.2077 | 2.5911 | 0.55777 | 20181101 | 10-22 |  | 11 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 23 | Taylor previously QC'ed
EN644 | 20190821 | L2 | 2 | 4 |  |  | b | 279.94 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 296.8 | 169.8 | 0.79233 | 296.0077 | 169.0077 | 2.5001 | 0.38844 | 20181101 | 10-22 |  | 12 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 24 | Taylor previously QC'ed
EN644 | 20190821 | L2 | 2 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 71.42 | 45.75 | 0.72767 | 70.6923 | 45.0223 | 0.52953 | 0.24534 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L2 | 2 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 77.71 | 47.21 | 0.72767 | 76.9823 | 46.4823 | 0.62917 | 0.17084 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L2 | 2 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 158.6 | 96.57 | 0.80267 | 157.7973 | 95.7673 | 1.2796 | 0.36866 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L2 | 2 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 156.1 | 92.83 | 0.80267 | 155.2973 | 92.0273 | 1.3052 | 0.27871 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L2 | 2 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 389.6 | 269.2 | 0.73967 | 388.8603 | 268.4603 | 2.4837 | 2.1368 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | rather odd that phaeo = chl
EN644 | 20190821 | L2 | 2 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 439.5 | 161 | 0.72767 | 438.7723 | 160.2723 | 5.745 | -2.9866 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | Looks like there was a transcribing error. Could Ra have been 261 and not 161? Chl very high poor replication and very negative phaoe (very extreme Rb/Ra ratio)
EN644 | 20190821 | L2 | 2 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 356.8 | 211.5 | 0.73967 | 356.0603 | 210.7603 | 2.9973 | 0.63006 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | rep a&b very different, could water have been incorrectly been poured more on this filter and not the other?
EN644 | 20190821 | L2 | 2 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 75.85 | 47.02 | 0.72767 | 75.1223 | 46.2923 | 0.59472 | 0.20201 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | clearly bad chl value, could this sample water have accidentally been poured on rep a filter?
EN644 | 20190821 | L2 | 2 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 141.6 | 89.55 | 0.80267 | 140.7973 | 88.7473 | 1.0737 | 0.45371 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | in lab NB 10-80 this C2N2a sample is written within cast 1 samples, could this be a typo and mean to be C1N2a?? Probably not but keep this in mind until all samples have been run and entered
EN644 | 20190821 | L2 | 2 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 147 | 94.88 | 0.72767 | 146.2723 | 94.1523 | 1.0752 | 0.54529 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L2 | 2 | 17 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 52.85 | 32.54 | 0.72767 | 52.1223 | 31.8123 | 0.41896 | 0.12856 | 20201113 | 11-00 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L2 | 2 | 17 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 49.75 | 30.76 | 0.72767 | 49.0223 | 30.0323 | 0.39173 | 0.12515 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L2 | 2 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 172.9 | 106.1 | 0.72767 | 172.1723 | 105.3723 | 1.378 | 0.43557 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication and phaeo/chl ratio very different between the replicates
EN644 | 20190821 | L2 | 2 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 142.5 | 94.03 | 0.72767 | 141.7723 | 93.3023 | 0.99986 | 0.60596 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication and phaeo/chl ratio very different between the replicates
EN644 | 20190821 | L2 | 2 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 42.41 | 26.16 | 0.72767 | 41.6823 | 25.4323 | 0.33521 | 0.1025 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor clh replication but seems like dynamic cast with high variability in a lot of samples; yeah DF thinks this is within bio variation
EN644 | 20190821 | L2 | 2 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 59.52 | 35.63 | 0.72767 | 58.7923 | 34.9023 | 0.49281 | 0.10789 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor clh replication but seems like dynamic cast with high variability in a lot of samples; yeah DF thinks this is within bio variation
EN644 | 20190821 | L3 | 5 | 16 |  |  | a | 280.31 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 314.1 | 194.7 | 0.79233 | 313.3077 | 193.9077 | 2.3474 | 0.96235 | 20181101 | 10-22 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 25; odd that >5 values are larger than both WSW values, but this water was from a different niskin; QF was 1
EN644 | 20190821 | L3 | 5 | 16 |  |  | b | 280.47 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 322.6 | 198.6 | 0.79233 | 321.8077 | 197.8077 | 2.4364 | 0.93795 | 20181101 | 10-22 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 26; odd that >5 values are larger than both WSW values, but this water was from a different niskin; QF was 1
EN644 | 20190821 | L3 | 5 | 12 |  |  | a | 281.05 | 5 | 5 | 1 | 0 | 0.5 | Cal_110119 | 1.8682 | 0.5122 | 345.3 | 215.2 | 0.79233 | 344.5077 | 214.4077 | 5.102 | 2.198 | 20181101 | 10-22 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 27 | somewhat poor chl replication (21%), phaeos don't match, dilution additional sourc eof error; Taylor previously QC'ed
EN644 | 20190821 | L3 | 5 | 12 |  |  | b | 280.09 | 5 | 5 | 1 | 0 | 0.5 | Cal_110119 | 1.8682 | 0.5122 | 391.2 | 230 | 0.79233 | 390.4077 | 229.2077 | 6.3432 | 1.4874 | 20181101 | 10-22 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 28 | somewhat poor chl replication (21%), phaeos don't match, dilution additional sourc eof error; Taylor previously QC'ed
EN644 | 20190821 | L3 | 5 | 9 |  |  | a | 279.94 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 80.44 | 53.69 | 0.79233 | 79.6477 | 52.8977 | 0.52659 | 0.37749 | 20181101 | 10-22 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 29
EN644 | 20190821 | L3 | 5 | 9 |  |  | b | 280.02 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 71.06 | 47.84 | 0.79233 | 70.2677 | 47.0477 | 0.45697 | 0.34689 | 20181101 | 10-22 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 30
EN644 | 20190821 | L3 | 5 | 16 |  |  | a | 280.16 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 267.6 | 167.7 | 0.79233 | 266.8077 | 166.9077 | 1.965 | 0.88534 | 20181101 | 10-22 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 31; values are very close to WSW values but this was was from a different niskin, QF was 1, DF changed QF from a 2 back to 1
EN644 | 20190821 | L3 | 5 | 16 |  |  | b | 280.66 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 251.1 | 156.6 | 0.79233 | 250.3077 | 155.8077 | 1.8555 | 0.80057 | 20181101 | 10-22 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 32; values are very close to WSW values but this was was from a different niskin, QF was 1; DF changed QF from a 2 back to 1
EN644 | 20190821 | L3 | 5 | 12 |  |  | a | 279.91 | 20 | 5 | 1 | 0 | 0.5 | Cal_110119 | 1.8682 | 0.5122 | 335.2 | 201.9 | 0.79233 | 334.4077 | 201.1077 | 5.2487 | 1.6263 | 20181101 | 10-22 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 33 | somewhat poor chl replication and value is larger than WSW value
EN644 | 20190821 | L3 | 5 | 12 |  |  | b | 280.43 | 20 | 5 | 1 | 0 | 0.5 | Cal_110119 | 1.8682 | 0.5122 | 367.4 | 215 | 0.79233 | 366.6077 | 214.2077 | 5.9897 | 1.3196 | 20181101 | 10-22 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 34 | somewhat poor chl replication and value is larger than WSW value
EN644 | 20190821 | L3 | 5 | 9 |  |  | a | 280.59 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 32.17 | 21.19 | 0.79233 | 31.3777 | 20.3977 | 0.21565 | 0.13216 | 20181101 | 10-22 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 35 | somewhat poor replication (24% range); Taylor previously QC'ed
EN644 | 20190821 | L3 | 5 | 9 |  |  | b | 279.89 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 26.13 | 17.53 | 0.79233 | 25.3377 | 16.7377 | 0.16933 | 0.11679 | 20181101 | 10-22 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 36 | somewhat poor replication (24% range); Taylor previously QC'ed
EN644 | 20190821 | L3 | 5 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 328.4 | 222.3 | 0.72767 | 327.6723 | 221.5723 | 2.1887 | 1.6248 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L3 | 5 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 319.4 | 222 | 0.72767 | 318.6723 | 221.2723 | 2.0092 | 1.7991 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L3 | 5 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 385.8 | 200.9 | 0.73967 | 385.0603 | 200.1603 | 3.8142 | -0.36926 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | negative phaeo but not horribly negative so leave
EN644 | 20190821 | L3 | 5 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 420.5 | 122.4 | 0.72767 | 419.7723 | 121.6723 | 6.1493 | -4.0553 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | appears very likely to be transcribing error when reading Ra, seems likely should be 222.4 not 122.4; poor chl replication and very high negative phaeo
EN644 | 20190821 | L3 | 5 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 101 | 69.26 | 0.72767 | 100.2723 | 68.5323 | 0.65475 | 0.52476 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | pretty high phaeo for depth?
EN644 | 20190821 | L3 | 5 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 102.9 | 68.76 | 0.72767 | 102.1723 | 68.0323 | 0.70426 | 0.46664 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L3 | 5 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 147.4 | 99.05 | 0.80267 | 146.5973 | 98.2473 | 0.99739 | 0.69354 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L3 | 5 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 149.8 | 102.1 | 0.72767 | 149.0723 | 101.3723 | 0.98398 | 0.76073 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L3 | 5 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 162.3 | 106.9 | 0.72767 | 161.5723 | 106.1723 | 1.1428 | 0.68451 | 20201113 | 11-00 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | slight increase in chl at depth matches cast fluor profile
EN644 | 20190821 | L3 | 5 | 4 |  |  | b | 290 | 0 |  |  |  |  | Cal_111320b_12mm | 1.8343 | 0.5442 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 9 | 0
EN644 | 20190821 | L3 | 5 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 44.04 | 32.03 | 0.72767 | 43.3123 | 31.3023 | 0.24775 | 0.29099 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L3 | 5 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 48.04 | 34.45 | 0.72767 | 47.3123 | 33.7223 | 0.28034 | 0.30005 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L3 | 5 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 24.53 | 15.87 | 0.72767 | 23.8023 | 15.1423 | 0.17864 | 0.081971 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | lab NB says repb but repb already entered so enter as rep a
EN644 | 20190821 | L3 | 5 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 20.24 | 12.82 | 0.80267 | 19.4373 | 12.0173 | 0.15306 | 0.053766 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L6 | 7 | 17 |  |  | a | 280.31 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 94.37 | 61.99 | 0.79233 | 93.5777 | 61.1977 | 0.63658 | 0.40797 | 20181101 | 10-22 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 37
EN644 | 20190821 | L6 | 7 | 17 |  |  | b | 280.47 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 101.6 | 65.05 | 0.79233 | 100.8077 | 64.2577 | 0.71815 | 0.37801 | 20181101 | 10-22 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 38
EN644 | 20190821 | L6 | 7 | 9 |  |  | a | 280.43 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 246.1 | 151.1 | 0.79233 | 245.3077 | 150.3077 | 1.8669 | 0.69756 | 20181101 | 10-22 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 39
EN644 | 20190821 | L6 | 7 | 9 |  |  | b | 279.91 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 246.4 | 152.4 | 0.79233 | 245.6077 | 151.6077 | 1.8506 | 0.74076 | 20181101 | 10-22 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 40
EN644 | 20190821 | L6 | 7 | 5 |  |  | a | 279.94 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 78.95 | 49.73 | 0.79233 | 78.1577 | 48.9377 | 0.57521 | 0.26118 | 20181101 | 10-22 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 41
EN644 | 20190821 | L6 | 7 | 5 |  |  | b | 279.89 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 88.69 | 55.07 | 0.79233 | 87.8977 | 54.2777 | 0.66195 | 0.26588 | 20181101 | 10-22 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 42
EN644 | 20190821 | L6 | 7 | 17 |  |  | a | 280.16 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 93.29 | 58.53 | 0.79233 | 92.4977 | 57.7377 | 0.68373 | 0.30229 | 20181101 | 10-22 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 43 | larger than one of >5 reps; QF was 1; DF thinks >5 is basically equal to >20 so all chl signal is coming from >20; thus changed QF back to 1
EN644 | 20190821 | L6 | 7 | 17 |  |  | b | 280.66 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 98.3 | 62.57 | 0.79233 | 97.5077 | 61.7777 | 0.70156 | 0.35157 | 20181101 | 10-22 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 44 | larger than one of >5 reps; QF was 1; DF thinks >5 is basically equal to >20 so all chl signal is coming from >20; thus changed QF back to 1
EN644 | 20190821 | L6 | 7 | 9 |  |  | a | 281.05 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 59.72 | 36.89 | 0.79233 | 58.9277 | 36.0977 | 0.44765 | 0.16686 | 20181101 | 10-22 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 45
EN644 | 20190821 | L6 | 7 | 9 |  |  | b | 280.09 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 64.56 | 39.34 | 0.79233 | 63.7677 | 38.5477 | 0.4962 | 0.16226 | 20181101 | 10-22 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 46
EN644 | 20190821 | L6 | 7 | 5 |  |  | a | 280.02 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 48.5 | 29.72 | 0.79233 | 47.7077 | 28.9277 | 0.36959 | 0.12467 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 47 | Taylor previously QC'ed
EN644 | 20190821 | L6 | 7 | 5 |  |  | b | 280.59 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 59.39 | 36.26 | 0.79233 | 58.5977 | 35.4677 | 0.45427 | 0.1505 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 48 | Taylor previously QC'ed
EN644 | 20190821 | L6 | 7 | 13 |  |  | a | 281.08 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 88.6 | 54.15 | 0.79233 | 87.8077 | 53.3577 | 0.67542 | 0.23282 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 49
EN644 | 20190821 | L6 | 7 | 13 |  |  | b | 280.17 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 97.05 | 58.79 | 0.79233 | 96.2577 | 57.9977 | 0.75255 | 0.23787 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 50
EN644 | 20190821 | L6 | 7 | 13 |  |  | a | 279.76 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 50.45 | 29.63 | 0.79233 | 49.6577 | 28.8377 | 0.41012 | 0.083065 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 51
EN644 | 20190821 | L6 | 7 | 13 |  |  | b | 279.67 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 48.41 | 28.2 | 0.79233 | 47.6177 | 27.4077 | 0.39823 | 0.070648 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 52
EN644 | 20190821 | L6 | 7 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 137.8 | 86.83 | 0.80267 | 136.9973 | 86.0273 | 1.0514 | 0.42918 | 20201113 | 10-81 |  | 1 |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0 | labNB written C7N11a but process of elimination should be C7N17a
EN644 | 20190821 | L6 | 7 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 143.5 | 89.28 | 0.73967 | 142.7603 | 88.5403 | 1.1185 | 0.40538 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
EN644 | 20190821 | L6 | 7 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 176.8 | 106.5 | 0.73967 | 176.0603 | 105.7603 | 1.4502 | 0.37005 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L6 | 7 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 164.3 | 100.7 | 0.72767 | 163.5723 | 99.9723 | 1.312 | 0.40864 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L6 | 7 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 262.9 | 163.5 | 0.72767 | 262.1723 | 162.7723 | 2.0505 | 0.75099 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | both tissue capsules labeled repa so just arbitrarily choose 1 as rep b
EN644 | 20190821 | L6 | 7 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 261.4 | 164.7 | 0.72767 | 260.6723 | 163.9723 | 1.9948 | 0.82734 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | both tissue capsules labeled repa so just arbitrarily choose 1 as rep b
EN644 | 20190821 | L6 | 7 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 436.6 | 285.5 | 0.72767 | 435.8723 | 284.7723 | 3.117 | 1.7842 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor replication
EN644 | 20190821 | L6 | 7 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 394.2 | 286.8 | 0.73967 | 393.4603 | 286.0603 | 2.2155 | 2.7079 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor replication and this chl lower than adjacent water and phaeo double chl so seems like this one is the error
EN644 | 20190821 | L6 | 7 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 442.2 | 272.6 | 0.73967 | 441.4603 | 271.8603 | 3.4986 | 1.1804 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L6 | 7 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 435.3 | 264.9 | 0.80267 | 434.4973 | 264.0973 | 3.5151 | 1.0303 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L6 | 7 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 146.5 | 94.43 | 0.72767 | 145.7723 | 93.7023 | 1.0741 | 0.53858 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190821 | L6 | 7 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 141.4 | 87.27 | 0.72767 | 140.6723 | 86.5423 | 1.1166 | 0.37285 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190821 | L6 | 7 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 44.35 | 30.37 | 0.72767 | 43.6223 | 29.6423 | 0.28839 | 0.22179 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L6 | 7 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 40.64 | 29.07 | 0.80267 | 39.8373 | 28.2673 | 0.23867 | 0.24784 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L6 | 7 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 25.26 | 20.7 | 0.80267 | 24.4573 | 19.8973 | 0.094066 | 0.24839 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L6 | 7 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 26.19 | 21.51 | 0.72767 | 25.4623 | 20.7823 | 0.096541 | 0.26114 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L6 | 7 | 17 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 25.44 | 16.02 | 0.73967 | 24.7003 | 15.2803 | 0.19432 | 0.068668 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L6 | 7 | 17 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 27.66 | 17.1 | 0.73967 | 26.9203 | 16.3603 | 0.21784 | 0.06374 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L6 | 7 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 401.3 | 240.8 | 0.80267 | 400.4973 | 239.9973 | 3.3109 | 0.8197 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | awful high value for <10 (greater than either WSW at this depth), are we sure this was <10?
EN644 | 20190821 | L6 | 7 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 351.4 | 224.9 | 0.72767 | 350.6723 | 224.1723 | 2.6095 | 1.2487 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very high chl for <10. are we sure this was filtered?
EN644 | 20190821 | L6 | 7 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 85.92 | 56.81 | 0.72767 | 85.1923 | 56.0823 | 0.60049 | 0.36473 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190821 | L6 | 7 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 84.02 | 53.31 | 0.72767 | 83.2923 | 52.5823 | 0.6335 | 0.27149 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 21 |  |  | a | 280.66 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 24.04 | 14.99 | 0.79233 | 23.2477 | 14.1977 | 0.1777 | 0.064333 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 53 | Taylor previously QC'ed
EN644 | 20190822 | L7 | 8 | 21 |  |  | b | 280.31 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.97 | 12.66 | 0.79233 | 19.1777 | 11.8677 | 0.14371 | 0.058851 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 54 | Taylor previously QC'ed
EN644 | 20190822 | L7 | 8 | 10 |  |  | a | 279.91 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 144.8 | 88.33 | 0.79233 | 144.0077 | 87.5377 | 1.1118 | 0.3845 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 55
EN644 | 20190822 | L7 | 8 | 10 |  |  | b | 280.09 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 143.8 | 87.68 | 0.79233 | 143.0077 | 86.8877 | 1.1042 | 0.38004 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 56
EN644 | 20190822 | L7 | 8 | 8 |  |  | a | 279.89 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 45.7 | 30.2 | 0.79233 | 44.9077 | 29.4077 | 0.30518 | 0.19752 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 57
EN644 | 20190822 | L7 | 8 | 8 |  |  | b | 280.02 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 46.21 | 30.58 | 0.79233 | 45.4177 | 29.7877 | 0.3076 | 0.20136 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 58
EN644 | 20190822 | L7 | 8 | 21 |  |  | a | 280.16 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 18.05 | 11.14 | 0.79233 | 17.2577 | 10.3477 | 0.13592 | 0.040793 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 59 | Taylor previously QC'ed
EN644 | 20190822 | L7 | 8 | 21 |  |  | b | 280.47 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 15.66 | 9.522 | 0.79233 | 14.8677 | 8.7297 | 0.1206 | 0.028315 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 60 | Taylor previously QC'ed
EN644 | 20190822 | L7 | 8 | 10 |  |  | a | 281.05 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 101.1 | 59.24 | 0.79233 | 100.3077 | 58.4477 | 0.82078 | 0.1742 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 61 | somewhat poor chl replication (32% range) but both good to be less than 5's, should this be flag2? What's acceptable range of error for the size frac?; Taylor previously QC'ed, DF thinks replication is good for this size class, QF was 2
EN644 | 20190822 | L7 | 8 | 10 |  |  | b | 280.43 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 75.66 | 45.42 | 0.79233 | 74.8677 | 44.6277 | 0.59425 | 0.16715 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 62 | somewhat poor chl replication (32% range) but both good to be less than 5's, should this be flag2? What's acceptable range of error for the size frac?; Taylor previously QC'ed, DF thinks replication is good for this size class, QF was 2
EN644 | 20190822 | L7 | 8 | 8 |  |  | a | 280.59 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.91 | 13.1 | 0.79233 | 19.1177 | 12.3077 | 0.13375 | 0.076115 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 63
EN644 | 20190822 | L7 | 8 | 8 |  |  | b | 279.94 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 21.43 | 14.42 | 0.79233 | 20.6377 | 13.6277 | 0.138 | 0.094915 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 64
EN644 | 20190822 | L7 | 8 | 21 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 37.89 | 23.89 | 1.2035 | 36.6865 | 22.6865 | 0.2888 | 0.10166 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication
EN644 | 20190822 | L7 | 8 | 21 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 44.98 | 26.93 | 0.80267 | 44.1773 | 26.1273 | 0.37234 | 0.077331 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication
EN644 | 20190822 | L7 | 8 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 34.41 | 21.07 | 0.80267 | 33.6073 | 20.2673 | 0.27518 | 0.073636 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 35.07 | 21.5 | 0.80267 | 34.2673 | 20.6973 | 0.27993 | 0.076292 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 57.16 | 36.78 | 0.72767 | 56.4323 | 36.0523 | 0.42041 | 0.20009 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 53.69 | 33.85 | 0.80267 | 52.8873 | 33.0473 | 0.40927 | 0.15951 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 302.2 | 186.3 | 0.72767 | 301.4723 | 185.5723 | 2.3908 | 0.80303 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 305.6 | 186.5 | 0.80267 | 304.7973 | 185.6973 | 2.4569 | 0.73917 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 89.85 | 57.85 | 0.73967 | 89.1103 | 57.1103 | 0.66011 | 0.32281 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 98.76 | 63.46 | 0.80267 | 97.9573 | 62.6573 | 0.72818 | 0.3502 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 25.32 | 19.14 | 0.80267 | 24.5173 | 18.3373 | 0.12748 | 0.18812 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 26.81 | 19.69 | 0.80267 | 26.0073 | 18.8873 | 0.14687 | 0.17819 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 29.39 | 21.22 | 0.72767 | 28.6623 | 20.4923 | 0.16853 | 0.18416 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 28.45 | 20.17 | 0.73967 | 27.7103 | 19.4303 | 0.1708 | 0.16361 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 19.14 | 15.23 | 0.72767 | 18.4123 | 14.5023 | 0.080657 | 0.16894 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 19.18 | 14.87 | 0.73967 | 18.4403 | 14.1303 | 0.088909 | 0.15429 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 21.15 | 13.58 | 0.80267 | 20.3473 | 12.7773 | 0.15616 | 0.063752 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L7 | 8 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 20.99 | 13.87 | 0.72767 | 20.2623 | 13.1423 | 0.14687 | 0.079317 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 18 |  |  | a | 280.16 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 21.12 | 14.06 | 0.79233 | 20.3277 | 13.2677 | 0.13887 | 0.087709 | 20181101 | 10-24 |  | 1 | Erin J | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 65
EN644 | 20190822 | L8 | 9 | 18 |  |  | b | 280.47 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 21.24 | 14.42 | 0.79233 | 20.4477 | 13.6277 | 0.134 | 0.098469 | 20181101 | 10-24 |  | 1 | Erin J | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 66
EN644 | 20190822 | L8 | 9 | 14 |  |  | a | 281.08 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.32 | 13.28 | 0.79233 | 18.5277 | 12.4877 | 0.11842 | 0.094143 | 20181101 | 10-25 |  | 1 | Erin J | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 77
EN644 | 20190822 | L8 | 9 | 14 |  |  | b | 279.67 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.38 | 13.02 | 0.79233 | 18.5877 | 12.2277 | 0.12532 | 0.083864 | 20181101 | 10-25 |  | 1 | Erin J | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 78
EN644 | 20190822 | L8 | 9 | 8 |  |  | a | 279.91 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 36.74 | 25.26 | 0.79233 | 35.9477 | 24.4677 | 0.22601 | 0.19221 | 20181101 | 10-24 |  | 1 | Erin J | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 67
EN644 | 20190822 | L8 | 9 | 8 |  |  | b | 280.09 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 41.69 | 29.07 | 0.79233 | 40.8977 | 28.2777 | 0.2483 | 0.23474 | 20181101 | 10-24 |  | 1 | Erin J | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 68
EN644 | 20190822 | L8 | 9 | 5 |  |  | a | 280.59 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 21.98 | 15.35 | 0.79233 | 21.1877 | 14.5577 | 0.13021 | 0.11802 | 20181101 | 10-24 |  | 1 | Diana/Taylor | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 69 | Erin got seasick, so some confusion. Either 69 or 70 was missing 92 mL, seems clear TAR70 was missing the vol; Taylor previously QC'ed, DF thinks maybe these were switched with >20 fractions reps?
EN644 | 20190822 | L8 | 9 | 5 |  |  | b | 188 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 14.42 | 10.03 | 0.79233 | 13.6277 | 9.2377 | 0.12868 | 0.10641 | 20181101 | 10-24 |  | 1 | Diana/Taylor | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 70 | Erin got seasick, so some confusion. Either 69 or 70 was missing 92 mL, seems clear TAR70 was missing the vol; Taylor previously QC'ed, DF thinks maybe these were switched with >20 fractions reps?
EN644 | 20190822 | L8 | 9 | 18 |  |  | a | 280.66 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 18.62 | 12.8 | 0.79233 | 17.8277 | 12.0077 | 0.11428 | 0.090421 | 20181101 | 10-24 |  | 1 | Diana/Taylor | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 71 | poor chl %range 49% but low chl values and phaeo good match; Taylor previously QC'ed
EN644 | 20190822 | L8 | 9 | 18 |  |  | b | 280.31 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 27.67 | 18.08 | 0.79233 | 26.8777 | 17.2877 | 0.18854 | 0.10654 | 20181101 | 10-24 |  | 1 | Diana/Taylor | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 72 | this value greater than the >5's; poor chl %range 49% but low chl values and phaeo good match; Taylor previously QC'ed
EN644 | 20190822 | L8 | 9 | 14 |  |  | a | 275 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.06 | 8.506 | 0.79233 | 12.2677 | 7.7137 | 0.091259 | 0.042944 | 20181101 | 10-25 |  | 1 | Diana/Taylor | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 79 | Bottle wasn't filled to top, so estimated volume using filter cups and grad cylinder. 38% chl range; Taylor previously QC'ed
EN644 | 20190822 | L8 | 9 | 14 |  |  | b | 275 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 17.83 | 11.13 | 0.79233 | 17.0377 | 10.3377 | 0.13426 | 0.045592 | 20181101 | 10-25 |  | 1 | Diana/Taylor | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 80 | Bottle wasn't filled to top, so estimated volume using filter cups and grad cylinder. 38% chl range; Taylor previously QC'ed
EN644 | 20190822 | L8 | 9 | 8 |  |  | a | 280.43 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 14.81 | 10.18 | 0.79233 | 14.0177 | 9.3877 | 0.090985 | 0.069179 | 20181101 | 10-24 |  | 1 | Diana/Taylor | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 73 | poor replication 49% chl range but this one appear to have accurate vol filt as opposed to replicate; Taylor previously QC'ed
EN644 | 20190822 | L8 | 9 | 8 |  |  | b | 240 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 19.25 | 12.68 | 0.79233 | 18.4577 | 11.8877 | 0.15086 | 0.086126 | 20181101 | 10-24 |  | 1 | Diana/Taylor | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 74 | Bottle wasn't filled to top, so estimated volume using filter cups and grad cylinder. Poor chl replication (49% range); Taylor previously QC'ed
EN644 | 20190822 | L8 | 9 | 5 |  |  | a | 279.89 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 20.88 | 13.75 | 0.79233 | 20.0877 | 12.9577 | 0.14038 | 0.081116 | 20181101 | 10-24 |  | 1 | Diana/Taylor | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 75 | Taylor previously QC'ed; could these be switched with >5 reps? These values are larger than the >5 rep values
EN644 | 20190822 | L8 | 9 | 5 |  |  | b | 279.94 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 20.5 | 13.7 | 0.79233 | 19.7077 | 12.9077 | 0.13386 | 0.086743 | 20181101 | 10-24 |  | 1 | Diana/Taylor | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 76 | Taylor previously QC'ed; could these be switched with >5 reps? These values are larger than the >5 rep values
EN644 | 20190822 | L8 | 9 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 36.78 | 24.6 | 0.80267 | 35.9773 | 23.7973 | 0.25125 | 0.15832 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication but chl/phaeo ratio consistent so natural variation?
EN644 | 20190822 | L8 | 9 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 51.47 | 33.69 | 0.80267 | 50.6673 | 32.8873 | 0.36677 | 0.19925 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | chl a bit high relative to other surf waters; poor chl replication but chl/phaeo ratio consistent so natural variation?
EN644 | 20190822 | L8 | 9 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 31.95 | 21.62 | 0.73967 | 31.2103 | 20.8803 | 0.21309 | 0.14628 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 33.27 | 21.54 | 0.72767 | 32.5423 | 20.8123 | 0.24197 | 0.11623 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 33.45 | 21.77 | 0.80267 | 32.6473 | 20.9673 | 0.24094 | 0.11993 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 34.81 | 22.52 | 0.80267 | 34.0073 | 21.7173 | 0.25352 | 0.12025 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 109.4 | 73.69 | 0.80267 | 108.5973 | 72.8873 | 0.73664 | 0.51781 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 111 | 75.06 | 0.80267 | 110.1973 | 74.2573 | 0.74139 | 0.53665 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 35.13 | 22.85 | 0.80267 | 34.3273 | 22.0473 | 0.25332 | 0.12614 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 37.36 | 23.46 | 1.2035 | 36.1565 | 22.2565 | 0.28674 | 0.096319 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 10.63 | 7.897 | 0.72767 | 9.9023 | 7.1693 | 0.056378 | 0.067013 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 10.54 | 7.943 | 0.72767 | 9.8123 | 7.2153 | 0.053572 | 0.07061 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 7.734 | 6.103 | 1.2035 | 6.5305 | 4.8995 | 0.033645 | 0.05068 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | porr chl replication but very low values
EN644 | 20190822 | L8 | 9 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 7.545 | 5.482 | 1.2035 | 6.3415 | 4.2785 | 0.042557 | 0.03108 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication but very low values
EN644 | 20190822 | L8 | 9 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 9.158 | 7.098 | 0.80267 | 8.3553 | 6.2953 | 0.042495 | 0.065854 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high phaeo range
EN644 | 20190822 | L8 | 9 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.17 | 9.023 | 0.80267 | 10.3673 | 8.2203 | 0.044289 | 0.09719 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high phaeo range
EN644 | 20190822 | L8 | 9 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 17.75 | 11.74 | 0.73967 | 17.0103 | 11.0003 | 0.12398 | 0.065349 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 16.84 | 11.53 | 0.73967 | 16.1003 | 10.7903 | 0.10954 | 0.076174 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 72.29 | 48.36 | 0.73967 | 71.5503 | 47.6203 | 0.49364 | 0.32595 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 69.06 | 47.48 | 0.72767 | 68.3323 | 46.7523 | 0.44516 | 0.35949 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 16.72 | 11.06 | 0.72767 | 15.9923 | 10.3323 | 0.11676 | 0.061072 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L8 | 9 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 16.18 | 10.72 | 0.80267 | 15.3773 | 9.9173 | 0.11263 | 0.058055 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L9 | 12 | 18 |  |  | a | 280.66 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 64.78 | 45.74 | 0.79233 | 63.9877 | 44.9477 | 0.37385 | 0.39238 | 20181101 | 10-25 |  | 1 | Heather/Sam | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 81 | Taylor previously QC'ed; DF thinks these should be 20s
EN644 | 20190822 | L9 | 12 | 18 |  |  | b | 280.16 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 67.26 | 46.98 | 0.79233 | 66.4677 | 46.1877 | 0.39891 | 0.38986 | 20181101 | 10-25 |  | 1 | Heather/Sam | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 82 | Taylor previously QC'ed; DF thinks these should be 20s
EN644 | 20190822 | L9 | 12 | 10 |  |  | a | 280.09 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 25.52 | 18.21 | 0.79233 | 24.7277 | 17.4177 | 0.14382 | 0.1537 | 20181101 | 10-25 |  | 1 | Heather/Sam | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 83
EN644 | 20190822 | L9 | 12 | 10 |  |  | b | 280.43 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 28.16 | 20.64 | 0.79233 | 27.3677 | 19.8477 | 0.14778 | 0.19085 | 20181101 | 10-25 |  | 1 | Heather/Sam | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 84
EN644 | 20190822 | L9 | 12 | 7 |  |  | a | 280.59 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 16.57 | 12.03 | 0.79233 | 15.7777 | 11.2377 | 0.089165 | 0.10245 | 20181101 | 10-25 |  | 1 | Heather/Sam | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 85
EN644 | 20190822 | L9 | 12 | 7 |  |  | b | 279.89 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 16.26 | 12.02 | 0.79233 | 15.4677 | 11.2277 | 0.083482 | 0.10844 | 20181101 | 10-25 |  | 1 | Heather/Sam | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 86
EN644 | 20190822 | L9 | 12 | 18 |  |  | a | 280.31 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 64.36 | 41.4 | 0.79233 | 63.5677 | 40.6077 | 0.45138 | 0.24173 | 20181101 | 10-25 |  | 1 | Heather/Sam | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 87 | DF thinks these should be 5s
EN644 | 20190822 | L9 | 12 | 18 |  |  | b | 280.47 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 69 | 45.23 | 0.79233 | 68.2077 | 44.4377 | 0.46704 | 0.29101 | 20181101 | 10-25 |  | 1 | Heather/Sam | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 88 | DF thinks these should be 5s
EN644 | 20190822 | L9 | 12 | 10 |  |  | a | 281.05 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 13.79 | 10.04 | 0.79233 | 12.9977 | 9.2477 | 0.073529 | 0.083898 | 20181101 | 10-25 |  | 1 | Heather/Sam | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 89
EN644 | 20190822 | L9 | 12 | 10 |  |  | b | 279.91 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 11.78 | 8.691 | 0.79233 | 10.9877 | 7.8987 | 0.060815 | 0.074195 | 20181101 | 10-25 |  | 1 | Heather/Sam | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 90
EN644 | 20190822 | L9 | 12 | 7 |  |  | a | 279.94 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 10.61 | 8.148 | 0.79233 | 9.8177 | 7.3557 | 0.048466 | 0.07725 | 20181101 | 10-25 |  | 1 | Heather/Sam | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 91
EN644 | 20190822 | L9 | 12 | 7 |  |  | b | 280.02 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 8.815 | 6.772 | 0.79233 | 8.0227 | 5.9797 | 0.040206 | 0.061963 | 20181101 | 10-25 |  | 1 | Heather/Sam | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 92
EN644 | 20190822 | L9 | 12 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 93.47 | 62.2 | 0.80267 | 92.6673 | 61.3973 | 0.64505 | 0.41165 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L9 | 12 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 88.39 | 59.83 | 0.80267 | 87.5873 | 59.0273 | 0.58915 | 0.42677 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L9 | 12 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 25.23 | 16.52 | 0.72767 | 24.5023 | 15.7923 | 0.17967 | 0.092126 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190822 | L9 | 12 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 31.59 | 20.77 | 1.2035 | 30.3865 | 19.5665 | 0.2232 | 0.11356 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190822 | L9 | 12 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 26.76 | 17.4 | 0.80267 | 25.9573 | 16.5973 | 0.19308 | 0.092573 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L9 | 12 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 26.96 | 17.58 | 0.80267 | 26.1573 | 16.7773 | 0.1935 | 0.095258 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L9 | 12 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 31.29 | 21.09 | 0.72767 | 30.5623 | 20.3623 | 0.21041 | 0.14004 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L9 | 12 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 31.09 | 20.77 | 0.72767 | 30.3623 | 20.0423 | 0.21289 | 0.13206 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L9 | 12 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 72.04 | 49.4 | 0.80267 | 71.2373 | 48.5973 | 0.46703 | 0.36938 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L9 | 12 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 68.25 | 47.93 | 0.80267 | 67.4473 | 47.1273 | 0.41917 | 0.39193 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L9 | 12 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 51.17 | 37.06 | 1.2035 | 49.9665 | 35.8565 | 0.29107 | 0.32606 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L9 | 12 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 49.71 | 35.24 | 0.80267 | 48.9073 | 34.4373 | 0.29849 | 0.2942 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L9 | 12 | 5 |  |  | a | 280 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 19.15 | 13.91 | 0.72767 | 18.4223 | 13.1823 | 0.11195 | 0.12303 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L9 | 12 | 5 |  |  | b | 280 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 18.52 | 13.52 | 0.73967 | 17.7803 | 12.7803 | 0.10683 | 0.12099 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L9 | 12 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 8.663 | 6.215 | 0.80267 | 7.8603 | 5.4123 | 0.050498 | 0.042653 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication but very low values; would think phaeo ratio would have phaeo a little higher for this sample
EN644 | 20190822 | L9 | 12 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 7.371 | 5.91 | 0.80267 | 6.5683 | 5.1073 | 0.030138 | 0.057764 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication but very low values
EN644 | 20190822 | L9 | 12 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 8.273 | 6.645 | 1.2035 | 7.0695 | 5.4415 | 0.033583 | 0.06007 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190822 | L9 | 12 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 7.284 | 6.105 | 1.2035 | 6.0805 | 4.9015 | 0.024321 | 0.060038 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190822 | L9 | 12 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 21.23 | 13.54 | 0.80267 | 20.4273 | 12.7373 | 0.15863 | 0.060588 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L9 | 12 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 21.3 | 13.3 | 0.73967 | 20.5603 | 12.5603 | 0.16503 | 0.051147 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L10 | 13 | 18 |  |  | a | 280.66 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.501 | 2.337 | 0.79233 | 2.7087 | 1.5447 | 0.022855 | 0.003477 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 93 | poor chl replication (47% range) but very low values and both appropriately greater than the >20s; Taylor previously QC'ed
EN644 | 20190822 | L10 | 13 | 18 |  |  | b | 280.47 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.668 | 1.952 | 0.79233 | 1.8757 | 1.1597 | 0.014068 | 0.0057142 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 94 | poor chl replication (47% range) but very low values and both appropriately greater than the >20s; Taylor previously QC'ed
EN644 | 20190822 | L10 | 13 | 9 |  |  | a | 190 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 20.23 | 14.7 | 0.79233 | 19.4377 | 13.9077 | 0.16039 | 0.18982 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 95 | Only half bottle poured by accident. Vol filt entered as approx "1/2 bottle"; Taylor previously QC'ed, with adjusted volume replication is okay, QF was 3
EN644 | 20190822 | L10 | 13 | 9 |  |  | b | 280.09 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 29.24 | 21.31 | 0.79233 | 28.4477 | 20.5177 | 0.15602 | 0.19446 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 96 | Taylor previously QC'ed
EN644 | 20190822 | L10 | 13 | 5 |  |  | a | 279.94 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.473 | 5.637 | 0.79233 | 5.6807 | 4.8447 | 0.016457 | 0.066343 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 97 | very poor phaeo replication (123% range); Taylor previously QC'ed
EN644 | 20190822 | L10 | 13 | 5 |  |  | b | 280.59 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.604 | 2.71 | 0.79233 | 2.8117 | 1.9177 | 0.017558 | 0.015141 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 98 | very poor phaeo replication (123% range); Taylor previously QC'ed
EN644 | 20190822 | L10 | 13 | 18 |  |  | a | 280.31 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.601 | 1.558 | 0.79233 | 0.80867 | 0.76567 | 0.00084536 | 0.012223 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 99 | poor replication, readings within noise of fluorometer, basically value of zero; Taylor previously QC'ed
EN644 | 20190822 | L10 | 13 | 18 |  |  | b | 280.16 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.163 | 0.811 | 0.79233 | 0.37067 | 0.018667 | 0.0069239 | -0.0066051 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 100 | poor replication, toss since negative phaeo, readings within noise of fluorometer, basically value of zero; Taylor previously QC'ed
EN644 | 20190822 | L10 | 13 | 9 |  |  | a | 281.05 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 10.12 | 7.021 | 0.79233 | 9.3277 | 6.2287 | 0.060765 | 0.045269 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 105
EN644 | 20190822 | L10 | 13 | 9 |  |  | b | 280.43 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 11.37 | 7.923 | 0.79233 | 10.5777 | 7.1307 | 0.067738 | 0.05392 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 106
EN644 | 20190822 | L10 | 13 | 5 |  |  | a | 280.02 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.1 | 1.888 | 0.79233 | 1.3077 | 1.0957 | 0.0041722 | 0.014549 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 101 | poor chl replication value (63% range) except within noise of fluorometer and basically value of zero; Taylor previously QC'ed
EN644 | 20190822 | L10 | 13 | 5 |  |  | b | 279.89 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.363 | 1.955 | 0.79233 | 1.5707 | 1.1627 | 0.0080332 | 0.011842 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 102 | poor chl replication value (63% range) except within noise of fluorometer and basically value of zero; Taylor previously QC'ed
EN644 | 20190822 | L10 | 13 | 14 |  |  | a | 279.67 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 5.091 | 3.404 | 0.79233 | 4.2987 | 2.6117 | 0.033242 | 0.011438 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 103 | Poor replication (91% range) and this value lower than the >20, "Bottles were re-arranged prior to recording-- may or may not be correct volume associated with the tube." are vol filt uneven and inaccuratelly recorded?; Taylor previously QC'ed
EN644 | 20190822 | L10 | 13 | 14 |  |  | b | 279.76 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 15.86 | 11.35 | 0.79233 | 15.0677 | 10.5577 | 0.088839 | 0.091718 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 104 | Poor replication (91% range), Bottles were re-arranged prior to recording-- may or may not be correct volume associated with the tube. are vol filt uneven and inaccuratelly recorded?; Taylor previously QC'ed
EN644 | 20190822 | L10 | 13 | 14 |  |  | a | 281.08 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.661 | 1.16 | 0.79233 | 0.86867 | 0.36767 | 0.0098225 | -0.0035642 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 107 | Poor replication, negative phaeo, questionable volume, toss sample. Bottles were re-arranged prior to recording-- may or may not be correct volume associated with the tube. are vol filt uneven and inaccuratelly recorded?; Taylor previously QC'ed
EN644 | 20190822 | L10 | 13 | 14 |  |  | b | 280.17 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 9.74 | 7.408 | 0.79233 | 8.9477 | 6.6157 | 0.045869 | 0.067107 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 108 | Poor replication, questionable volume filtered. Bottles were re-arranged prior to recording-- may or may not be correct volume associated with the tube. are vol filt uneven and inaccuratelly recorded?; Taylor previously QC'ed, value larger than one of the >5 reps
EN644 | 20190822 | L10 | 13 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 17.84 | 11.61 | 0.72767 | 17.1123 | 10.8823 | 0.12852 | 0.05878 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L10 | 13 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 18.69 | 11.9 | 0.72767 | 17.9623 | 11.1723 | 0.14007 | 0.052219 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L10 | 13 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 21.67 | 13.55 | 0.73967 | 20.9303 | 12.8103 | 0.1675 | 0.052974 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L10 | 13 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 19.43 | 12.68 | 0.73967 | 18.6903 | 11.9403 | 0.13924 | 0.066262 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L10 | 13 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 36.36 | 23.94 | 1.2035 | 35.1565 | 22.7365 | 0.25621 | 0.13511 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L10 | 13 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 34.32 | 22.13 | 0.73967 | 33.5803 | 21.3903 | 0.25146 | 0.11669 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L10 | 13 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 129.7 | 88.97 | 0.80267 | 128.8973 | 88.1673 | 0.8402 | 0.67724 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L10 | 13 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 122.9 | 85.71 | 0.80267 | 122.0973 | 84.9073 | 0.76717 | 0.69416 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L10 | 13 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 24.67 | 9.337 | 1.2035 | 23.4665 | 8.1335 | 0.3163 | -0.17631 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | negative phaeo and poor replication, how does it compare to CTD fluor? Is there a transcribing error?
EN644 | 20190822 | L10 | 13 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 10.65 | 7.839 | 0.73967 | 9.9103 | 7.0993 | 0.057987 | 0.064199 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | very poor match to replicate, this was run out of order in lab, could it be a typo in lab of what cast?
EN644 | 20190822 | L10 | 13 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 8.548 | 3.843 | 1.2035 | 7.3445 | 2.6395 | 0.097057 | -0.051629 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | looks very clear that C13N3a and C13N2a were mixed up in tissue capsules. These values originally from tiss cap labeled C13N2a; neg phaeo, QF was 1 but DF changed to 2
EN644 | 20190822 | L10 | 13 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 8.302 | 3.831 | 1.2035 | 7.0985 | 2.6275 | 0.09223 | -0.047008 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | neg phaeo, QF was 1 but DF changed to 2
EN644 | 20190822 | L10 | 13 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 2.88 | 2.649 | 0.80267 | 2.0773 | 1.8463 | 0.0047652 | 0.027012 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | looks very clear that C13N3a and C13N2a were mixed up in tissue capsules so switch values. These values from a tiss capsule labeled C13N3a
EN644 | 20190822 | L10 | 13 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.718 | 3.3 | 0.80267 | 2.9153 | 2.4973 | 0.0086227 | 0.034359 | 20201113 | 10-81 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190822 | L10 | 13 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 15.29 | 10.12 | 1.2035 | 14.0865 | 8.9165 | 0.10665 | 0.046812 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L10 | 13 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 15 | 9.683 | 1.2035 | 13.7965 | 8.4795 | 0.10968 | 0.036258 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190822 | L10 | 13 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 16.17 | 10.85 | 0.73967 | 15.4303 | 10.1103 | 0.10974 | 0.064265 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190822 | L10 | 13 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 19.17 | 12.39 | 1.2035 | 17.9665 | 11.1865 | 0.13986 | 0.052669 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190822 | L10 | 13 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 110.3 | 74.47 | 0.73967 | 109.5603 | 73.7303 | 0.73912 | 0.52985 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | tissue capsule labeled C13N9 but accurate that sample is from N8, see cruise log sheet; high range of replicates
EN644 | 20190822 | L10 | 13 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 118.5 | 69.89 | 1.2035 | 117.2965 | 68.6865 | 1.0027 | 0.17941 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | tissue capsule labeled C13N9 but accurate that sample is from N8, see cruise log sheet; high range of replicates
EN644 | 20190823 | L11 | 17 | 18 |  |  | a | 280.66 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.193 | 1.737 | 0.79233 | 1.4007 | 0.94467 | 0.0089536 | 0.0071503 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 109 | poor phaeo replication (45% range) but values basically within noise of fluorometer and actually good chl agreement; Taylor previously QC'ed
EN644 | 20190823 | L11 | 17 | 18 |  |  | b | 280.16 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.307 | 1.923 | 0.79233 | 1.5147 | 1.1307 | 0.0075533 | 0.011756 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 110 | poor phaeo replication (45% range) but values basically within noise of fluorometer and actually good chl agreement; Taylor previously QC'ed
EN644 | 20190823 | L11 | 17 | 10 |  |  | a | 280.09 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 11.79 | 8.939 | 0.79233 | 10.9977 | 8.1467 | 0.056094 | 0.083067 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 111 | somewhat poor chl replication (30% range) but low values; Taylor previously QC'ed
EN644 | 20190823 | L11 | 17 | 10 |  |  | b | 281.05 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 14.82 | 10.91 | 0.79233 | 14.0277 | 10.1177 | 0.076667 | 0.095572 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 112 | somewhat poor chl replication (30% range) but low values; Taylor previously QC'ed
EN644 | 20190823 | L11 | 17 | 7 |  |  | a | 280.02 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 10.19 | 7.791 | 0.79233 | 9.3977 | 6.9987 | 0.047212 | 0.072368 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 113 | somewhat poor chl replication (36% range) but low values; Taylor previously QC'ed
EN644 | 20190823 | L11 | 17 | 7 |  |  | b | 280.59 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 8.641 | 6.974 | 0.79233 | 7.8487 | 6.1817 | 0.03274 | 0.072666 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 114 | somewhat poor chl replication (36% range) but low values; Taylor previously QC'ed
EN644 | 20190823 | L11 | 17 | 18 |  |  | a | 280.31 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.73 | 1.21 | 0.79233 | 0.93767 | 0.41767 | 0.010223 | -0.0030941 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 115 | negative phaeo so toss, and in noise of fluorometer; Taylor previously QC'ed
EN644 | 20190823 | L11 | 17 | 18 |  |  | b | 280.47 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.648 | 2.043 | 0.79233 | 1.8557 | 1.2507 | 0.011887 | 0.0094475 | 20181101 | 10-25 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 116 | larger than >5 values but basically a value of zero anwyays; Taylor previously QC'ed, QF was 1
EN644 | 20190823 | L11 | 17 | 10 |  |  | a | 279.91 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 4.454 | 3.091 | 0.79233 | 3.6617 | 2.2987 | 0.026834 | 0.012456 | 20181101 | 10-26 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 117
EN644 | 20190823 | L11 | 17 | 10 |  |  | b | 280.43 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 4.418 | 3.134 | 0.79233 | 3.6257 | 2.3417 | 0.025232 | 0.014719 | 20181101 | 10-26 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 118
EN644 | 20190823 | L11 | 17 | 7 |  |  | a | 279.94 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.125 | 4.901 | 0.79233 | 5.3327 | 4.1087 | 0.024095 | 0.046126 | 20181101 | 10-26 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 119
EN644 | 20190823 | L11 | 17 | 7 |  |  | b | 279.89 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 5.486 | 4.418 | 0.79233 | 4.6937 | 3.6257 | 0.021028 | 0.040949 | 20181101 | 10-26 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 120
EN644 | 20190823 | L11 | 17 | 14 |  |  | a | 280.17 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 3.578 | 2.848 | 0.79233 | 2.7857 | 2.0557 | 0.014359 | 0.020746 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 121 | poor phaeo replication (65% range) but very low values in the noise and good chl replication; Taylor previously QC'ed
EN644 | 20190823 | L11 | 17 | 14 |  |  | b | 279.67 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 2.86 | 2.18 | 0.79233 | 2.0677 | 1.3877 | 0.013399 | 0.01034 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 122 | poor phaeo replication (65% range) but very low values in the noise and good chl replication; Taylor previously QC'ed
EN644 | 20190823 | L11 | 17 | 14 |  |  | a | 279.76 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 6.831 | 5.538 | 0.79233 | 6.0387 | 4.7457 | 0.02547 | 0.055691 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 123 | poor replicaiton, signal very low within noise of fluorometer but value double what the >5 values are, did all volume get poured in 1 funnel?; Taylor previously QC'ed, DF thinks possible all volume went onto this filter
EN644 | 20190823 | L11 | 17 | 14 |  |  | b | 281.06 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 1.647 | 1.169 | 0.79233 | 0.85467 | 0.37667 | 0.0093722 | -0.0029602 | 20181101 | 10-23 |  | 1 | MJM | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 124 | poor replication and phaeo negative so toss, did all vol get poured in the other replicate funnel?; Taylor previously QC'ed, DF thinks all volume went onto rep a filter
EN644 | 20190823 | L11 | 17 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 13.19 | 8.491 | 1.2035 | 11.9865 | 7.2875 | 0.096933 | 0.028491 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190823 | L11 | 17 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.94 | 7.723 | 0.73967 | 11.2003 | 6.9833 | 0.08699 | 0.033199 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190823 | L11 | 17 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 17.88 | 12.08 | 1.2035 | 16.6765 | 10.8765 | 0.11965 | 0.067549 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190823 | L11 | 17 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 17.73 | 11.51 | 0.73967 | 16.9903 | 10.7703 | 0.12831 | 0.057058 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190823 | L11 | 17 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 88.19 | 63.12 | 1.2035 | 86.9865 | 61.9165 | 0.51716 | 0.54848 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190823 | L11 | 17 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 89 | 60.67 | 0.73967 | 88.2603 | 59.9303 | 0.5844 | 0.44705 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190823 | L11 | 17 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 38.59 | 29.05 | 0.73967 | 37.8503 | 28.3103 | 0.1968 | 0.29045 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190823 | L11 | 17 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 38.89 | 30.21 | 0.73967 | 38.1503 | 29.4703 | 0.17906 | 0.32816 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190823 | L11 | 17 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 26.27 | 19.77 | 0.73967 | 25.5303 | 19.0303 | 0.13409 | 0.19344 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190823 | L11 | 17 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 23.3 | 17.24 | 0.73967 | 22.5603 | 16.5003 | 0.12501 | 0.15898 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190823 | L11 | 17 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.023 | 3.267 | 0.73967 | 3.2833 | 2.5273 | 0.015595 | 0.027903 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190823 | L11 | 17 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.91 | 3.083 | 0.73967 | 3.1703 | 2.3433 | 0.01706 | 0.023271 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190823 | L11 | 17 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.012 | 2.79 | 0.73967 | 2.2723 | 2.0503 | 0.0045795 | 0.030709 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication but within noise of fluorometer and basically values of zero
EN644 | 20190823 | L11 | 17 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.27 | 2.472 | 1.2035 | 2.0665 | 1.2685 | 0.016462 | 0.0053705 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication but within noise of fluorometer and basically values of zero
EN644 | 20190823 | L11 | 17 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.091 | 2.68 | 0.72767 | 2.3633 | 1.9523 | 0.0084783 | 0.025123 | 20201113 | 11-00 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high chl z-score but chl basically zero so any differences result in high shore
EN644 | 20190823 | L11 | 17 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 2.035 | 2.012 | 0.73967 | 1.2953 | 1.2723 | 0.00047445 | 0.021424 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high chl z-score but chl basically zero so any differences result in high shore
EN644 | 20190823 | L11 | 17 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 14.92 | 9.652 | 0.73967 | 14.1803 | 8.9123 | 0.10867 | 0.044719 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190823 | L11 | 17 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 13.59 | 8.718 | 0.73967 | 12.8503 | 7.9783 | 0.1005 | 0.036812 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190824 | L5 | 22 | 20 |  |  | a | 280.66 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 178.9 | 107.5 | 0.79233 | 178.1077 | 106.7077 | 1.4019 | 0.41712 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 137 | somewhat poor phaeo replication (38% range); Taylor previously QC'ed; variation is within bio variation
EN644 | 20190824 | L5 | 22 | 20 |  |  | b | 280.16 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 207.4 | 128.6 | 0.79233 | 206.6077 | 127.8077 | 1.55 | 0.63264 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 138 | somewhat poor phaeo replication (38% range); Taylor previously QC'ed, value larger than both of the WSW reps, QF was 1; within bio variation and seems like all chl signal is just >5 hence the close value to WSW…DF changed this QF back to 1
EN644 | 20190824 | L5 | 22 | 14 |  |  | a | 280.09 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 233.5 | 144.6 | 0.79233 | 232.7077 | 143.8077 | 1.7491 | 0.70739 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 139
EN644 | 20190824 | L5 | 22 | 14 |  |  | b | 280.43 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 237.2 | 147.8 | 0.79233 | 236.4077 | 147.0077 | 1.7568 | 0.75131 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 140
EN644 | 20190824 | L5 | 22 | 9 |  |  | a | 279.89 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 72.29 | 45.5 | 0.79233 | 71.4977 | 44.7077 | 0.52747 | 0.23677 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 141
EN644 | 20190824 | L5 | 22 | 9 |  |  | b | 280.59 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 76.75 | 47.65 | 0.79233 | 75.9577 | 46.8577 | 0.57152 | 0.22747 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 142
EN644 | 20190824 | L5 | 22 | 20 |  |  | a | 280.47 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 122.8 | 73.98 | 0.79233 | 122.0077 | 73.1877 | 0.95923 | 0.28925 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 143
EN644 | 20190824 | L5 | 22 | 20 |  |  | b | 280.31 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 130.7 | 76.62 | 0.79233 | 129.9077 | 75.8277 | 1.0632 | 0.23107 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 144
EN644 | 20190824 | L5 | 22 | 14 |  |  | a | 281.05 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 246.6 | 145.7 | 0.79233 | 245.8077 | 144.9077 | 1.9784 | 0.48841 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 145 | value larger than both >5 values, QF was 1
EN644 | 20190824 | L5 | 22 | 14 |  |  | b | 279.91 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 211 | 124.3 | 0.79233 | 210.2077 | 123.5077 | 1.7069 | 0.40418 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 146
EN644 | 20190824 | L5 | 22 | 9 |  |  | a | 279.94 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 56.19 | 34.34 | 0.79233 | 55.3977 | 33.5477 | 0.43013 | 0.14323 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 147
EN644 | 20190824 | L5 | 22 | 9 |  |  | b | 280.02 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 53.85 | 32.75 | 0.79233 | 53.0577 | 31.9577 | 0.41525 | 0.13079 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 148
EN644 | 20190824 | L5 | 22 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 226.4 | 151.4 | 0.73967 | 225.6603 | 150.6603 | 1.5471 | 1.0459 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high phaeo range
EN644 | 20190824 | L5 | 22 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 216.3 | 136.7 | 0.73967 | 215.5603 | 135.9603 | 1.642 | 0.69798 | 20201113 | 10-83 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high phaeo range
EN644 | 20190824 | L5 | 22 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 292.6 | 183.6 | 0.73967 | 291.8603 | 182.8603 | 2.2485 | 0.89869 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190824 | L5 | 22 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 318.8 | 185.1 | 1.2035 | 317.5965 | 183.8965 | 2.758 | 0.407 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190824 | L5 | 22 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 116.4 | 75.3 | 0.73967 | 115.6603 | 74.5603 | 0.84783 | 0.43542 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high phaeo range
EN644 | 20190824 | L5 | 22 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 105.4 | 64.82 | 1.2035 | 104.1965 | 63.6165 | 0.8371 | 0.25779 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high phaeo range
EN644 | 20190824 | L5 | 22 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 43.36 | 30.06 | 0.73967 | 42.6203 | 29.3203 | 0.27436 | 0.23027 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor chl replication
EN644 | 20190824 | L5 | 22 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 50.81 | 33.17 | 1.2035 | 49.6065 | 31.9665 | 0.36389 | 0.18629 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor chl replication
EN644 | 20190824 | L5 | 22 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 28.22 | 19.75 | 1.2035 | 27.0165 | 18.5465 | 0.17472 | 0.14448 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor chl replication
EN644 | 20190824 | L5 | 22 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 24.41 | 18.15 | 0.72767 | 23.6823 | 17.4223 | 0.12913 | 0.17072 | 20201113 | 11-01 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | somewhat poor chl replication
EN644 | 20190824 | L5 | 22 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 36.17 | 30.72 | 0.73967 | 35.4303 | 29.9803 | 0.11243 | 0.40356 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190824 | L5 | 22 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 37.45 | 31.52 | 0.73967 | 36.7103 | 30.7803 | 0.12233 | 0.40743 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190824 | L5 | 22 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 18.2 | 11.7 | 1.2035 | 16.9965 | 10.4965 | 0.13409 | 0.046569 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190824 | L5 | 22 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 19.01 | 12.16 | 1.2035 | 17.8065 | 10.9565 | 0.14131 | 0.047266 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190824 | L4 | 23 | 18 |  |  | a | 280.31 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 341.2 | 223.1 | 0.79233 | 340.4077 | 222.3077 | 2.3218 | 1.4727 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 149
EN644 | 20190824 | L4 | 23 | 18 |  |  | b | 280.66 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 368.8 | 242.3 | 0.79233 | 368.0077 | 241.5077 | 2.4838 | 1.6332 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 150
EN644 | 20190824 | L4 | 23 | 13 |  |  | a | 281.05 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 482.9 | 322.2 | 0.79233 | 482.1077 | 321.4077 | 3.151 | 2.3205 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 151 | both samples diluted but non-dilution values matchup better; Looks like Taylor previously QC'ed but QC_d was 0; vale larger than one of the WSW reps
EN644 | 20190824 | L4 | 23 | 13 |  |  | b | 280.09 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 484.5 | 309.9 | 0.79233 | 483.7077 | 309.1077 | 3.4353 | 1.8449 | 20181101 | 10-26 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 152 | both samples diluted but non-dilution values matchup better; Looks like Taylor previously QC'ed but QC_d was 0; value larger than one of the WSW rep
EN644 | 20190824 | L4 | 23 | 5 |  |  | a | 280.02 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 76.77 | 52.25 | 0.79233 | 75.9777 | 51.4577 | 0.48255 | 0.39666 | 20181101 | 10-26 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 153 | poor replication (60% chl range); Taylor previously QC'ed
EN644 | 20190824 | L4 | 23 | 5 |  |  | b | 279.89 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 45.93 | 32.76 | 0.79233 | 45.1377 | 31.9677 | 0.25931 | 0.28715 | 20181101 | 10-26 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 154 | poor replication (60% chl range); Taylor previously QC'ed, value smaller than one of >20 rep values
EN644 | 20190824 | L4 | 23 | 10 |  |  | a | 281.08 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 182.2 | 116.7 | 0.79233 | 181.4077 | 115.9077 | 1.2842 | 0.68877 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 155 | good replication, values very close to >20 reps
EN644 | 20190824 | L4 | 23 | 10 |  |  | b | 280.17 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 191.8 | 123.7 | 0.79233 | 191.0077 | 122.9077 | 1.3395 | 0.75941 | 20181101 | 10-23 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 156 | good replication, values very close to >20 reps
EN644 | 20190824 | L4 | 23 | 18 |  |  | a | 280.47 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 264 | 170.7 | 0.79233 | 263.2077 | 169.9077 | 1.8332 | 1.0652 | 20181101 | 10-24 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 157
EN644 | 20190824 | L4 | 23 | 18 |  |  | b | 280.16 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 267.3 | 172 | 0.79233 | 266.5077 | 171.2077 | 1.8746 | 1.0493 | 20181101 | 10-24 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 158
EN644 | 20190824 | L4 | 23 | 13 |  |  | a | 280.43 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 372.1 | 242 | 0.79233 | 371.3077 | 241.2077 | 2.5566 | 1.5587 | 20181101 | 10-24 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 159
EN644 | 20190824 | L4 | 23 | 13 |  |  | b | 279.91 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 373 | 235.7 | 0.79233 | 372.2077 | 234.9077 | 2.7031 | 1.3121 | 20181101 | 10-24 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 160
EN644 | 20190824 | L4 | 23 | 5 |  |  | a | 280.59 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 40.94 | 27.1 | 0.79233 | 40.1477 | 26.3077 | 0.27182 | 0.17677 | 20181101 | 10-24 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 161 | Taylor previously QC'ed
EN644 | 20190824 | L4 | 23 | 5 |  |  | b | 279.94 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 35.13 | 23.84 | 0.79233 | 34.3377 | 23.0477 | 0.22225 | 0.17166 | 20181101 | 10-24 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 162 | Taylor previously QC'ed
EN644 | 20190824 | L4 | 23 | 10 |  |  | a | 279.67 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 166.4 | 101.6 | 0.79233 | 165.6077 | 100.8077 | 1.2769 | 0.44771 | 20181101 | 10-24 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 163 | good replication, values very close to >5 reps
EN644 | 20190824 | L4 | 23 | 10 |  |  | b | 279.76 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 177.2 | 109.3 | 0.79233 | 176.4077 | 108.5077 | 1.3375 | 0.51819 | 20181101 | 10-24 |  | 1 | Erin | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 164 | good replication, values very close to >5 reps
EN644 | 20190824 | L4 | 23 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 409.4 | 255.2 | 1.2035 | 408.1965 | 253.9965 | 3.1809 | 1.1906 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190824 | L4 | 23 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 394.9 | 251.2 | 0.73967 | 394.1603 | 250.4603 | 2.9643 | 1.3463 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190824 | L4 | 23 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 407.7 | 216.2 | 1.2035 | 406.4965 | 214.9965 | 3.9504 | -0.25007 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | could this Ra be a writtrn typo and should be 316.2 it was originally a 3 but overwwritten? Very poor replication.
EN644 | 20190824 | L4 | 23 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 426.8 | 343.7 | 1.2035 | 425.5965 | 342.4965 | 1.7142 | 4.1805 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | very poor replication. Lab NB originally has Rb=526.8 but then overwritten with 426.8 could the 5 been correct?
EN644 | 20190824 | L4 | 23 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 230.4 | 145.2 | 1.2035 | 229.1965 | 143.9965 | 1.7575 | 0.72077 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190824 | L4 | 23 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 227.3 | 142.6 | 0.73967 | 226.5603 | 141.8603 | 1.7472 | 0.69431 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190824 | L4 | 23 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 109.4 | 73.61 | 1.2035 | 108.1965 | 72.4065 | 0.73829 | 0.50789 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190824 | L4 | 23 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 94.12 | 65.16 | 1.2035 | 92.9165 | 63.9565 | 0.5974 | 0.50335 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190824 | L4 | 23 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 27.36 | 21.07 | 1.2035 | 26.1565 | 19.8665 | 0.12975 | 0.21217 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190824 | L4 | 23 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 23.03 | 18.23 | 0.73967 | 22.2903 | 17.4903 | 0.099017 | 0.20201 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190824 | L4 | 23 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 25.04 | 20.67 | 0.73967 | 24.3003 | 19.9303 | 0.090146 | 0.25287 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190824 | L4 | 23 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 25.49 | 20.78 | 1.2035 | 24.2865 | 19.5765 | 0.09716 | 0.23977 | 20201113 | 10-79 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190824 | L4 | 23 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 27.07 | 17.77 | 0.73967 | 26.3303 | 17.0303 | 0.19184 | 0.10126 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190824 | L4 | 23 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 26.57 | 17.62 | 0.73967 | 25.8303 | 16.8803 | 0.18462 | 0.1059 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190824 | L4 | 23 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 98.77 | 60.94 | 0.73967 | 98.0303 | 60.2003 | 0.78037 | 0.25573 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | tissue capsule and labNB originally had values written as C23N10<10b but swapping made excellent match. Likely filters got mixed up
EN644 | 20190824 | L4 | 23 | 13 |  |  | b | 280 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 309.4 | 196.3 | 1.2035 | 308.1965 | 195.0965 | 2.4164 | 1.0613 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | where did this high value come from? No 10 filter? Mixed up sample? Looks like a match for WSW values. Rb originally had written 409 but that definitely makes no sense
EN644 | 20190824 | L4 | 23 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 53.19 | 35.52 | 0.73967 | 52.4503 | 34.7803 | 0.36451 | 0.2341 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN644 | 20190824 | L4 | 23 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 54.09 | 36.08 | 0.73967 | 53.3503 | 35.3403 | 0.37152 | 0.23672 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | tissue cpasule and labNB originally had Rb 54 written as C23N13<10a but switching gave almost perfect match with N10's so mfilter must have been msxed up
EN644 | 20190825 | MVCO | 29 | 8 |  |  | a | 280.31 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 83.04 | 53.15 | 0.79233 | 82.2477 | 52.3577 | 0.58763 | 0.30604 | 20181101 | 10-26 |  | 1 | MJM/EDJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 165 | values smaller than >20 reps, QF was 1; DF thinks all >5 and >20 reps were switched for this cast
EN644 | 20190825 | MVCO | 29 | 8 |  |  | b | 280.47 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 77.94 | 50.19 | 0.79233 | 77.1477 | 49.3977 | 0.54524 | 0.29742 | 20181101 | 10-26 |  | 1 | MJM/EDJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 166 | values smaller than >20 reps, QF was 1; DF thinks all >5 and >20 reps were switched for this cast
EN644 | 20190825 | MVCO | 29 | 4 |  |  | a | 280.09 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 92.71 | 59.57 | 0.79233 | 91.9177 | 58.7777 | 0.65203 | 0.352 | 20181101 | 10-26 |  | 1 | MJM/EDJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 167 | values smaller than >20 reps, QF was 1; DF thinks all >5 and >20 reps were switched for this cast
EN644 | 20190825 | MVCO | 29 | 4 |  |  | b | 280.43 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 95.28 | 61.98 | 0.79233 | 94.4877 | 61.1877 | 0.65438 | 0.38955 | 20181101 | 10-26 |  | 1 | MJM/EDJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 168 | values smaller than >20 reps, QF was 1; DF thinks all >5 and >20 reps were switched for this cast
EN644 | 20190825 | MVCO | 29 | 2 |  |  | a | 279.89 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 115.8 | 75.22 | 0.79233 | 115.0077 | 74.4277 | 0.79898 | 0.47329 | 20181101 | 10-26 |  | 1 | MJM/EDJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 169 | values smaller than >20 reps, QF was 1; DF thinks all >5 and >20 reps were switched for this cast
EN644 | 20190825 | MVCO | 29 | 2 |  |  | b | 279.94 | 5 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 136.6 | 87.2 | 0.79233 | 135.8077 | 86.4077 | 0.97247 | 0.50433 | 20181101 | 10-26 |  | 1 | MJM/EDJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 170 | values smaller than >20 reps, QF was 1; DF thinks all >5 and >20 reps were switched for this cast
EN644 | 20190825 | MVCO | 29 | 8 |  |  | a | 280.66 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 112.5 | 73.15 | 0.79233 | 111.7077 | 72.3577 | 0.77264 | 0.46085 | 20181101 | 10-24 |  | 1 | MJM/EDJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 171 | values larger than >5 reps, QF was 1; DF thinks all >5 and >20 reps were switched for this cast
EN644 | 20190825 | MVCO | 29 | 8 |  |  | b | 280.16 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 123.9 | 82.23 | 0.79233 | 123.1077 | 81.4377 | 0.81965 | 0.57111 | 20181101 | 10-24 |  | 1 | MJM/EDJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 172 | values larger than >5 reps, QF was 1; DF thinks all >5 and >20 reps were switched for this cast
EN644 | 20190825 | MVCO | 29 | 4 |  |  | a | 281.05 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 135.3 | 88.37 | 0.79233 | 134.5077 | 87.5777 | 0.9202 | 0.57068 | 20181101 | 10-24 |  | 1 | MJM/EDJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 173 | values larger than >5 reps, QF was 1; DF thinks all >5 and >20 reps were switched for this cast
EN644 | 20190825 | MVCO | 29 | 4 |  |  | b | 279.91 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 137.4 | 88.32 | 0.79233 | 136.6077 | 87.5277 | 0.96627 | 0.52983 | 20181101 | 10-24 |  | 1 | MJM/EDJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 174 | values larger than >5 reps, QF was 1; DF thinks all >5 and >20 reps were switched for this cast
EN644 | 20190825 | MVCO | 29 | 2 |  |  | a | 280.02 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 170.3 | 111.9 | 0.79233 | 169.5077 | 111.1077 | 1.1493 | 0.74909 | 20181101 | 10-24 |  | 1 | MJM/EDJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 175 | values larger than >5 reps, QF was 1; DF thinks all >5 and >20 reps were switched for this cast
EN644 | 20190825 | MVCO | 29 | 2 |  |  | b | 280.59 | 20 | 5 | 1 | 0 | 1 | Cal_110119 | 1.8682 | 0.5122 | 182 | 118.8 | 0.79233 | 181.2077 | 118.0077 | 1.2412 | 0.77095 | 20181101 | 10-24 |  | 1 | MJM/EDJ | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 176 | values larger than >5 reps, QF was 1; DF thinks all >5 and >20 reps were switched for this cast
EN644 | 20190825 | MVCO | 29 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 191.9 | 122.7 | 0.73967 | 191.1603 | 121.9603 | 1.4275 | 0.67156 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190825 | MVCO | 29 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 197.7 | 121.3 | 1.2035 | 196.4965 | 120.0965 | 1.576 | 0.49096 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190825 | MVCO | 29 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 223.1 | 135.4 | 1.2035 | 221.8965 | 134.1965 | 1.8091 | 0.50053 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190825 | MVCO | 29 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 205 | 131.8 | 0.73967 | 204.2603 | 131.0603 | 1.51 | 0.74566 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190825 | MVCO | 29 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 296.4 | 184.5 | 1.2035 | 295.1965 | 183.2965 | 2.3083 | 0.84637 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190825 | MVCO | 29 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 301.8 | 190.2 | 1.2035 | 300.5965 | 188.9965 | 2.3021 | 0.95066 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190825 | MVCO | 29 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 71.31 | 47.39 | 0.73967 | 70.5703 | 46.6503 | 0.49343 | 0.30946 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication; just high variance? Consistent phaeo
EN644 | 20190825 | MVCO | 29 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 84.96 | 55.28 | 0.73967 | 84.2203 | 54.5403 | 0.61225 | 0.32644 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | poor chl replication; just high variance? Consistent phaeo
EN644 | 20190825 | MVCO | 29 | 4 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 80.5 | 51.52 | 1.2035 | 79.2965 | 50.3165 | 0.59781 | 0.26818 | 20201113 | 10-80 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication; just high variance? Consistent phaeo
EN644 | 20190825 | MVCO | 29 | 4 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 63.3 | 42.21 | 0.73967 | 62.5603 | 41.4703 | 0.43505 | 0.27869 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | poor chl replication; just high variance? Consistent phaeo
EN644 | 20190825 | MVCO | 29 | 2 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 86.41 | 57.34 | 0.73967 | 85.6703 | 56.6003 | 0.59967 | 0.37447 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN644 | 20190825 | MVCO | 29 | 2 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 87.27 | 58.78 | 0.73967 | 86.5303 | 58.0403 | 0.58771 | 0.41122 | 20201113 | 10-82 |  | 1 |  | TH | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR38 | 20190920 | L3 | B1 |  |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 105.8 | 68.4 | 0.46683 | 105.3332 | 67.9332 | 0.79454 | 0.41521 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3 | B1 |  |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 107.1 | 67.43 | 0.719 | 106.381 | 66.711 | 0.84276 | 0.34522 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3 | B1 |  |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 42.4 | 28.8 | 0.719 | 41.681 | 28.081 | 0.28892 | 0.21114 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3 | B1 |  |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 37.74 | 25.65 | 0.55533 | 37.1847 | 25.0947 | 0.25684 | 0.19004 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3 | 3 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 140.1 | 105.7 | 0.46683 | 139.6332 | 105.2332 | 0.7308 | 1.1432 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | phaeo>chl
AR38 | 20190920 | L3 | 3 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 137.9 | 104.4 | 0.47433 | 137.4257 | 103.9257 | 0.71169 | 1.139 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | phaeo>chl
AR38 | 20190920 | L3 | 3 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 219.5 | 144 | 0.7165 | 218.7835 | 143.2835 | 1.6039 | 0.94763 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3 | 3 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 216.1 | 144 | 0.7165 | 215.3835 | 143.2835 | 1.5317 | 1.0199 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3 | 3 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 201.8 | 129.8 | 0.47433 | 201.3257 | 129.3257 | 1.5296 | 0.77342 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3 | 3 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 195.3 | 127.9 | 0.7165 | 194.5835 | 127.1835 | 1.4319 | 0.833 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3 | 3 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 121.9 | 79.81 | 0.47433 | 121.4257 | 79.3357 | 0.89417 | 0.51863 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3 | 3 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 108.2 | 70.82 | 0.719 | 107.481 | 70.101 | 0.79411 | 0.45424 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3 | 3 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 86.64 | 56.22 | 0.47433 | 86.1657 | 55.7457 | 0.64625 | 0.34646 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3 | 3 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 87.32 | 56.03 | 0.46683 | 86.8532 | 55.5632 | 0.66474 | 0.32473 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3 | 3 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.65 | 24.92 | 0.55533 | 36.0947 | 24.3647 | 0.2492 | 0.18469 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3 | 3 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 35.72 | 24.26 | 0.55533 | 35.1647 | 23.7047 | 0.24346 | 0.17867 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3.5 | 4 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 102.3 | 72.73 | 0.719 | 101.581 | 72.011 | 0.62819 | 0.65417 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phaeo >chl; depth only 55m
AR38 | 20190920 | L3.5 | 4 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 101.3 | 71.97 | 0.7165 | 100.5835 | 71.2535 | 0.6231 | 0.64578 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phaeo >chl; depth only 55m
AR38 | 20190920 | L3.5 | 4 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 244.5 | 164.4 | 0.7165 | 243.7835 | 163.6835 | 1.7017 | 1.2132 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3.5 | 4 | 4 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR38 | 20190920 | L3.5 | 4 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 355 | 223.2 | 0.7165 | 354.2835 | 222.4835 | 2.8 | 1.162 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3.5 | 4 | 6 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 196.9 | 125 | 0.47433 | 196.4257 | 124.5257 | 3.0549 | 1.3801 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3.5 | 4 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 265.1 | 166.4 | 0.7165 | 264.3835 | 165.6835 | 2.0968 | 0.85365 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3.5 | 4 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 263.8 | 166.9 | 0.47433 | 263.3257 | 166.4257 | 2.0586 | 0.90511 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3.5 | 4 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 286.8 | 181.1 | 0.46683 | 286.3332 | 180.6332 | 2.2455 | 0.97117 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3.5 | 4 | 10 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 158.1 | 98.47 | 0.55533 | 157.5447 | 97.9147 | 2.5336 | 0.9537 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3.5 | 4 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 74.14 | 50.98 | 0.47433 | 73.6657 | 50.5057 | 0.49202 | 0.40738 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190920 | L3.5 | 4 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 68.06 | 46.82 | 0.719 | 67.341 | 46.101 | 0.45123 | 0.36973 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L4 | 5 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 55.18 | 43.57 | 0.719 | 54.461 | 42.851 | 0.24665 | 0.51644 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae >chl; depth only 64m
AR38 | 20190921 | L4 | 5 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 54.48 | 42.79 | 0.46683 | 54.0132 | 42.3232 | 0.24835 | 0.50534 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae >chl; depth only 64m
AR38 | 20190921 | L4 | 5 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 53.98 | 40.46 | 0.46683 | 53.5132 | 39.9932 | 0.28722 | 0.42497 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae >chl; depth only 64m
AR38 | 20190921 | L4 | 5 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.98 | 53.89 | 0.719 | 66.261 | 53.171 | 0.27809 | 0.66877 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae >chl; depth only 64m
AR38 | 20190921 | L4 | 5 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 262.8 | 164 | 0.719 | 262.081 | 163.281 | 2.0989 | 0.80875 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L4 | 5 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 295.1 | 185.1 | 0.47433 | 294.6257 | 184.6257 | 2.3369 | 0.95091 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L4 | 5 | 9 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 234.4 | 145.6 | 0.719 | 233.681 | 144.881 | 3.773 | 1.3871 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L4 | 5 | 9 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 258.2 | 157.3 | 0.46683 | 257.7332 | 156.8332 | 4.2871 | 1.2986 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L4 | 5 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 147.4 | 90.4 | 0.719 | 146.681 | 89.681 | 3.6328 | 1.1583 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | sample was diluted; first reading was 426.4
AR38 | 20190921 | L4 | 5 | 11 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 219.3 | 133.2 | 0.46683 | 218.8332 | 132.7332 | 3.6583 | 1.0691 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L4 | 5 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 74.49 | 50.34 | 0.46683 | 74.0232 | 49.8732 | 0.51305 | 0.37508 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L4 | 5 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 78.1 | 53.4 | 0.55533 | 77.5447 | 52.8447 | 0.52473 | 0.41632 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L5 | 7 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.69 | 32.19 | 0.7165 | 39.9735 | 31.4735 | 0.18058 | 0.3799 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phaeo >chl
AR38 | 20190921 | L5 | 7 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 43.32 | 33.85 | 0.55533 | 42.7647 | 33.2947 | 0.20118 | 0.39172 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phaeo >chl
AR38 | 20190921 | L5 | 7 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 47.19 | 34.42 | 0.47433 | 46.7157 | 33.9457 | 0.27129 | 0.33321 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | suspicious phaeo>chl depth only 55m
AR38 | 20190921 | L5 | 7 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 48.79 | 36.08 | 0.80517 | 47.9848 | 35.2748 | 0.27002 | 0.35815 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | suspicious phaeo>chl depth only 55m
AR38 | 20190921 | L5 | 7 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 221.6 | 136.7 | 0.46683 | 221.1332 | 136.2332 | 1.8036 | 0.62238 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L5 | 7 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 231.8 | 143.4 | 0.55533 | 231.2447 | 142.8447 | 1.878 | 0.66576 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L5 | 7 | 9 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 233.8 | 139.3 | 0.46683 | 233.3332 | 138.8332 | 4.0152 | 0.92947 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L5 | 7 | 9 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 228.4 | 136.7 | 0.47433 | 227.9257 | 136.2257 | 3.8962 | 0.95557 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L5 | 7 | 11 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 229.8 | 140 | 0.47433 | 229.3257 | 139.5257 | 3.8155 | 1.1538 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | filter capsule reads N7C11 chlb, so we think it should have been C7N11 chlb
AR38 | 20190921 | L5 | 7 | 11 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 237.4 | 143.5 | 0.719 | 236.681 | 142.781 | 3.9897 | 1.0956 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | filter capsule reads N7C11 chlb, so we think it should have been C7N11 chlb
AR38 | 20190921 | L5 | 7 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.2 | 43.2 | 0.47433 | 63.7257 | 42.7257 | 0.44613 | 0.31472 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | filter capsule reads N7C11 chlb, so we think it should have been C7N11 chlb
AR38 | 20190921 | L5 | 7 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 72.93 | 47.98 | 0.47433 | 72.4557 | 47.5057 | 0.53005 | 0.31593 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | filter capsule reads N7C11 chlb, so we think it should have been C7N11 chlb
AR38 | 20190921 | L6 | 8 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.47 | 33.64 | 0.7165 | 39.7535 | 32.9235 | 0.1451 | 0.4412 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L6 | 8 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.82 | 33.87 | 0.719 | 40.101 | 33.151 | 0.14765 | 0.4427 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L6 | 8 | 5 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR38 | 20190921 | L6 | 8 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 35.78 | 28.29 | 0.47433 | 35.3057 | 27.8157 | 0.15912 | 0.33622 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | phaeo >chl need to check depth
AR38 | 20190921 | L6 | 8 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 47.98 | 36.05 | 0.47433 | 47.5057 | 35.5757 | 0.25344 | 0.38008 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | phaeo >chl need to check depth
AR38 | 20190921 | L6 | 8 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 35.12 | 28.43 | 0.7165 | 34.4035 | 27.7135 | 0.14212 | 0.35139 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3
AR38 | 20190921 | L6 | 8 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 35.31 | 27.66 | 0.46683 | 34.8432 | 27.1932 | 0.16252 | 0.32173 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L6 | 8 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.83 | 31.07 | 0.47433 | 40.3557 | 30.5957 | 0.20734 | 0.3375 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L6 | 8 | 12 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 123.9 | 78.41 | 0.55533 | 123.3447 | 77.8547 | 1.9328 | 0.84004 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 |  | cruise logsheets Niskins and correspondign depth does not match CTD btl file. Sample depth too questionable to use data. Flag as 3 and throw out
AR38 | 20190921 | L6 | 8 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 248.2 | 160.3 | 0.719 | 247.481 | 159.581 | 1.8674 | 0.97442 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 |  | cruise logsheets Niskins and correspondign depth does not match CTD btl file. Sample depth too questionable to use data. Flag as 3 and throw out
AR38 | 20190921 | L6 | 8 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 156.6 | 101 | 0.47433 | 156.1257 | 100.5257 | 1.1812 | 0.60896 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 |  | cruise logsheets Niskins and correspondign depth does not match CTD btl file. Sample depth too questionable to use data. Flag as 3 and throw out
AR38 | 20190921 | L6 | 8 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 173.6 | 110.7 | 0.47433 | 173.1257 | 110.2257 | 1.3363 | 0.62662 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 |  | cruise logsheets Niskins and correspondign depth does not match CTD btl file. Sample depth too questionable to use data. Flag as 3 and throw out
AR38 | 20190921 | L6 | 8 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 149.8 | 95.87 | 0.47433 | 149.3257 | 95.3957 | 1.1457 | 0.55309 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | for <10 a&b matching <10 filter unclear; CTD btl file depth does not match cruise hand written log sheets. Because this is surface feel more comfortable using these samples as surface depth but flag as 2 = questionable
AR38 | 20190921 | L6 | 8 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 143.2 | 92.52 | 0.47433 | 142.7257 | 92.0457 | 1.0767 | 0.56247 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | for <10 a&b matching <10 filter unclear; CTD btl file depth does not match cruise hand written log sheets. Because this is surface feel more comfortable using these samples as surface depth but flag as 2 = questionable
AR38 | 20190921 | L6 | 8 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.17 | 42.36 | 0.55533 | 62.6147 | 41.8047 | 0.44209 | 0.30236 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | for <10 a&b matching <10 filter unclear; CTD btl file depth does not match cruise hand written log sheets. Because this is surface feel more comfortable using these samples as surface depth but flag as 2 = questionable
AR38 | 20190921 | L6 | 8 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.38 | 33.26 | 0.719 | 48.661 | 32.541 | 0.34246 | 0.23703 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | for <10 a&b matching <10 filter unclear; CTD btl file depth does not match cruise hand written log sheets. Because this is surface feel more comfortable using these samples as surface depth but flag as 2 = questionable
AR38 | 20190921 | L7 | 9 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.25 | 14.18 | 0.55533 | 16.6947 | 13.6247 | 0.06522 | 0.17741 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L7 | 9 | 1 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR38 | 20190921 | L7 | 9 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.19 | 15.61 | 0.55533 | 18.6347 | 15.0547 | 0.076055 | 0.19204 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L7 | 9 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.11 | 15.52 | 0.719 | 18.391 | 14.801 | 0.076267 | 0.18731 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L7 | 9 | 5 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR38 | 20190921 | L7 | 9 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 39.71 | 27.39 | 0.80517 | 38.9048 | 26.5848 | 0.26173 | 0.21169 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L7 | 9 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 72.41 | 52.36 | 0.55533 | 71.8547 | 51.8047 | 0.42595 | 0.49658 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae>chl; depth 52m
AR38 | 20190921 | L7 | 9 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.29 | 50.15 | 0.55533 | 69.7347 | 49.5947 | 0.42786 | 0.45531 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae>chl; depth 52m
AR38 | 20190921 | L7 | 9 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 245.1 | 163.3 | 0.80517 | 244.2948 | 162.4948 | 1.7378 | 1.1559 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L7 | 9 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 241.1 | 155.1 | 0.7165 | 240.3835 | 154.3835 | 1.827 | 0.92223 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L7 | 9 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 139 | 92.42 | 0.80517 | 138.1948 | 91.6148 | 0.98956 | 0.6419 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L7 | 9 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 144.2 | 95.41 | 0.46683 | 143.7332 | 94.9432 | 1.0365 | 0.65422 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L7 | 9 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 231.1 | 148.4 | 0.719 | 230.381 | 147.681 | 1.7569 | 0.87298 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L7 | 9 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 240.2 | 151.5 | 0.7165 | 239.4835 | 150.7835 | 1.8844 | 0.80076 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L7 | 9 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.81 | 33.43 | 0.46683 | 45.3432 | 32.9632 | 0.263 | 0.324 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L7 | 9 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.35 | 27.07 | 0.719 | 37.631 | 26.351 | 0.23964 | 0.22962 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | note on logsheet C8N16 unclear what sample type talkin about "found on floor 1hr later"
AR38 | 20190921 | L9 | 10 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.33 | 9.261 | 0.80517 | 10.5248 | 8.4558 | 0.043955 | 0.10663 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L9 | 10 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.36 | 9.009 | 0.719 | 9.641 | 8.29 | 0.028701 | 0.11893 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L9 | 10 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.84 | 9.846 | 0.47433 | 12.3657 | 9.3717 | 0.063606 | 0.10328 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L9 | 10 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.25 | 9.008 | 0.55533 | 11.6947 | 8.4527 | 0.068874 | 0.08165 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L9 | 10 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 30.4 | 21.93 | 0.55533 | 29.8447 | 21.3747 | 0.17994 | 0.2007 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | N5 didn’t fire use N6
AR38 | 20190921 | L9 | 10 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.25 | 25.57 | 0.46683 | 32.7832 | 25.1032 | 0.16316 | 0.28388 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L9 | 10 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 181.1 | 119.4 | 0.80517 | 180.2948 | 118.5948 | 1.3108 | 0.80115 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L9 | 10 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 160.8 | 109.3 | 0.719 | 160.081 | 108.581 | 1.0941 | 0.83951 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L9 | 10 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 158.7 | 105.4 | 0.7165 | 157.9835 | 104.6835 | 1.1323 | 0.73187 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L9 | 10 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 141.3 | 92.65 | 0.47433 | 140.8257 | 92.1757 | 1.0335 | 0.60792 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L9 | 10 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.97 | 42.59 | 0.47433 | 63.4957 | 42.1157 | 0.4542 | 0.29579 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L9 | 10 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 68.19 | 45.53 | 0.47433 | 67.7157 | 45.0557 | 0.4814 | 0.32095 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L9 | 10 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.4 | 25.76 | 0.7165 | 37.6835 | 25.0435 | 0.26853 | 0.17744 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L9 | 10 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.79 | 23.08 | 0.47433 | 33.3157 | 22.6057 | 0.22753 | 0.17503 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.063 | 4.453 | 0.719 | 4.344 | 3.734 | 0.012959 | 0.053536 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.662 | 4.492 | 0.7165 | 3.9455 | 3.7755 | 0.0036115 | 0.063622 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.856 | 4.122 | 0.46683 | 4.3892 | 3.6552 | 0.015593 | 0.049497 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.843 | 3.817 | 0.55533 | 3.2877 | 3.2617 | 0.00055235 | 0.057531 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.933 | 2.45 | 0.47433 | 2.4587 | 1.9757 | 0.010261 | 0.024921 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.583 | 2.359 | 0.719 | 1.864 | 1.64 | 0.0047587 | 0.024446 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.407 | 2.201 | 0.719 | 1.688 | 1.482 | 0.0043763 | 0.022015 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.133 | 1.848 | 0.55533 | 1.5777 | 1.2927 | 0.0060546 | 0.016965 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.47 | 2.954 | 0.55533 | 2.9147 | 2.3987 | 0.010962 | 0.031753 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.864 | 2.17 | 0.7165 | 2.1475 | 1.4535 | 0.014744 | 0.01114 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.61 | 14.8 | 0.7165 | 19.8935 | 14.0835 | 0.12343 | 0.12737 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3
AR38 | 20190921 | L10 | 11 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.83 | 15.74 | 0.80517 | 20.0248 | 14.9348 | 0.10813 | 0.15782 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.63 | 39.15 | 0.7165 | 51.9135 | 38.4335 | 0.28637 | 0.39805 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3
AR38 | 20190921 | L10 | 11 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 56.69 | 42.45 | 0.80517 | 55.8848 | 41.6448 | 0.30252 | 0.43909 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 138.1 | 96.76 | 0.80517 | 137.2948 | 95.9548 | 0.87824 | 0.83051 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 153.7 | 106.9 | 0.7165 | 152.9835 | 106.1835 | 0.99423 | 0.89667 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 60.85 | 40.06 | 0.719 | 60.131 | 39.341 | 0.44167 | 0.25891 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 62.22 | 41 | 0.46683 | 61.7532 | 40.5332 | 0.4508 | 0.271 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 17 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.06 | 22.51 | 0.47433 | 32.5857 | 22.0357 | 0.22413 | 0.16828 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L10 | 11 | 17 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.94 | 25.22 | 0.46683 | 36.4732 | 24.7532 | 0.24898 | 0.19182 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.187 | 2.085 | 0.47433 | 1.7127 | 1.6107 | 0.0021669 | 0.026516 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 7 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR38 | 20190921 | L11 | 12 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.093 | 1.759 | 0.46683 | 1.6262 | 1.2922 | 0.0070956 | 0.015915 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.606 | 2.131 | 0.719 | 1.887 | 1.412 | 0.010091 | 0.015054 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 1.811 | 1.743 | 0.47433 | 1.3367 | 1.2687 | 0.0014446 | 0.021148 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.301 | 2.404 | 0.80517 | 1.4958 | 1.5988 | -0.0021882 | 0.03066 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Ra larger than Rb. Emily notified when analyzed. Value you so small that even though negative chl it's basicaly zero. Water depth 290m
AR38 | 20190921 | L11 | 12 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.455 | 1.983 | 0.46683 | 1.9882 | 1.5162 | 0.010027 | 0.016972 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.96 | 11.01 | 0.55533 | 14.4047 | 10.4547 | 0.083915 | 0.10226 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.6 | 11.3 | 0.46683 | 15.1332 | 10.8332 | 0.091351 | 0.10156 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 67.55 | 49.87 | 0.719 | 66.831 | 49.151 | 0.3756 | 0.49967 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.02 | 51.98 | 0.46683 | 69.5532 | 51.5132 | 0.38325 | 0.53409 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 87.31 | 59.38 | 0.7165 | 86.5935 | 58.6635 | 0.59335 | 0.45132 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 83.13 | 56.39 | 0.7165 | 82.4135 | 55.6735 | 0.56807 | 0.42335 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76.3 | 50.13 | 0.55533 | 75.7447 | 49.5747 | 0.55596 | 0.32685 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76.84 | 50.1 | 0.7165 | 76.1235 | 49.3835 | 0.56807 | 0.31134 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 51.53 | 35.19 | 0.55533 | 50.9747 | 34.6347 | 0.34713 | 0.26964 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190921 | L11 | 12 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 71.38 | 47.13 | 0.7165 | 70.6635 | 46.4135 | 0.51517 | 0.31135 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L4.5 | 13 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 50.42 | 39.26 | 0.719 | 49.701 | 38.541 | 0.23709 | 0.44925 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | filter touched after filtering
AR38 | 20190922 | L4.5 | 13 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 54.4 | 43.05 | 0.80517 | 53.5948 | 42.2448 | 0.24112 | 0.51117 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L4.5 | 13 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 203 | 126.4 | 0.741 | 202.259 | 125.659 | 1.6273 | 0.6104 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L4.5 | 13 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 164.9 | 104.7 | 0.46683 | 164.4332 | 104.2332 | 1.2789 | 0.57726 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L4.5 | 13 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 342.9 | 215.1 | 0.7165 | 342.1835 | 214.3835 | 2.715 | 1.1027 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L4.5 | 13 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 350.8 | 213.3 | 0.80517 | 349.9948 | 212.4948 | 2.9211 | 0.86298 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L4.5 | 13 | 8 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 208.4 | 129.3 | 0.46683 | 207.9332 | 128.8332 | 3.3609 | 1.2276 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L4.5 | 13 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 179.9 | 111.7 | 0.741 | 179.159 | 110.959 | 4.3466 | 1.5812 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L4.5 | 13 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.24 | 53.97 | 0.719 | 79.521 | 53.251 | 0.55809 | 0.3902 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L4.5 | 13 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.25 | 43.46 | 0.7165 | 65.5335 | 42.7435 | 0.48416 | 0.27701 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L4 | 14 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 54.61 | 41.84 | 0.55533 | 54.0547 | 41.2847 | 0.27129 | 0.4639 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae>chl; depth 58m
AR38 | 20190922 | L4 | 14 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 55.4 | 41.98 | 0.55533 | 54.8447 | 41.4247 | 0.2851 | 0.45259 | 20220519 | 13-11 | 20230307 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae>chl; depth 58m; there were two rep a, so I filled in one of the rep a under rep b
AR38 | 20190922 | L4 | 14 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 268.4 | 160.4 | 0.7165 | 267.6835 | 159.6835 | 6.8832 | 1.6477 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | sample was diluted; first reading was 769.8
AR38 | 20190922 | L4 | 14 | 3 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 397.7 | 240.7 | 0.47433 | 397.2257 | 240.2257 | 6.6707 | 1.8851 | 20220519 | 13-10 | 20230303 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Tissue capsule had 2 rep a and I assume it will be rep b
AR38 | 20190922 | L4 | 14 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 251.3 | 151.7 | 0.80517 | 250.4948 | 150.8948 | 2.1159 | 0.57118 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L4 | 14 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 259.9 | 157.3 | 0.741 | 259.159 | 156.559 | 2.1797 | 0.60831 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L4 | 14 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 51.15 | 35.53 | 0.741 | 50.409 | 34.789 | 0.33184 | 0.28768 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | tissue capsule was labeld C1N5b but after all samples run process of elimination and matching Rb/Ra values look like the 4 of C14 was forgotten to be written. Feel comfortable assuming this is this sample
AR38 | 20190922 | L4 | 14 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 51.61 | 34.21 | 0.80517 | 50.8048 | 33.4048 | 0.36965 | 0.22522 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | d2a | 18 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 174.6 | 129 | 0.719 | 173.881 | 128.281 | 0.96874 | 1.3157 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | d2a | 18 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 164.7 | 120.2 | 0.46683 | 164.2332 | 119.7332 | 0.94537 | 1.1868 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | d2a | 18 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 111.8 | 75.24 | 0.80517 | 110.9948 | 74.4348 | 0.77669 | 0.54883 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | d2a | 18 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 120.6 | 80.82 | 0.7165 | 119.8835 | 80.1035 | 0.8451 | 0.58137 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | d2a | 18 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 142.1 | 94.61 | 0.7165 | 141.3835 | 93.8935 | 1.0089 | 0.66315 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | d2a | 18 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 181.1 | 116 | 0.46683 | 180.6332 | 115.5332 | 1.383 | 0.67439 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | d2a | 18 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 149.6 | 91.8 | 0.80517 | 148.7948 | 90.9948 | 1.2279 | 0.3925 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | d2a | 18 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 148.2 | 91.66 | 0.80517 | 147.3948 | 90.8548 | 1.2012 | 0.41678 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | d2a | 18 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.38 | 21.76 | 0.80517 | 32.5748 | 20.9548 | 0.24686 | 0.1263 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | d2a | 18 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.87 | 22.69 | 0.7165 | 34.1535 | 21.9735 | 0.25876 | 0.13255 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L2 | 19 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 225.9 | 155.7 | 0.80517 | 225.0948 | 154.8948 | 1.4914 | 1.267 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phaeo quite high! But I guess bottom water at 40m is reasonable
AR38 | 20190922 | L2 | 19 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 243.7 | 159.9 | 0.741 | 242.959 | 159.159 | 1.7803 | 1.054 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L2 | 19 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 110.2 | 72.48 | 0.719 | 109.481 | 71.761 | 0.80134 | 0.47657 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L2 | 19 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 111.9 | 73.65 | 0.7165 | 111.1835 | 72.9335 | 0.8126 | 0.48619 | 20220519 | 13-07 | 20230223 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L2 | 19 | 5 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 199.3 | 123.2 | 0.46683 | 198.8332 | 122.7332 | 3.2334 | 1.1378 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L2 | 19 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 344.8 | 216.7 | 0.719 | 344.081 | 215.981 | 2.7214 | 1.1248 | 20220519 | 13-08 | 20230224 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L2 | 19 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 201.3 | 123 | 0.46683 | 200.8332 | 122.5332 | 1.6634 | 0.51862 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L2 | 19 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 202.8 | 127.3 | 0.80517 | 201.9948 | 126.4948 | 1.6039 | 0.64866 | 20220519 | 13-02 | 20230216 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L2 | 19 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.09 | 24.84 | 0.46683 | 37.6232 | 24.3732 | 0.28149 | 0.15255 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR38 | 20190922 | L2 | 19 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 39.93 | 25.94 | 0.46683 | 39.4632 | 25.4732 | 0.29721 | 0.15641 | 20220519 | 13-09 | 20230302 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR39A | 20190926 | L6 | 1 | 3 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33 | Cal_111320b_12mm | 1.8343 | 0.5442 | 192.9 | 120.2 | 1.02 | 191.88 | 119.18 | 2.496 | 0.91791 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 01/01; Inshore SurfM deploy
AR39A | 20190926 | L6 | 1 | 4 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33 | Cal_111320b_12mm | 1.8343 | 0.5442 | 165.6 | 110.09 | 1.0175 | 164.5825 | 109.0725 | 1.9058 | 1.2186 | 20201113 | 11-15 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 01/02 | Pre dilution values are 462.5 and 306.0
AR39A | 20190926 | L6 | 1 | 5 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 267.9 | 174.8 | 1.0175 | 266.8825 | 173.7825 | 0.52691 | 0.29369 | 20201113 | 11-15 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 01/03
AR39A | 20190926 | L6 | 1 | 6 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 199.7 | 132 | 1.0175 | 198.6825 | 130.9825 | 0.38352 | 0.23556 | 20201113 | 11-15 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 01/04
AR39A | 20190926 | L6 | 1 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 206 | 133.6 | 1.02 | 204.98 | 132.58 | 0.41015 | 0.21649 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 01/05
AR39A | 20190926 | L6 | 1 | 8 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 221.3 | 144.6 | 1.0175 | 220.2825 | 143.5825 | 0.43451 | 0.24413 | 20201113 | 11-15 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 01/06
AR39A | 20190926 | L8 | 2 | 5 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 0.33 | Cal_111320b_12mm | 1.8343 | 0.5442 | 239.5 | 159.6 | 1.0175 | 238.4825 | 158.5825 | 1.3716 | 0.8997 | 20201113 | 11-15 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 02/01; Central SurfM recover
AR39A | 20190926 | L8 | 2 | 6 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 0.33 | Cal_111320b_12mm | 1.8343 | 0.5442 | 230.5 | 154.6 | 1.0175 | 229.4825 | 153.5825 | 1.303 | 0.89676 | 20201113 | 11-15 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 02/02
AR39A | 20190926 | L8 | 2 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 138 | 92.53 | 1.0175 | 136.9825 | 91.5125 | 0.25734 | 0.17478 | 20201113 | 11-15 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 02/03
AR39A | 20190926 | L8 | 2 | 8 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 138.6 | 89.86 | 1.0175 | 137.5825 | 88.8425 | 0.27611 | 0.1438 | 20201113 | 11-15 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 02/04
AR39A | 20190927 |  | 3 | 5 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 214 | 154.8 | 1.02 | 212.98 | 153.78 | 0.33537 | 0.39147 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 03/01; Glider deploy 340&379
AR39A | 20190927 |  | 3 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 216.1 | 148.7 | 1.02 | 215.08 | 147.68 | 0.76364 | 0.63237 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 03/02
AR39A | 20190927 |  | 3 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 185.1 | 118.7 | 1.0175 | 184.0825 | 117.6825 | 0.3758 | 0.1799 | 20201113 | 11-15 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 03/03
AR39A | 20190927 |  | 3 | 8 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 171.8 | 108.9 | 1.02 | 170.78 | 107.88 | 0.35633 | 0.15356 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 03/04
AR39A | 20190927 | L8 | 4 | 5 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 0.1665 | Cal_111320b_12mm | 1.8343 | 0.5442 | 162.3 | 107.1 | 1.0175 | 161.2825 | 106.0825 | 1.8781 | 1.1333 | 20201113 | 11-16 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/01; Central Surf M deploy | First added 5ml before measuring, and then diluted .333 after that. Before second dilution values were 454.1 and 298.8
AR39A | 20190927 | L8 | 4 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33 | Cal_111320b_12mm | 1.8343 | 0.5442 | 161.8 | 109.3 | 1.0175 | 160.7825 | 108.2825 | 1.8025 | 1.2993 | 20201113 | 11-16 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/02 | Pre dilution values are 465.6 and 302.9
AR39A | 20190927 | L8 | 4 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 134.5 | 84.18 | 1.02 | 133.48 | 83.16 | 0.28479 | 0.10789 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/03
AR39A | 20190927 | L8 | 4 | 8 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 144.4 | 92.89 | 1.0175 | 143.3825 | 91.8725 | 0.2918 | 0.14243 | 20201113 | 11-16 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/04
AR39A | 20190928 | L10 | 5 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 81.62 | 54.57 | 1.0175 | 80.6025 | 53.5525 | 0.15324 | 0.099876 | 20201113 | 11-16 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/01; Offshore Surf Moor recover
AR39A | 20190928 | L10 | 5 | 8 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 80.06 | 53.57 | 1.0175 | 79.0425 | 52.5525 | 0.15007 | 0.098322 | 20201113 | 11-16 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/02
AR39A | 20190929 | L10 | 6 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 87.11 | 57.47 | 1.02 | 86.09 | 56.45 | 0.16791 | 0.098899 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/01; Offshore SurfM deploy; OOI logsheet says N8 but incorrect-should be N7. Note says N5 didn't confirm fire so called subsiquent N +1 for N6-9 but bottle file came out correct N5-8 can confirm by comparing depths of logsheet to btl file
AR39A | 20190929 | L10 | 6 | 8 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 88.44 | 57.59 | 1.02 | 87.42 | 56.57 | 0.17477 | 0.092611 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/02; OOI logsheet says N9 but incorrect-should be N8. Note says N5 didn't confirm fire so called subsiquent N +1 for N6-9 but bottle file came out correct N5-8 can confirm by comparing depths of logsheet to btl file
AR39A | 20190930 | L12 | 7 | 5 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 195.6 | 138.1 | 1.02 | 194.58 | 137.08 | 0.32574 | 0.32217 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/01; UpstreamOffshore
AR39A | 20190930 | L12 | 7 | 6 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 196.2 | 137.3 | 1.02 | 195.18 | 136.28 | 0.33367 | 0.31046 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/02 | Read tissue capsule as 17, not 7; But, there is no cast 17 and the numbers match with cast 7.
AR39A | 20190930 | L12 | 7 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 169 | 110.9 | 1.02 | 167.98 | 109.88 | 0.32883 | 0.19003 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/03
AR39A | 20190930 | L12 | 7 | 8 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 163.3 | 104.8 | 1.02 | 162.28 | 103.78 | 0.3314 | 0.15911 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/04 | Read tissue capsule as 17, not 7; But, there is no cast 17 and the numbers match with cast 7.
AR39A | 20190930 | L13 | 13 | 1 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 145.9 | 106.5 | 1.0175 | 144.8825 | 105.4825 | 0.4464 | 0.55072 | 20201113 | 11-16 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 13/01; Upstream Inshore
AR39A | 20190930 | L13 | 13 | 3 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 230.4 | 160.4 | 1.02 | 229.38 | 159.38 | 0.7931 | 0.71351 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 13/02
AR39A | 20190930 | L13 | 13 | 4 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 214 | 149.9 | 1.02 | 212.98 | 148.88 | 0.72557 | 0.68046 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 13/03
AR39A | 20190930 | L13 | 13 | 5 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 0.5 | Cal_111320b_12mm | 1.8343 | 0.5442 | 325.8 | 225.1 | 1.02 | 324.78 | 224.08 | 1.1409 | 0.97729 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 13/04
AR39A | 20190930 | L13 | 13 | 6 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 0.333 | Cal_111320b_12mm | 1.8343 | 0.5442 | 213.7 | 143.2 | 1.0175 | 212.6825 | 142.1825 | 1.1993 | 0.81874 | 20201113 | 11-16 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 13/05
AR39A | 20190930 | L13 | 13 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 206.2 | 132.1 | 1.0175 | 205.1825 | 131.0825 | 0.41978 | 0.19978 | 20201113 | 11-16 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 13/06
AR39A | 20191002 |  | 14 | 5 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 143.4 | 91.76 | 1.0175 | 142.3825 | 90.7425 | 0.29254 | 0.13635 | 20201113 | 11-16 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 14/01; AS-1 AUV deploy | Possinly too many drops of acid.
AR39A | 20191002 |  | 14 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 72.23 | 48.83 | 1.02 | 71.21 | 47.81 | 0.26512 | 0.18682 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 14/02
AR39A | 20191002 |  | 14 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 125.5 | 84.89 | 1.02 | 124.48 | 83.87 | 0.22984 | 0.1662 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 14/03
AR39A | 20191002 |  | 14 | 8 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 128 | 85.6 | 1.02 | 126.98 | 84.58 | 0.2402 | 0.15957 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 14/04
AR39A | 20191002 |  | 20 | 1 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 97.69 | 70.55 | 1.02 | 96.67 | 69.53 | 0.15375 | 0.17488 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 20/01; AS-7
AR39A | 20191002 |  | 20 | 3 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 119.7 | 87.18 | 1.02 | 118.68 | 86.16 | 0.18423 | 0.22301 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 20/02
AR39A | 20191002 |  | 20 | 4 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 123.9 | 88.6 | 1.02 | 122.88 | 87.58 | 0.19979 | 0.21377 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 20/03
AR39A | 20191002 |  | 20 | 5 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 373.2 | 242.1 | 1.0175 | 372.1825 | 241.0825 | 0.74268 | 0.39679 | 20201113 | 11-16 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 20/04
AR39A | 20191002 |  | 20 | 6 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 353.8 | 229.7 | 1.02 | 352.78 | 228.68 | 0.70303 | 0.37782 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 20/05
AR39A | 20191002 |  | 20 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 150.5 | 95.81 | 1.0175 | 149.4825 | 94.7925 | 0.30982 | 0.13822 | 20201113 | 11-16 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 20/06
AR39B | 20191005 | L1 | 1 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 195.8 | 125.3 | 0.71483 | 195.0852 | 124.5852 | 1.4543 | 0.68992 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L1 | 1 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 227.7 | 147.7 | 0.70517 | 226.9948 | 146.9948 | 1.6503 | 0.87964 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L1 | 1 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 155.8 | 95.37 | 0.70517 | 155.0948 | 94.6648 | 1.2466 | 0.38269 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L1 | 1 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 139.8 | 84.88 | 0.71483 | 139.0852 | 84.1652 | 1.1329 | 0.31564 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L1 | 1 | 4 |  |  | a | 260 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 54.08 | 34.18 | 0.70517 | 53.3748 | 33.4748 | 0.45787 | 0.18474 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L1 | 1 | 4 |  |  | b | 260 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 56.38 | 36.42 | 0.71483 | 55.6652 | 35.7052 | 0.45925 | 0.22617 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L2 | 2 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 133.4 | 93.68 | 0.70517 | 132.6948 | 92.9748 | 0.81936 | 0.78082 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L2 | 2 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 115.6 | 80.56 | 0.71483 | 114.8852 | 79.8452 | 0.72282 | 0.65139 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L2 | 2 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 147.9 | 103.4 | 0.71483 | 147.1852 | 102.6852 | 0.91797 | 0.84934 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high range of phaeo:chl ratio but pretty dynamic water on CTD trace
AR39B | 20191005 | L2 | 2 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 143.1 | 106.6 | 0.71483 | 142.3852 | 105.8852 | 0.75294 | 1.0694 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high range of phaeo:chl ratio but pretty dynamic water on CTD trace
AR39B | 20191005 | L2 | 2 | 4 |  |  | a | 279.94 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 79.71 | 51.46 | 0.71483 | 78.9952 | 50.7452 | 0.6037 | 0.30106 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 3 | Taylor previously QC'ed
AR39B | 20191005 | L2 | 2 | 4 |  |  | b | 280.59 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 87.78 | 57.4 | 0.65567 | 87.1243 | 56.7443 | 0.64771 | 0.36166 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 4 | Taylor previously QC'ed
AR39B | 20191005 | L2 | 2 | 4 |  |  | a | 279.91 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 20.39 | 13.5 | 0.65567 | 19.7343 | 12.8443 | 0.14725 | 0.081778 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 9; very poor replication, could water have been poured unevently between replicates? | Taylor previously QC'ed
AR39B | 20191005 | L2 | 2 | 4 |  |  | b | 280.02 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 114.1 | 71.13 | 0.71483 | 113.3852 | 70.4152 | 0.918 | 0.3371 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 10; COULD THIS BE TAR7 SUPPOSED TO BE C2N11>20A?? very poor replication, could water have been poured unevently between replicates?; DF thinks this should be TAR 7 | Taylor previously QC'ed
AR39B | 20191005 | L2 | 2 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 337.4 | 201.2 | 0.70517 | 336.6948 | 200.4948 | 2.8096 | 0.6411 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR39B | 20191005 | L2 | 2 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 323.4 | 197.5 | 0.70517 | 322.6948 | 196.7948 | 2.5971 | 0.78989 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR39B | 20191005 | L2 | 2 | 7 |  |  | a | 279.7 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 147 | 90.25 | 0.65567 | 146.3443 | 89.5943 | 1.2138 | 0.38501 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 5; poor chl replication, just natural variance? | Taylor previously QC'ed
AR39B | 20191005 | L2 | 2 | 7 |  |  | b | 280.53 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 108.2 | 67.26 | 0.65567 | 107.5443 | 66.6043 | 0.87304 | 0.31198 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 6; poor chl replication, just natural variance | Taylor previously QC'ed, Could this be switched with >20 rep b?, QF was 1; DF thinks could be switched with >20b
AR39B | 20191005 | L2 | 2 | 7 |  |  | a | 279.91 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 133.2 | 79.66 | 0.71483 | 132.4852 | 78.9452 | 1.1443 | 0.26344 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 11; poor chl replication, just natural variance?; pretty high values for >20 but base of mixed layer so maybe makes sense? | Taylor previously QC'ed
AR39B | 20191005 | L2 | 2 | 7 |  |  | b | 281.05 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 178.7 | 106.5 | 0.71483 | 177.9852 | 105.7852 | 1.5368 | 0.34183 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR 12; poor chl replication, just natural variance? ; pretty high values for >20 but base of mixed layer so maybe makes sense? Could this be switched with >5 rep b? QF was 1; DF thinks this should be >5 rep b | Taylor previously QC'ed
AR39B | 20191005 | L2 | 2 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 268.5 | 194.1 | 0.71483 | 267.7852 | 193.3852 | 1.5348 | 1.7936 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | high chl range and this replicate phaeo is greater than chl which is odd at surface; other rep does have high phaeo but still less than chl
AR39B | 20191005 | L2 | 2 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 306.9 | 193.4 | 0.70517 | 306.1948 | 192.6948 | 2.3413 | 0.97512 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high chl range but this replicate with phaeo pattern seems more reasonable
AR39B | 20191005 | L2 | 2 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 104.5 | 65.29 | 0.71483 | 103.7852 | 64.5752 | 0.80884 | 0.30256 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L2 | 2 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 116.2 | 74.19 | 0.70517 | 115.4948 | 73.4848 | 0.8666 | 0.39814 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L2 | 2 | 11 |  |  | a | 280.07 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 180.3 | 110.7 | 0.71483 | 179.5852 | 109.9852 | 1.4866 | 0.47341 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 1
AR39B | 20191005 | L2 | 2 | 11 |  |  | b | 280.31 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 188.5 | 114 | 0.71483 | 187.7852 | 113.2852 | 1.5899 | 0.42719 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 2
AR39B | 20191005 | L2 | 2 | 11 |  |  | a | 280.24 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 38.83 | 23.32 | 0.65567 | 38.1743 | 22.6643 | 0.33109 | 0.072569 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 7; COULD THIS BE TAR10 SUPPOSED TO BE C2N4>20B?? very poor chl replication and significantly lower than >5 filter, would think value would be a bit higher?; DF thinks this should be TAR10 | Taylor previously QC'ed
AR39B | 20191005 | L2 | 2 | 11 |  |  | b | 280.66 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 152 | 69.56 | 0.65567 | 151.3443 | 68.9043 | 1.7572 | -0.53183 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 8; negative phaeo, chl value greater than both >5 but could be natural variability | Taylor previously QC'ed
AR39B | 20191005 | L3 | 3 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 54.68 | 32.6 | 0.71483 | 53.9652 | 31.8852 | 0.45548 | 0.093296 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | extremely high range of phaeo replicates and also phaeo:chl ratio but unclear if one is more correc than other; bottom water of interesting CTD water mass
AR39B | 20191005 | L3 | 3 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 73.49 | 53.16 | 0.70517 | 72.7848 | 52.4548 | 0.41938 | 0.48342 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | in LabNB could this Rb be a duplicate reading of previous Ra and then this Ra actually be the true Rb of this sample?? extremely high range of phaeo replicates and also phaeo:chl ratio but unclear if one is more correc than other; bottom water of interesting CTD water mass
AR39B | 20191005 | L3 | 3 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 132.5 | 89.47 | 0.70517 | 131.7948 | 88.7648 | 0.88764 | 0.64008 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L3 | 3 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 147.3 | 94.7 | 0.71483 | 146.5852 | 93.9852 | 1.0851 | 0.53251 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L3 | 3 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 96.29 | 60.04 | 0.70517 | 95.5848 | 59.3348 | 0.74778 | 0.27342 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L3 | 3 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 87.39 | 53.91 | 0.70517 | 86.6848 | 53.2048 | 0.69064 | 0.22506 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L3 | 3 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 77.86 | 48.4 | 0.71483 | 77.1452 | 47.6852 | 0.60771 | 0.21299 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L3 | 3 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 79.89 | 48.96 | 0.71483 | 79.1752 | 48.2452 | 0.63804 | 0.1923 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L3 | 3 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 56.04 | 36.14 | 0.70517 | 55.3348 | 35.4348 | 0.41051 | 0.19936 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high range of phaeo replicates
AR39B | 20191005 | L3 | 3 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 49.41 | 30.97 | 0.70517 | 48.7048 | 30.2648 | 0.38039 | 0.1405 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high range of phaeo replicates
AR39B | 20191005 | L4 | 4 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 55.26 | 44.03 | 0.71483 | 54.5452 | 43.3152 | 0.23166 | 0.51384 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high range of chl replicates
AR39B | 20191005 | L4 | 4 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 32.6 | 24.73 | 0.71483 | 31.8852 | 24.0152 | 0.16235 | 0.25098 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high range of chl replicates
AR39B | 20191005 | L4 | 4 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 107 | 73.92 | 0.70517 | 106.2948 | 73.2148 | 0.68239 | 0.5777 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L4 | 4 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 110 | 73.83 | 0.70517 | 109.2948 | 73.1248 | 0.74613 | 0.51241 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L4 | 4 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 63.87 | 40.81 | 0.70517 | 63.1648 | 40.1048 | 0.47569 | 0.21455 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L4 | 4 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 68.67 | 44.28 | 0.70517 | 67.9648 | 43.5748 | 0.50313 | 0.24683 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L4 | 4 | 5 |  |  | a | 280.59 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 23.98 | 16.54 | 0.77367 | 23.2063 | 15.7663 | 0.15862 | 0.12183 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 17 | Taylor previously QC'ed
AR39B | 20191005 | L4 | 4 | 5 |  |  | b | 279.92 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 24.76 | 16.06 | 0.77367 | 23.9863 | 15.2863 | 0.18593 | 0.086635 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 18 | Taylor previously QC'ed
AR39B | 20191005 | L4 | 4 | 5 |  |  | a | 280.02 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 16.59 | 10.58 | 0.71483 | 15.8752 | 9.8652 | 0.1284 | 0.047444 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 23 | Taylor previously QC'ed
AR39B | 20191005 | L4 | 4 | 5 |  |  | b | 280.91 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 59.51 | 36.86 | 0.65567 | 58.8543 | 36.2043 | 0.48235 | 0.16092 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR24; chl much higher than most values for the entire cast, including WSW, poor replication; could it be one giant cell made it into the sample? DF thinks this sample should be tossed given WSW and 5 values. DF changed QF to 3 | Taylor previously QC'ed
AR39B | 20191005 | L4 | 4 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 35.95 | 22.23 | 0.71483 | 35.2352 | 21.5152 | 0.28302 | 0.087273 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L4 | 4 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 47.7 | 30.38 | 0.71483 | 46.9852 | 29.6652 | 0.35729 | 0.15328 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L4 | 4 | 8 |  |  | a | 279.7 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 15.26 | 10.52 | 0.65567 | 14.6043 | 9.8643 | 0.10138 | 0.074646 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR 15; could water have been poured unequally between duplicates?; could this be the >20 put in the wrong vial? 5a more simliar to 20b. That leaves 5's mismatched though | Taylor previously QC'ed
AR39B | 20191005 | L4 | 4 | 8 |  |  | b | 280.43 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 87.05 | 53.28 | 0.65567 | 86.3943 | 52.6243 | 0.7204 | 0.21622 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR 16; this is double either extracted WSW; CTD fluor ~0.6 on upcast; could the water have been unequally pored between duplicates? Sig higher than WSW so toss | Taylor previously QC'ed
AR39B | 20191005 | L4 | 4 | 8 |  |  | a | 280.09 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 159.8 | 95.08 | 0.65567 | 159.1443 | 94.4243 | 1.3823 | 0.30032 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR 21; something clearly wrong with this sample. Chl way too high. Triple any WSW values. Checked lab NB | Taylor previously QC'ed
AR39B | 20191005 | L4 | 4 | 8 |  |  | b | 279.91 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 20.17 | 13.05 | 0.65567 | 19.5143 | 12.3943 | 0.15217 | 0.068838 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR 22 | Taylor previously QC'ed
AR39B | 20191005 | L4 | 4 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 48.31 | 31.39 | 0.70517 | 47.6048 | 30.6848 | 0.34903 | 0.17908 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L4 | 4 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 44.83 | 28.67 | 0.71483 | 44.1152 | 27.9552 | 0.33336 | 0.14778 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L4 | 4 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 37.82 | 24.27 | 0.70517 | 37.1148 | 23.5648 | 0.27952 | 0.12606 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L4 | 4 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 35.7 | 22.42 | 0.71483 | 34.9852 | 21.7052 | 0.27395 | 0.09962 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L4 | 4 | 12 |  |  | a | 280.07 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 21.95 | 14.46 | 0.65567 | 21.2943 | 13.8043 | 0.15999 | 0.086023 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 13 | Taylor previously QC'ed
AR39B | 20191005 | L4 | 4 | 12 |  |  | b | 280.1 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 124.4 | 76.07 | 0.65567 | 123.7443 | 75.4143 | 1.0322 | 0.31161 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR 14; poor chl replication and this is 3x higher than either extracted WSW; within reasonable range of CTD fluor but no other extracted chl that high so throw out | Taylor previously QC'ed
AR39B | 20191005 | L4 | 4 | 12 |  |  | a | 280.66 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 8.036 | 5.557 | 0.65567 | 7.3803 | 4.9013 | 0.05284 | 0.034324 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 19 | Taylor previously QC'ed
AR39B | 20191005 | L4 | 4 | 12 |  |  | b | 280.24 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.5 | 7.957 | 0.65567 | 10.8443 | 7.3013 | 0.075632 | 0.054407 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR 20 | Taylor previously QC'ed
AR39B | 20191005 | L5 | 5 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 65.14 | 49.1 | 0.71483 | 64.4252 | 48.3852 | 0.33088 | 0.50187 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L5 | 5 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 56.1 | 40.41 | 0.71483 | 55.3852 | 39.6952 | 0.32366 | 0.35953 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L5 | 5 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 41.19 | 28.8 | 0.71483 | 40.4752 | 28.0852 | 0.25559 | 0.22778 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L5 | 5 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 43.82 | 31.62 | 0.71483 | 43.1052 | 30.9052 | 0.25167 | 0.28024 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L5 | 5 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 74.39 | 49.53 | 0.70517 | 73.6752 | 48.8152 | 0.51282 | 0.32733 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L5 | 5 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 60.54 | 39.11 | 0.71483 | 59.8252 | 38.3952 | 0.44207 | 0.21875 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L5 | 5 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 39.14 | 24.64 | 0.71483 | 38.4252 | 23.9252 | 0.29911 | 0.11266 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L5 | 5 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 42.69 | 27.43 | 0.71483 | 41.9752 | 26.7152 | 0.31479 | 0.145 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L5 | 5 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 46.72 | 30.8 | 0.70517 | 46.0148 | 30.0948 | 0.32841 | 0.18955 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L5 | 5 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 39.43 | 25.25 | 0.71483 | 38.7152 | 24.5352 | 0.29251 | 0.12976 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L5 | 5 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 39.31 | 25.23 | 0.70517 | 38.6048 | 24.5248 | 0.29045 | 0.13165 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L5 | 5 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 39.52 | 25.86 | 0.71483 | 38.8052 | 25.1452 | 0.28178 | 0.15099 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L6 | 6 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 40.77 | 32.08 | 0.70517 | 40.0648 | 31.3748 | 0.17926 | 0.36073 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L6 | 6 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 39.27 | 29.2 | 0.70517 | 38.5648 | 28.4948 | 0.20773 | 0.28269 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L6 | 6 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 20.56 | 15.05 | 0.71483 | 19.8452 | 14.3352 | 0.11366 | 0.13306 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L6 | 6 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 20.78 | 15.09 | 0.71483 | 20.0652 | 14.3752 | 0.11738 | 0.13003 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L6 | 6 | 4 |  |  | a | 280.16 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 9.468 | 6.672 | 0.65567 | 8.8123 | 6.0163 | 0.059703 | 0.04748 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR31 | Taylor previously QC'ed
AR39B | 20191005 | L6 | 6 | 4 |  |  | b | 280.07 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.17 | 7.84 | 0.65567 | 10.5143 | 7.1843 | 0.071128 | 0.056905 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR32 | Taylor previously QC'ed
AR39B | 20191005 | L6 | 6 | 4 |  |  | a | 280.24 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.958 | 3.357 | 0.65567 | 4.3023 | 2.7013 | 0.034176 | 0.013935 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR39; high range chl replicates but very low concentrations | Taylor previously QC'ed
AR39B | 20191005 | L6 | 6 | 4 |  |  | b | 280.66 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.598 | 2.509 | 0.71483 | 2.8832 | 1.7942 | 0.023212 | 0.0086949 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR40; high range of chl replicates but very low concentrations | Taylor previously QC'ed
AR39B | 20191005 | L6 | 6 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 56.21 | 38.69 | 0.70517 | 55.5048 | 37.9848 | 0.36141 | 0.29234 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L6 | 6 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 52.49 | 35.74 | 0.71483 | 51.7752 | 35.0252 | 0.34553 | 0.25729 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L6 | 6 | 9 |  |  | a | 279.89 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 15.53 | 10.71 | 0.65567 | 14.8743 | 10.0543 | 0.10302 | 0.076274 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR29 | Taylor previously QC'ed
AR39B | 20191005 | L6 | 6 | 9 |  |  | b | 279.91 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 17.13 | 11.64 | 0.65567 | 16.4743 | 10.9843 | 0.11733 | 0.078532 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR30 | Taylor previously QC'ed
AR39B | 20191005 | L6 | 6 | 9 |  |  | a | 279.94 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 5.201 | 3.838 | 0.65567 | 4.5453 | 3.1823 | 0.029127 | 0.027612 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR37 | Taylor previously QC'ed; very poor replication (117%)
AR39B | 20191005 | L6 | 6 | 9 |  |  | b | 280.24 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 14.14 | 8.851 | 0.65567 | 13.4843 | 8.1953 | 0.1129 | 0.033058 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR38 | Taylor previously QC'ed; very poor replication (117%) but chl values are low … this is higher than one of the reps
AR39B | 20191005 | L6 | 6 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 62.84 | 40.26 | 0.70517 | 62.1348 | 39.5548 | 0.46579 | 0.21498 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L6 | 6 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 68.87 | 44.18 | 0.70517 | 68.1648 | 43.4748 | 0.50932 | 0.23892 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L6 | 6 | 14 |  |  | a | 280.53 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.9 | 8.133 | 0.71483 | 11.1852 | 7.4182 | 0.080331 | 0.051653 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR27; 5s and 20s look like replicates could have been mxied up? All values extremely low, some in noise of fluroometer | Taylor previously QC'ed
AR39B | 20191005 | L6 | 6 | 14 |  |  | b | 281.05 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 6.144 | 4.229 | 0.65567 | 5.4883 | 3.5733 | 0.040761 | 0.022697 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR28; could this be a >20 replicate? All values extremely low, some in noise of fluroometer | Taylor previously QC'ed; DF thinks this was switched with >20 rep a; QF was a 1 and DF changed to 2
AR39B | 20191005 | L6 | 6 | 14 |  |  | a | 279.91 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 10.14 | 6.455 | 0.65567 | 9.4843 | 5.7993 | 0.078756 | 0.024654 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR35; could this be a >5 replicate? Looks like they might have been mixed up?All values extremely low, some in noise of fluroometer | Taylor previously QC'ed; DF thinks this was switched with >5 rep b; QF was a 1 and DF changed to a 2
AR39B | 20191005 | L6 | 6 | 14 |  |  | b | 280.09 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 5.132 | 3.459 | 0.65567 | 4.4763 | 2.8033 | 0.035732 | 0.014223 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR36; 5s and 20s look like replicates might have been mixed up?All values extremely low, some in noise of fluroometer | Taylor previously QC'ed
AR39B | 20191005 | L6 | 6 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 61.32 | 39.49 | 0.70517 | 60.6148 | 38.7848 | 0.45032 | 0.2172 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L6 | 6 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 60 | 36.85 | 0.71483 | 59.2852 | 36.1352 | 0.47755 | 0.14437 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L6 | 6 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 43.6 | 26.74 | 0.71483 | 42.8852 | 26.0252 | 0.3478 | 0.10012 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L6 | 6 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 51.8 | 33.04 | 0.71483 | 51.0852 | 32.3252 | 0.38699 | 0.16936 | 20201113 | 10-55 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L6 | 6 | 20 |  |  | a | 280.41 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 18.95 | 12.18 | 0.71483 | 18.2352 | 11.4652 | 0.14443 | 0.059644 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR25; higher range of chl replicates (27%) | Taylor previously QC'ed
AR39B | 20191005 | L6 | 6 | 20 |  |  | b | 279.67 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 13.87 | 8.761 | 0.71483 | 13.1552 | 8.0462 | 0.10928 | 0.034313 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR26; higher range of chl replicates (27%) | Taylor previously QC'ed
AR39B | 20191005 | L6 | 6 | 20 |  |  | a | 281.59 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 10.41 | 6.784 | 0.65567 | 9.7543 | 6.1283 | 0.077033 | 0.031592 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR33 | Taylor previously QC'ed
AR39B | 20191005 | L6 | 6 | 20 |  |  | b | 281.08 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 9.993 | 6.536 | 0.65567 | 9.3373 | 5.8803 | 0.073576 | 0.030842 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR34 | Taylor previously QC'ed
AR39B | 20191005 | L6 | 7 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 238.4 | 154.5 | 0.70517 | 237.6948 | 153.7948 | 1.7307 | 0.91622 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 07/01; UpstreamInshore
AR39B | 20191005 | L6 | 7 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 238.1 | 153 | 0.70517 | 237.3948 | 152.2948 | 1.7555 | 0.86565 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 07/02; UpstreamInshore
AR39B | 20191005 | L6 | 7 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 276.3 | 178.7 | 0.70517 | 275.5948 | 177.9948 | 2.0133 | 1.0501 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 07/03; UpstreamInshore
AR39B | 20191005 | L6 | 7 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 270 | 174.1 | 0.70517 | 269.2948 | 173.3948 | 1.9783 | 1.006 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 07/04; UpstreamInshore
AR39B | 20191005 | L6 | 7 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 214 | 133.9 | 0.70517 | 213.2948 | 133.1948 | 1.6523 | 0.64006 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 07/05; UpstreamInshore
AR39B | 20191005 | L6 | 7 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 230.1 | 142.9 | 0.70517 | 229.3948 | 142.1948 | 1.7988 | 0.6485 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 07/06; UpstreamInshore
AR39B | 20191005 | L8 | 10 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 26.44 | 18.65 | 0.70517 | 25.7348 | 17.9448 | 0.1607 | 0.14815 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L8 | 10 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 27.06 | 19.18 | 0.70517 | 26.3548 | 18.4748 | 0.16255 | 0.15542 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L8 | 10 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 19.53 | 14.65 | 0.71483 | 18.8152 | 13.9352 | 0.10067 | 0.13917 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L8 | 10 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 20.9 | 15.51 | 0.70517 | 20.1948 | 14.8048 | 0.11119 | 0.14362 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L8 | 10 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 52.58 | 37.46 | 0.70517 | 51.8748 | 36.7548 | 0.3119 | 0.32068 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L8 | 10 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 44.12 | 31.5 | 0.71483 | 43.4052 | 30.7852 | 0.26033 | 0.26951 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L8 | 10 | 7 |  |  | a | 277 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 13.51 | 9.398 | 0.65567 | 12.8543 | 8.7423 | 0.088805 | 0.06872 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR47
AR39B | 20191005 | L8 | 10 | 7 |  |  | b | 289.02 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 13.77 | 9.665 | 0.65567 | 13.1143 | 9.0093 | 0.084967 | 0.070618 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR48
AR39B | 20191005 | L8 | 10 | 7 |  |  | a | 280 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 7.175 | 5.0975 | 0.71483 | 6.4602 | 4.3827 | 0.044386 | 0.033737 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR55 | Taylor previously QC'ed
AR39B | 20191005 | L8 | 10 | 7 |  |  | b | 277 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 7.356 | 4.847 | 0.71483 | 6.6412 | 4.1322 | 0.054186 | 0.02027 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR56 | Taylor previously QC'ed
AR39B | 20191005 | L8 | 10 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 64.61 | 42.25 | 0.71483 | 63.8932 | 41.5332 | 0.46125 | 0.25357 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L8 | 10 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 74.83 | 49.12 | 0.71683 | 74.1152 | 48.4052 | 0.53036 | 0.30274 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L8 | 10 | 9 |  |  | a | 279.67 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 23.01 | 14.62 | 0.71483 | 22.2952 | 13.9052 | 0.17947 | 0.068695 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR45 | Taylor previously QC'ed
AR39B | 20191005 | L8 | 10 | 9 |  |  | b | 281.59 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 21.91 | 14.11 | 0.65567 | 21.2543 | 13.4543 | 0.16571 | 0.07277 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR46 | Taylor previously QC'ed
AR39B | 20191005 | L8 | 10 | 9 |  |  | a | 278 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 12.67 | 7.796 | 0.71483 | 11.9552 | 7.0812 | 0.10488 | 0.022251 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR53 | Taylor previously QC'ed
AR39B | 20191005 | L8 | 10 | 9 |  |  | b | 277 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 9.804 | 6.171 | 0.71683 | 9.0872 | 5.4542 | 0.07846 | 0.019816 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR54 | Taylor previously QC'ed
AR39B | 20191005 | L8 | 10 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 75.73 | 49.51 | 0.71683 | 75.0132 | 48.7932 | 0.54088 | 0.2989 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L8 | 10 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 63.49 | 40.39 | 0.71483 | 62.7752 | 39.6752 | 0.47652 | 0.20633 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L8 | 10 | 14 |  |  | a | 279.91 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 20.93 | 13.64 | 0.71483 | 20.2152 | 12.9252 | 0.1558 | 0.07467 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR43 | Taylor previously QC'ed
AR39B | 20191005 | L8 | 10 | 14 |  |  | b | 276 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 16.53 | 10.75 | 0.71483 | 15.8152 | 10.0352 | 0.12528 | 0.056195 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR44 | Taylor previously QC'ed
AR39B | 20191005 | L8 | 10 | 14 |  |  | a | 277 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 13.16 | 8.075 | 0.71483 | 12.4452 | 7.3602 | 0.10982 | 0.022802 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR51 | Taylor previously QC'ed
AR39B | 20191005 | L8 | 10 | 14 |  |  | b | 277 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 14.38 | 9.395 | 0.71483 | 13.6652 | 8.6802 | 0.10766 | 0.048746 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR52 | Taylor previously QC'ed
AR39B | 20191005 | L8 | 10 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 67.97 | 43.94 | 0.71483 | 67.2552 | 43.2252 | 0.4957 | 0.24824 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L8 | 10 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 79.32 | 52.39 | 0.70517 | 78.6148 | 51.6848 | 0.55552 | 0.33402 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L8 | 10 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 64.57 | 42.47 | 0.71683 | 63.8532 | 41.7532 | 0.45589 | 0.26272 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L8 | 10 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 61.76 | 39.22 | 0.71683 | 61.0432 | 38.5032 | 0.46497 | 0.19771 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L8 | 10 | 21 |  |  | a | 280.31 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 20.06 | 13.28 | 0.71483 | 19.3452 | 12.5652 | 0.1447 | 0.079038 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR41 | Taylor previously QC'ed
AR39B | 20191005 | L8 | 10 | 21 |  |  | b | 280.16 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 20.25 | 13.06 | 0.71483 | 19.5352 | 12.3452 | 0.15353 | 0.066406 | 20201113 | 10-50 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR42 | Taylor previously QC'ed
AR39B | 20191005 | L8 | 10 | 21 |  |  | a | 280.66 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.04 | 7.051 | 0.71483 | 10.3252 | 6.3362 | 0.085025 | 0.027655 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR49 | Taylor previously QC'ed
AR39B | 20191005 | L8 | 10 | 21 |  |  | b | 280.07 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 12.74 | 7.917 | 0.71483 | 12.0252 | 7.2022 | 0.10302 | 0.025332 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR50 | Taylor previously QC'ed
AR39B | 20191005 | L7 | 11 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 15.83 | 12.92 | 0.71683 | 15.1132 | 12.2032 | 0.060029 | 0.15 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L7 | 11 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 15.04 | 11.99 | 0.71683 | 14.3232 | 11.2732 | 0.062917 | 0.1311 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L7 | 11 | 6 |  |  | a | 270 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 31.94 | 23.31 | 0.71683 | 31.2232 | 22.5932 | 0.19121 | 0.22644 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | subtract 20ml from chla leaky funnel
AR39B | 20191005 | L7 | 11 | 6 |  |  | b | 195 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 22.25 | 15.84 | 0.70517 | 21.5448 | 15.1348 | 0.19665 | 0.19074 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | found 95ml in one of chl bottles, pretty sure from chlb; once samples run, adjust chlb volume and definitely from that one
AR39B | 20191005 | L7 | 11 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 75.94 | 50.87 | 0.70517 | 75.2348 | 50.1648 | 0.51716 | 0.34623 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L7 | 11 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 62.85 | 41.11 | 0.71483 | 62.1352 | 40.3952 | 0.44846 | 0.24677 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L7 | 11 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 82.55 | 53.87 | 0.71683 | 81.8332 | 53.1532 | 0.59162 | 0.32319 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L7 | 11 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 83.58 | 54.73 | 0.71683 | 82.8632 | 54.0132 | 0.59513 | 0.33448 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L7 | 11 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 72.64 | 47.98 | 0.71683 | 71.9232 | 47.2632 | 0.5087 | 0.30474 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L7 | 11 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 75.78 | 49.57 | 0.71683 | 75.0632 | 48.8532 | 0.54067 | 0.30014 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L7 | 11 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 53.16 | 33.93 | 0.70517 | 52.4548 | 33.2248 | 0.39669 | 0.17514 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | These WSW values are smaller than <10s
AR39B | 20191005 | L7 | 11 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 56.39 | 36.98 | 0.71683 | 55.6732 | 36.2632 | 0.4004 | 0.22372 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | These WSW values are smaller than <10s
AR39B | 20191005 | L7 | 11 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 67.54 | 43.36 | 0.71683 | 66.8232 | 42.6432 | 0.4988 | 0.23513 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | These <10 values are larger than WSW…was the water pre-filtered?
AR39B | 20191005 | L7 | 11 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 74.99 | 49.6 | 0.71683 | 74.2732 | 48.8832 | 0.52376 | 0.31757 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | These <10 values are larger than WSW…was the water pre-filtered?
AR39B | 20191005 | L9 | 12 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 17.43 | 12.43 | 0.70517 | 16.7248 | 11.7248 | 0.10314 | 0.098653 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L9 | 12 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 19.29 | 13.94 | 0.71683 | 18.5732 | 13.2232 | 0.11036 | 0.11722 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L9 | 12 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 20.62 | 14.96 | 0.71683 | 19.9032 | 14.2432 | 0.11676 | 0.12838 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L9 | 12 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 21.98 | 16.14 | 0.70517 | 21.2748 | 15.4348 | 0.12047 | 0.14518 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L9 | 12 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 27.05 | 20.39 | 0.71483 | 26.3352 | 19.6752 | 0.13739 | 0.20124 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L9 | 12 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 30 | 22.86 | 0.71683 | 29.2832 | 22.1432 | 0.14729 | 0.23382 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L9 | 12 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 44.36 | 33.94 | 0.71683 | 43.6432 | 33.2232 | 0.21495 | 0.35685 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L9 | 12 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 44.53 | 33.63 | 0.71683 | 43.8132 | 32.9132 | 0.22485 | 0.34162 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L9 | 12 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 60.08 | 38.61 | 0.71483 | 59.3652 | 37.8952 | 0.44289 | 0.20932 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L9 | 12 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 68.42 | 45.23 | 0.71683 | 67.7032 | 44.5132 | 0.47837 | 0.28774 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L9 | 12 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 71.13 | 47.16 | 0.71683 | 70.4132 | 46.4432 | 0.49446 | 0.30486 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L9 | 12 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 72.75 | 48.17 | 0.71683 | 72.0332 | 47.4532 | 0.50705 | 0.30966 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L9 | 12 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 58.49 | 38.23 | 0.70517 | 57.7848 | 37.5248 | 0.41793 | 0.2279 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L9 | 12 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 55.06 | 37.14 | 0.71683 | 54.3432 | 36.4232 | 0.36966 | 0.25721 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L11 | 13 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 19.08 | 14.45 | 0.71683 | 18.3632 | 13.7332 | 0.09551 | 0.14085 | 20201113 | 10-53 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L11 | 13 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 19.89 | 14.89 | 0.70517 | 19.1848 | 14.1848 | 0.10314 | 0.14099 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L11 | 13 | 4 |  |  | a | 279.67 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.903 | 3.752 | 0.71483 | 4.1882 | 3.0372 | 0.02462 | 0.029583 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR63 | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 4 |  |  | b | 279.76 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.882 | 3.998 | 0.71483 | 4.1672 | 3.2832 | 0.018903 | 0.039672 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR64 | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 4 |  |  | a | 280.41 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 2.282 | 1.912 | 0.71483 | 1.5672 | 1.1972 | 0.0078936 | 0.013415 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR71; high range of replicates but values basically zero and in noise of fluorometer | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 4 |  |  | b | 281.08 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 2.273 | 1.605 | 0.71483 | 1.5582 | 0.89017 | 0.014217 | 0.0015897 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR72; high range of replicates but values basically zero and in noise of fluorometer | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 48.66 | 32.88 | 0.71683 | 47.9432 | 32.1632 | 0.32552 | 0.22804 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L11 | 13 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 48.63 | 32.88 | 0.71683 | 47.9132 | 32.1632 | 0.3249 | 0.22866 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L11 | 13 | 7 |  |  | a | 279.89 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 8.978 | 6.223 | 0.71483 | 8.2632 | 5.5082 | 0.058884 | 0.039341 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR61 | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 7 |  |  | b | 280.94 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 12.01 | 7.997 | 0.71483 | 11.2952 | 7.2822 | 0.085452 | 0.043923 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR62 | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 7 |  |  | a | 279.91 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 5.425 | 3.873 | 0.65567 | 4.7693 | 3.2173 | 0.033169 | 0.0242 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR69; range of replicates acceptable for >20 size frac | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 7 |  |  | b | 279.94 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 6 | 4.372 | 0.71483 | 5.2852 | 3.6572 | 0.03479 | 0.030415 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR70; range of replicates acceptable for >20 size frac | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 34.22 | 22.95 | 0.71683 | 33.5032 | 22.2332 | 0.23248 | 0.15017 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L11 | 13 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 32.88 | 21.71 | 0.71683 | 32.1632 | 20.9932 | 0.23042 | 0.13089 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L11 | 13 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 35.45 | 23.25 | 0.71683 | 34.7332 | 22.5332 | 0.25167 | 0.13615 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L11 | 13 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 33.25 | 21.92 | 0.71683 | 32.5332 | 21.2032 | 0.23372 | 0.13121 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L11 | 13 | 12 |  |  | a | 279.91 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 6.243 | 4.151 | 0.65567 | 5.5873 | 3.4953 | 0.04471 | 0.017616 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR59 | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 12 |  |  | b | 280.09 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 5.314 | 3.932 | 0.65567 | 4.6583 | 3.2763 | 0.029517 | 0.028867 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR60 | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 12 |  |  | a | 281.05 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.477 | 2.921 | 0.65567 | 2.8213 | 2.2653 | 0.011835 | 0.028395 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR67; range of replicates expected for larger size frac; odd that phaeo>chl at shallower depth but maybe just 1 large chunk of detritus | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 12 |  |  | b | 280.53 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.225 | 2.373 | 0.65567 | 2.5693 | 1.7173 | 0.018169 | 0.012386 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR68; range of replicates expected for larger size frac | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 32.77 | 22.1 | 0.71683 | 32.0532 | 21.3832 | 0.22011 | 0.14792 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L11 | 13 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 32.3 | 21.46 | 0.71683 | 31.5832 | 20.7432 | 0.22361 | 0.1334 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L11 | 13 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 26.83 | 18.11 | 0.71683 | 26.1132 | 17.3932 | 0.17988 | 0.11947 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L11 | 13 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 28.42 | 19.14 | 0.71683 | 27.7032 | 18.4232 | 0.19143 | 0.12565 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L11 | 13 | 15 |  |  | a | 280.31 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 5.661 | 3.864 | 0.71483 | 4.9462 | 3.1492 | 0.038351 | 0.017723 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR57 | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 15 |  |  | b | 280.24 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 5.805 | 4.125 | 0.65567 | 5.1493 | 3.4693 | 0.035863 | 0.025927 | 20201113 | 10-52 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR58 | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 15 |  |  | a | 280.16 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 2.979 | 2.139 | 0.71483 | 2.2642 | 1.4242 | 0.017937 | 0.0074356 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR65; high range phaeo but values within range of fluorometer | Taylor previously QC'ed
AR39B | 20191005 | L11 | 13 | 15 |  |  | b | 280.66 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.411 | 2.618 | 0.71483 | 2.6962 | 1.9032 | 0.016903 | 0.016943 | 20201113 | 10-51 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR66; high range phaeo but values within range of fluorometer | Taylor previously QC'ed
AR39B | 20191005 | L10 | 14 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 5.377 | 4.29 | 0.71683 | 4.6602 | 3.5732 | 0.022423 | 0.039074 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L10 | 14 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 5.455 | 4.205 | 0.70517 | 4.7498 | 3.4998 | 0.025786 | 0.03445 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
AR39B | 20191005 | L10 | 14 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 18.25 | 13.92 | 0.71683 | 17.5332 | 13.2032 | 0.089321 | 0.13792 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L10 | 14 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 19.49 | 15.05 | 0.71683 | 18.7732 | 14.3332 | 0.09159 | 0.1551 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L10 | 14 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 62.3 | 41.74 | 0.71683 | 61.5832 | 41.0232 | 0.42412 | 0.28192 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L10 | 14 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 65.15 | 42.98 | 0.70517 | 64.4448 | 42.2748 | 0.45733 | 0.27026 | 20201113 | 10-57 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L10 | 14 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 82.56 | 53.1 | 0.71483 | 81.8452 | 52.3852 | 0.60771 | 0.29388 | 20201113 | 10-56 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L10 | 14 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 81.65 | 54.03 | 0.71683 | 80.9332 | 53.3132 | 0.56976 | 0.34781 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L10 | 14 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 81.31 | 53.6 | 0.71683 | 80.5932 | 52.8832 | 0.57161 | 0.33855 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L10 | 14 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 79.17 | 51.84 | 0.71683 | 78.4532 | 51.1232 | 0.56378 | 0.3161 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR39B | 20191005 | L10 | 14 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 68.24 | 45.36 | 0.71683 | 67.5232 | 44.6432 | 0.47198 | 0.29637 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high range of chl replicates
AR39B | 20191005 | L10 | 14 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 60.81 | 43.45 | 0.71683 | 60.0932 | 42.7332 | 0.35811 | 0.37737 | 20201113 | 10-54 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | high range of chl replicates; this one phaeo/chl ratio a little high? And raction of <10 to WSW a bit low compared to previous close station surf water
AR39B | 20191013 | L10 | 15 | 3 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 6.073 | 5.788 | 1.02 | 5.053 | 4.768 | 0.0016145 | 0.020921 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | OOI 15/01; Offshore; phaeo is higher than chl but chl value is so low and within floruometer noise
AR39B | 20191013 | L10 | 15 | 4 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 6.151 | 5.398 | 1.0175 | 5.1335 | 4.3805 | 0.0042658 | 0.016439 | 20201113 | 11-15 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | OOI 15/02; Offshore; phaeo is higher than chl but chl value is so low and within floruometer noise
AR39B | 20191013 | L10 | 15 | 5 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 204.5 | 137.5 | 1.02 | 203.48 | 136.48 | 0.3792 | 0.26526 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 15/03; Offshore
AR39B | 20191013 | L10 | 15 | 6 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 197.3 | 133.7 | 1.02 | 196.28 | 132.68 | 0.36029 | 0.26681 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 15/04; Offshore
AR39B | 20191013 | L10 | 15 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 203.9 | 136.3 | 1.02 | 202.88 | 135.28 | 0.38295 | 0.25644 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 15/05; Offshore
AR39B | 20191013 | L10 | 15 | 8 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 200.5 | 133.2 | 1.02 | 199.48 | 132.18 | 0.38126 | 0.24349 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 15/06; Offshore
AR39B | 20191014 | L12 | 16 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 280.3 | 180.1 | 1.0175 | 279.2825 | 179.0825 | 0.56763 | 0.2788 | 20201113 | 11-15 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 16/01; Upstream Offshore
AR39B | 20191014 | L12 | 16 | 7 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 142.8 | 93.68 | 1.02 | 141.78 | 92.66 | 0.55653 | 0.31938 | 20201113 | 11-17 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 16/02; filter damaged before freezing; UpstreamOffshore
EN649 | 20200201 | L1 | 1 | 2 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 205.7 | 123.7 | 0.743 | 204.957 | 122.957 | 3.4841 | 0.89513 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200201 | L1 | 1 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 199.1 | 119.8 | 0.69 | 198.41 | 119.11 | 3.2717 | 0.82831 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200201 | L1 | 1 | 2 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 111.6 | 71.77 | 0.743 | 110.857 | 71.027 | 0.84616 | 0.41868 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200201 | L1 | 1 | 2 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 115.6 | 74.67 | 0.69 | 114.91 | 73.98 | 0.84432 | 0.42894 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | Fluor turned off before Ra recorded, turned fluor back on with vial still in fluor, might have affected Ra reading? Looks okay
EN649 | 20200201 | L1 | 1 | 2 |  |  | a | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 327.8 | 192.1 | 0.85967 | 326.9403 | 191.2403 | 2.7518 | 0.4838 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR5
EN649 | 20200201 | L1 | 1 | 2 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 384 | 223.8 | 0.85967 | 383.1403 | 222.9403 | 3.2597 | 0.52508 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR6
EN649 | 20200201 | L1 | 1 | 2 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 185.2 | 105.5 | 0.93933 | 184.2607 | 104.5607 | 1.6273 | 0.1539 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR11 | poor replication, could >20 N2 and N7 be mixed up?; Taylor previously QC'ed, DF think this switch could have occurred but how wuld we determine which one is which? Like which value is for N7 and which value is for N2; QF was 1
EN649 | 20200201 | L1 | 1 | 2 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 130.5 | 73.63 | 0.93933 | 129.5607 | 72.6907 | 1.1611 | 0.077134 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR12 | poor replication, could >20 N2 and N7 be mixed up?; Taylor previously QC'ed, DF think this switch could have occurred but how would we determine which one is which? Like which value is for N7 and which value is for N2; QF was 1
EN649 | 20200201 | L1 | 1 | 7 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 207.2 | 123 | 0.72767 | 206.4723 | 122.2723 | 3.4738 | 0.735 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200201 | L1 | 1 | 7 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 208.7 | 123.6 | 0.72767 | 207.9723 | 122.8723 | 3.511 | 0.71852 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200201 | L1 | 1 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 110.1 | 70.11 | 0.72767 | 109.3723 | 69.3823 | 0.82493 | 0.3692 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | poor replication
EN649 | 20200201 | L1 | 1 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 187.5 | 116 | 0.72767 | 186.7723 | 115.2723 | 1.4749 | 0.50901 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication; note during extraction not sure if zooplankton on filter - didn't pick off
EN649 | 20200201 | L1 | 1 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 350.6 | 210.4 | 0.85967 | 349.7403 | 209.5403 | 2.8625 | 0.70696 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR3
EN649 | 20200201 | L1 | 1 | 7 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 349.6 | 207.2 | 0.85967 | 348.7403 | 206.3403 | 2.8975 | 0.60546 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR4
EN649 | 20200201 | L1 | 1 | 7 |  |  | a | 292 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 114.4 | 66.48 | 0.85967 | 113.5403 | 65.6203 | 0.98175 | 0.1399 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR9 | poor replication, could >20 N2 and N7 be mixed up?; Taylor previously QC'ed, DF think this switch could have occurred; but how would we determine which one is which? Like which value is for N7 and which value is for N2; QF was 1
EN649 | 20200201 | L1 | 1 | 7 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 191.7 | 109 | 0.93933 | 190.7607 | 108.0607 | 1.6828 | 0.15176 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR10 | poor replication, could >20 N2 and N7 be mixed up?; Taylor previously QC'ed, DF think this switch could have occurred; but how would we determine which one is which? Like which value is for N7 and which value is for N2; QF was 1
EN649 | 20200201 | L1 | 1 | 13 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 210.9 | 127 | 0.743 | 210.157 | 126.257 | 3.4615 | 0.88454 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication; QF was 1
EN649 | 20200201 | L1 | 1 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 359.4 | 214.7 | 0.72767 | 358.6723 | 213.9723 | 2.9849 | 0.69772 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication; QF was 1
EN649 | 20200201 | L1 | 1 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 122.9 | 77.04 | 0.743 | 122.157 | 76.297 | 0.94602 | 0.36712 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication; QF was 1
EN649 | 20200201 | L1 | 1 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 186.9 | 115.7 | 0.72767 | 186.1723 | 114.9723 | 1.4687 | 0.51003 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication; QF was 1
EN649 | 20200201 | L1 | 1 | 13 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 303.2 | 180 | 0.85967 | 302.3403 | 179.1403 | 2.5068 | 0.53437 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR1
EN649 | 20200201 | L1 | 1 | 13 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 303 | 177.4 | 0.85967 | 302.1403 | 176.5403 | 2.5644 | 0.4429 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR2
EN649 | 20200201 | L1 | 1 | 13 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 156.2 | 90.91 | 0.85967 | 155.3403 | 90.0503 | 1.333 | 0.20094 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR7
EN649 | 20200201 | L1 | 1 | 13 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 202.4 | 120.1 | 0.85967 | 201.5403 | 119.2403 | 1.6803 | 0.35088 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR8
EN649 | 20200201 | MVCO | 2 | 1 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 156.3 | 92.91 | 0.69 | 155.61 | 92.22 | 2.2319 | 0.69693 | [6 empty] | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | is calibration correct? And when was sample analyzed? DF had flagged as questionable but Taylor thinks within reason so change back to 1
EN649 | 20200201 | MVCO | 2 | 1 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 166.2 | 98.85 | 0.743 | 165.457 | 98.107 | 2.8616 | 0.63254 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN649 | 20200201 | MVCO | 2 | 1 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 90.64 | 57.68 | 0.743 | 89.897 | 56.937 | 0.70021 | 0.31371 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200201 | MVCO | 2 | 2 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 82.84 | 52.41 | 0.69 | 82.15 | 51.72 | 0.62772 | 0.26242 | 20201113 | 10-96 | 20200818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | Niskin 1 ran out of water for <10b so sampled form N2 which makes <10s not true replicates; Cast was missing from tissue capsule but N2<10b written. Process of elimination plus reasonable rep match
EN649 | 20200201 | MVCO | 2 | 1 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 289.8 | 172 | 0.85967 | 288.9403 | 171.1403 | 2.4052 | 0.51017 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR157
EN649 | 20200201 | MVCO | 2 | 1 |  |  | b | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 271.2 | 160.9 | 0.85967 | 270.3403 | 160.0403 | 2.2368 | 0.471 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR158
EN649 | 20200201 | MVCO | 2 | 1 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 144.9 | 83.14 | 0.93933 | 143.9607 | 82.2007 | 1.261 | 0.13929 | 20201113 | 10-78 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR163
EN649 | 20200201 | MVCO | 2 | 1 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 152 | 87.14 | 0.93933 | 151.0607 | 86.2007 | 1.3198 | 0.14365 | 20201113 | 10-78 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR164
EN649 | 20200201 | MVCO | 2 | 3 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 140.3 | 84.48 | 0.69 | 139.61 | 83.79 | 1.9654 | 0.69574 | [6 empty] | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200201 | MVCO | 2 | 3 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_012618_12mm | 1.902 | 0.48424 | 140.9 | 84.53 | 0.69 | 140.21 | 83.84 | 1.9848 | 0.67796 | [6 empty] | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200201 | MVCO | 2 | 3 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 75.47 | 48.12 | 0.72767 | 74.7423 | 47.3923 | 0.56419 | 0.25148 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | floating dark particle; picked particle off filter during extraction think it was zooop
EN649 | 20200201 | MVCO | 2 | 3 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 91.91 | 57.78 | 0.72767 | 91.1823 | 57.0523 | 0.70405 | 0.27787 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200201 | MVCO | 2 | 3 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 250.4 | 146.7 | 0.85967 | 249.5403 | 145.8403 | 2.1101 | 0.36583 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR155 | replication not great but within acceptable limits of size fractions
EN649 | 20200201 | MVCO | 2 | 3 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 226.1 | 134.5 | 0.85967 | 225.2403 | 133.6403 | 1.8639 | 0.40492 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR156 | replication not great but within acceptable limits of size fractions
EN649 | 20200201 | MVCO | 2 | 3 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 118.5 | 69.26 | 0.93933 | 117.5607 | 68.3207 | 1.0053 | 0.15848 | 20201113 | 10-78 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR161 | replication not great but within acceptable limits of size fractions
EN649 | 20200201 | MVCO | 2 | 3 |  |  | b | 292 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 149.6 | 88.49 | 0.93933 | 148.6607 | 87.5507 | 1.252 | 0.24453 | 20201113 | 10-78 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR162 | replication not great but within acceptable limits of size fractions
EN649 | 20200201 | MVCO | 2 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 250.6 | 150.8 | 0.69 | 249.91 | 150.11 | 2.0587 | 0.52481 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200201 | MVCO | 2 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 244 | 146.5 | 0.72767 | 243.2723 | 145.7723 | 2.0113 | 0.4976 | [6 empty] | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200201 | MVCO | 2 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 76.79 | 48.55 | 0.69 | 76.1 | 47.86 | 0.58255 | 0.24117 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200201 | MVCO | 2 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 63.31 | 39.84 | 0.72767 | 62.5823 | 39.1123 | 0.48415 | 0.18901 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200201 | MVCO | 2 | 6 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 219.9 | 130 | 0.85967 | 219.0403 | 129.1403 | 1.8293 | 0.36312 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR153
EN649 | 20200201 | MVCO | 2 | 6 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 244.5 | 145.3 | 0.85967 | 243.6403 | 144.4403 | 2.0254 | 0.4351 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR154
EN649 | 20200201 | MVCO | 2 | 6 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 132.5 | 76.99 | 0.93933 | 131.5607 | 76.0507 | 1.1334 | 0.16214 | 20201113 | 10-78 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR159
EN649 | 20200201 | MVCO | 2 | 6 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 132.8 | 76.3 | 0.93933 | 131.8607 | 75.3607 | 1.1536 | 0.13017 | 20201113 | 10-78 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR160
EN649 | 20200202 | L2 | 5 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.5 | Cal_111320b_12mm | 1.8343 | 0.5442 | 166.8 | 109.1 | 0.69 | 166.11 | 108.41 | 2.3805 | 1.3511 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | dilute by half with 5ml acetone before vortext etc
EN649 | 20200202 | L2 | 5 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 289.7 | 186.8 | 0.743 | 288.957 | 186.057 | 2.186 | 1.1272 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L2 | 5 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 248.7 | 157.2 | 0.69 | 248.01 | 156.51 | 1.8875 | 0.80617 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L2 | 5 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 273.1 | 169.6 | 0.743 | 272.357 | 168.857 | 2.1988 | 0.80819 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | Had a particle in extract
EN649 | 20200202 | L2 | 5 | 3 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 89.53 | 60.56 | 0.72767 | 88.8023 | 59.8323 | 0.59761 | 0.43216 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L2 | 5 | 3 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 85.67 | 57.78 | 0.72767 | 84.9423 | 57.0523 | 0.57533 | 0.40659 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L2 | 5 | 3 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 125 | 74.15 | 0.85967 | 124.1403 | 73.2903 | 1.0382 | 0.21026 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR13; logsheet notes there wer ea couple copepods on surf samples and did not pick off. This isn't a surf sample but was inccorectly labeled in logsheet comments as surf. Assume comment is for true surf samples
EN649 | 20200202 | L2 | 5 | 3 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 124.1 | 74.59 | 0.85967 | 123.2403 | 73.7303 | 1.0074 | 0.24428 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR14; logsheet notes there wer ea couple copepods on surf samples and did not pick off. This isn't a surf sample but was inccorectly labeled in logsheet comments as surf. Assume comment is for true surf samples
EN649 | 20200202 | L2 | 5 | 3 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 98.88 | 57.37 | 0.85967 | 98.0203 | 56.5103 | 0.84752 | 0.11512 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR19;during extraction sample too cold, some acetone bubbled out when added to vial, added a few more drops to match level of next sample tube; logsheet notes there wer ea couple copepods on surf samples and did not pick off. This isn't a surf sample but was inccorectly labeled in logsheet comments as surf. Assume comment is for true surf samples
EN649 | 20200202 | L2 | 5 | 3 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 90.23 | 53.37 | 0.85967 | 89.3703 | 52.5103 | 0.75258 | 0.14192 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR20; logsheet notes there wer ea couple copepods on surf samples and did not pick off. This isn't a surf sample but was inccorectly labeled in logsheet comments as surf. Assume comment is for true surf samples
EN649 | 20200202 | L2 | 5 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 228.7 | 140.4 | 0.75667 | 227.9433 | 139.6433 | 1.8215 | 0.5819 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L2 | 5 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 210 | 129.7 | 0.743 | 209.257 | 128.957 | 1.7059 | 0.59053 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L2 | 5 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 98.54 | 65.18 | 0.72767 | 97.8123 | 64.4523 | 0.68817 | 0.42112 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | ripped filter during extraction still with sample though
EN649 | 20200202 | L2 | 5 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 106.8 | 69.6 | 0.69 | 106.11 | 68.91 | 0.76738 | 0.41862 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L2 | 5 | 7 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 118.7 | 71.07 | 0.85967 | 117.8403 | 70.2103 | 0.9758 | 0.2243 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR15
EN649 | 20200202 | L2 | 5 | 7 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 135.3 | 78.64 | 0.93933 | 134.3607 | 77.7007 | 1.1529 | 0.1662 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR16
EN649 | 20200202 | L2 | 5 | 7 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 91.39 | 53.48 | 0.85967 | 90.5303 | 52.6203 | 0.77402 | 0.12235 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR21
EN649 | 20200202 | L2 | 5 | 7 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 102.4 | 59.42 | 0.85967 | 101.5403 | 58.5603 | 0.87455 | 0.11962 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR22
EN649 | 20200202 | L2 | 5 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 156 | 95.94 | 0.743 | 155.257 | 95.197 | 1.2759 | 0.41932 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L2 | 5 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 217.5 | 129.1 | 0.69 | 216.81 | 128.41 | 1.8236 | 0.3865 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L2 | 5 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 100.7 | 64.43 | 0.69 | 100.01 | 63.74 | 0.74819 | 0.34883 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L2 | 5 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 101.8 | 64.94 | 0.72767 | 101.0723 | 64.2123 | 0.76037 | 0.34479 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L2 | 5 | 12 |  |  | a | 290 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 183.9 | 110.8 | 0.93933 | 182.9607 | 109.8607 | 1.5079 | 0.38286 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR17; lost about 5ml when tightening loose manifold during ifltering; logsheet notes there wer ea couple copepods on surf samples and did not pick off excep comments of which is surface is incorrect, logsheet labeled N3 as surface
EN649 | 20200202 | L2 | 5 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 192.7 | 114.6 | 0.93933 | 191.7607 | 113.6607 | 1.5946 | 0.34159 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR18; logsheet notes there wer ea couple copepods on surf samples and did not pick off excep comments of which is surface is incorrect, logsheet labeled N3 as surface
EN649 | 20200202 | L2 | 5 | 12 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 153 | 90.41 | 0.85967 | 152.1403 | 89.5503 | 1.2779 | 0.24755 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR23; logsheet notes there wer ea couple copepods on surf samples and did not pick off excep comments of which is surface is incorrect, logsheet labeled N3 as surface
EN649 | 20200202 | L2 | 5 | 12 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 153.4 | 89.86 | 0.85967 | 152.5403 | 89.0003 | 1.2929 | 0.21804 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR24; logsheet notes there wer ea couple copepods on surf samples and did not pick off excep comments of which is surface is incorrect, logsheet labeled N3 as surface
EN649 | 20200202 | L3 | 6 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.5 | Cal_111320b_12mm | 1.8343 | 0.5442 | 158.9 | 99.22 | 0.69 | 158.21 | 98.53 | 2.4622 | 0.92937 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | dilute by half added 5ml acetone before vortex etc, when adding dilution tip of repipettor was missing a drop or 2 of its prime; poor replication
EN649 | 20200202 | L3 | 6 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.5 | Cal_111320b_12mm | 1.8343 | 0.5442 | 130.6 | 80.97 | 0.69 | 129.91 | 80.28 | 2.0476 | 0.7158 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | diluted by half, added 5ml acetone before vortex etc; poor replication
EN649 | 20200202 | L3 | 6 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 145.1 | 94.96 | 0.69 | 144.41 | 94.27 | 1.0343 | 0.58816 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L3 | 6 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 162.4 | 103.5 | 0.72767 | 161.6723 | 102.7723 | 1.215 | 0.55379 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L3 | 6 | 3 |  |  | a | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 143.5 | 90.71 | 0.85967 | 142.6403 | 89.8503 | 1.0705 | 0.44968 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR29 | did >5s and >20s get mixed up?; noDF doesn't think they got mixed up – chl signal is just mostly all in >20
EN649 | 20200202 | L3 | 6 | 3 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 126.7 | 81 | 0.85967 | 125.8403 | 80.1403 | 0.93307 | 0.4321 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR30 | did >5s and >20s get mixed up?; noDF doesn't think they got mixed up – chl signal is just mostly all in >20
EN649 | 20200202 | L3 | 6 | 3 |  |  | a | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 133.8 | 80.74 | 0.85967 | 132.9403 | 79.8803 | 1.0797 | 0.27645 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR35 | did >5s and >20s get mixed up?; noDF doesn't think they got mixed up – chl signal is just mostly all in >20
EN649 | 20200202 | L3 | 6 | 3 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 108 | 65.35 | 0.85967 | 107.1403 | 64.4903 | 0.8708 | 0.22778 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR36 | did >5s and >20s get mixed up?; noDF doesn't think they got mixed up – chl signal is just mostly all in >20
EN649 | 20200202 | L3 | 6 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 161.7 | 103.1 | 0.743 | 160.957 | 102.357 | 1.2449 | 0.57784 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L3 | 6 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 147.6 | 93.54 | 0.69 | 146.91 | 92.85 | 1.1152 | 0.48286 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L3 | 6 | 6 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 161.5 | 100.9 | 0.85967 | 160.6403 | 100.0403 | 1.2331 | 0.46528 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR27
EN649 | 20200202 | L3 | 6 | 6 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 113.1 | 72.06 | 0.85967 | 112.2403 | 71.2003 | 0.83792 | 0.37495 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR28 | lower than either >20 value ; QF was 1; could just be that most of the chl signal is in >20 fraction
EN649 | 20200202 | L3 | 6 | 6 |  |  | a | 292 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 131.6 | 79.92 | 0.93933 | 130.6607 | 78.9807 | 1.0588 | 0.29124 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR33
EN649 | 20200202 | L3 | 6 | 6 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 115.3 | 68.49 | 0.85967 | 114.4403 | 67.6303 | 0.95248 | 0.19566 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR34
EN649 | 20200202 | L3 | 6 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 164.6 | 104.3 | 0.69 | 163.91 | 103.61 | 1.2439 | 0.53933 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L3 | 6 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 162.8 | 104.1 | 0.743 | 162.057 | 103.357 | 1.247 | 0.59353 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L3 | 6 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 42.46 | 29.12 | 0.72767 | 41.7323 | 28.3923 | 0.27518 | 0.21347 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L3 | 6 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 42.99 | 30.29 | 0.72767 | 42.2623 | 29.5623 | 0.26198 | 0.24681 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L3 | 6 | 9 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 144.8 | 92.43 | 0.85967 | 143.9403 | 91.5703 | 1.0656 | 0.48895 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR25
EN649 | 20200202 | L3 | 6 | 9 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 140.1 | 87.68 | 0.85967 | 139.2403 | 86.8203 | 1.0703 | 0.40868 | 20201113 | 10-84 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR26
EN649 | 20200202 | L3 | 6 | 9 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 128.1 | 74.47 | 0.93933 | 127.1607 | 73.5307 | 1.095 | 0.15759 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR31 | higher than either >5 value; QF was 1; could just be that most of chl signal is in >20 fraction
EN649 | 20200202 | L3 | 6 | 9 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 125.9 | 77.14 | 0.93933 | 124.9607 | 76.2007 | 0.99555 | 0.30251 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR32
EN649 | 20200202 | L7 | 7 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 39.6 | 26.53 | 0.75667 | 38.8433 | 25.7733 | 0.26961 | 0.17397 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L7 | 7 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.92 | 27.9 | 0.743 | 40.177 | 27.157 | 0.2766 | 0.20701 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L7 | 7 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 31.23 | 21.76 | 0.72767 | 30.5023 | 21.0323 | 0.19535 | 0.16663 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L7 | 7 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 32.54 | 22.12 | 0.72767 | 31.8123 | 21.3923 | 0.21495 | 0.15323 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L7 | 7 | 6 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 41 | 27.57 | 0.69 | 40.31 | 26.88 | 0.2819 | 0.18884 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L7 | 7 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 39.84 | 26.77 | 0.72767 | 39.1123 | 26.0423 | 0.26961 | 0.1786 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L7 | 7 | 8 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 90.94 | 58.43 | 0.69 | 90.25 | 57.74 | 0.6824 | 0.3288 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | knocked fluor tube cover during first Ra reading but second reading is consistent 58.26
EN649 | 20200202 | L7 | 7 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 98.03 | 62.48 | 0.743 | 97.287 | 61.737 | 0.75524 | 0.34417 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L7 | 7 | 8 |  |  | a | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 18.79 | 11.08 | 0.93933 | 17.8507 | 10.1407 | 0.15635 | 0.015222 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR41 | did >5s and >20s get mixed up? >5 values are smaller than >20 values; QF was 1; DF thinks for this cast, 5s and 20s got mixed up
EN649 | 20200202 | L7 | 7 | 8 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 23.63 | 14.42 | 0.93933 | 22.6907 | 13.4807 | 0.18804 | 0.041596 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR42 | did >5s and >20s get mixed up? >5 values are smaller than >20 values; QF was 1; DF thinks for this cast, 5s and 20s got mixed up
EN649 | 20200202 | L7 | 7 | 8 |  |  | a | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 37.64 | 22.98 | 0.93933 | 36.7007 | 22.0407 | 0.2983 | 0.07588 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR47 | did >5s and >20s get mixed up? >5 values are smaller than >20 values; QF was 1; DF thinks for this cast, 5s and 20s got mixed up
EN649 | 20200202 | L7 | 7 | 8 |  |  | b | 295 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 34.07 | 20.79 | 0.93933 | 33.1307 | 19.8507 | 0.2693 | 0.066555 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR48 | did >5s and >20s get mixed up? >5 values are smaller than >20 values; QF was 1; DF thinks for this cast, 5s and 20s got mixed up
EN649 | 20200202 | L7 | 7 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 100.2 | 63.81 | 0.69 | 99.51 | 63.12 | 0.75067 | 0.33568 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L7 | 7 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 92.64 | 58.4 | 0.72767 | 91.9123 | 57.6723 | 0.70632 | 0.28627 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L7 | 7 | 11 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 18.27 | 11.22 | 0.93933 | 17.3307 | 10.2807 | 0.14394 | 0.031186 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR39 | >5 values are smaller than >20 values, potential mix-up?, QF was 1; DF thinks for this cast, 5s and 20s got mixed up
EN649 | 20200202 | L7 | 7 | 11 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 18.68 | 11.77 | 0.93933 | 17.7407 | 10.8307 | 0.1406 | 0.043266 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR40 | >5 values are smaller than >20 values, potential mix-up?, QF was 1; DF thinks for this cast, 5s and 20s got mixed up
EN649 | 20200202 | L7 | 7 | 11 |  |  | a | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 41.11 | 24.89 | 0.93933 | 40.1707 | 23.9507 | 0.33004 | 0.076563 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR45 | >5 values are smaller than >20 values, potential mix-up?, QF was 1; DF thinks for this cast, 5s and 20s got mixed up
EN649 | 20200202 | L7 | 7 | 11 |  |  | b | 292 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 27.47 | 17.45 | 0.93933 | 26.5307 | 16.5107 | 0.20528 | 0.076936 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR46 | >5 values are smaller than >20 values, potential mix-up?, QF was 1; DF thinks for this cast, 5s and 20s got mixed up
EN649 | 20200202 | L7 | 7 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 104.1 | 66.64 | 0.69 | 103.41 | 65.95 | 0.77274 | 0.36232 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L7 | 7 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 101.8 | 64.54 | 0.72767 | 101.0723 | 63.8123 | 0.76862 | 0.32965 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L7 | 7 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 114.3 | 73.21 | 0.743 | 113.557 | 72.467 | 0.87293 | 0.41755 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
EN649 | 20200202 | L7 | 7 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 111.7 | 72.86 | 0.81783 | 110.8822 | 72.0422 | 0.82513 | 0.45779 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
EN649 | 20200202 | L7 | 7 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 75.01 | 48.53 | 0.69 | 74.32 | 47.84 | 0.54624 | 0.27713 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200202 | L7 | 7 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 84.87 | 54.8 | 0.75667 | 84.1133 | 54.0433 | 0.6203 | 0.30984 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | LabNB says C7N1<10b but this is the only<10 on this cast; had to redo Rb with leftover volume from large test tube so particles in reading, possibly compromised
EN649 | 20200202 | L7 | 7 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 29.34 | 17.33 | 0.93933 | 28.4007 | 16.3907 | 0.24521 | 0.033999 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR37 | >5 values are smaller than >20 values, potential mix-up?, QF was 1; DF thinks for this cast, 5s and 20s got mixed up
EN649 | 20200202 | L7 | 7 | 16 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 22.65 | 14.04 | 0.93933 | 21.7107 | 13.1007 | 0.17519 | 0.047212 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR38 | >5 values are smaller than >20 values, potential mix-up?, QF was 1; DF thinks for this cast, 5s and 20s got mixed up
EN649 | 20200202 | L7 | 7 | 16 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 44.17 | 26.61 | 0.93933 | 43.2307 | 25.6707 | 0.35853 | 0.078765 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR43; Particles were visible in this solution. Pressed read three times for Rb and Ra to check if settling particles were affecting readings. Subsequent Rb=43.13, 43.67 and Ra=25.95, 26.17 so no big change. | >5 values are smaller than >20 values, potential mix-up?, QF was 1; DF thinks for this cast, 5s and 20s got mixed up
EN649 | 20200202 | L7 | 7 | 16 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 53.05 | 33.06 | 0.93933 | 52.1107 | 32.1207 | 0.40814 | 0.13902 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR44 | >5 values are smaller than >20 values, potential mix-up?, QF was 1; DF thinks for this cast, 5s and 20s got mixed up
EN649 | 20200203 | L8 | 8 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 10.59 | 8.194 | 0.72767 | 9.8623 | 7.4663 | 0.049426 | 0.079077 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.65 | 8.733 | 0.75667 | 10.8933 | 7.9763 | 0.060173 | 0.077107 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 20.43 | 15.21 | 0.75667 | 19.6733 | 14.4533 | 0.10768 | 0.14107 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 19.02 | 14.09 | 0.69 | 18.33 | 13.4 | 0.1017 | 0.12893 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 20.34 | 14.77 | 0.72767 | 19.6123 | 14.0423 | 0.1149 | 0.12678 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 21.37 | 15.43 | 0.75667 | 20.6133 | 14.6733 | 0.12253 | 0.13001 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 23.49 | 16.04 | 0.69 | 22.8 | 15.35 | 0.15368 | 0.11051 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 22.5 | 15.72 | 0.72767 | 21.7723 | 14.9923 | 0.13986 | 0.11817 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 14.47 | 10.41 | 0.75667 | 13.7133 | 9.6533 | 0.083752 | 0.082391 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.78 | 10.84 | 0.743 | 15.037 | 10.097 | 0.10495 | 0.074859 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 1.021 | 0.865 | 0.85967 | 0.16133 | 0.0053333 | 0.0031851 | -0.0030942 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR53 | no filter?? Couldn't see filter in tube when extracting and running; values so low assume filter never put in tube so throw out
EN649 | 20200203 | L8 | 8 | 6 |  |  | b | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.68 | 8.49 | 0.85967 | 10.8203 | 7.6303 | 0.064689 | 0.06441 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR54
EN649 | 20200203 | L8 | 8 | 6 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.936 | 3.563 | 0.85967 | 4.0763 | 2.7033 | 0.028033 | 0.018018 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR59
EN649 | 20200203 | L8 | 8 | 6 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 7.177 | 4.524 | 0.93933 | 6.2377 | 3.5847 | 0.053983 | 0.0068732 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR60
EN649 | 20200203 | L8 | 8 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 29.88 | 20.41 | 0.75667 | 29.1233 | 19.6533 | 0.19535 | 0.1429 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 27.53 | 18.44 | 0.72767 | 26.8023 | 17.7123 | 0.18751 | 0.11733 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 123.2 | 77.04 | 0.72767 | 122.4723 | 76.3123 | 0.95221 | 0.36119 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 135.3 | 85.94 | 0.75667 | 134.5433 | 85.1833 | 1.0182 | 0.44786 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 85.85 | 55.91 | 0.75667 | 85.0933 | 55.1533 | 0.61762 | 0.33162 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 91.75 | 59.69 | 0.75667 | 90.9933 | 58.9333 | 0.66135 | 0.35295 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 12 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 48.31 | 29.9 | 0.93933 | 47.3707 | 28.9607 | 0.37717 | 0.11786 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR51
EN649 | 20200203 | L8 | 8 | 12 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 42.78 | 26.86 | 0.85967 | 41.9203 | 26.0003 | 0.32394 | 0.11746 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR52
EN649 | 20200203 | L8 | 8 | 12 |  |  | a | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 27.76 | 16.36 | 0.93933 | 26.8207 | 15.4207 | 0.23197 | 0.029828 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR57
EN649 | 20200203 | L8 | 8 | 12 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 24.34 | 14.68 | 0.85967 | 23.4803 | 13.8203 | 0.19723 | 0.038194 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR58
EN649 | 20200203 | L8 | 8 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 153.6 | 97.58 | 0.75667 | 152.8433 | 96.8233 | 1.1556 | 0.51081 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 149.6 | 94.61 | 0.75667 | 148.8433 | 93.8533 | 1.1344 | 0.48094 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 134.5 | 85.23 | 0.72767 | 133.7723 | 84.5023 | 1.0164 | 0.438 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 143.3 | 90.23 | 0.743 | 142.557 | 89.487 | 1.1274 | 0.46614 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 89.8 | 58.17 | 0.69 | 89.11 | 57.48 | 0.65248 | 0.3368 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 100.3 | 64.4 | 0.743 | 99.557 | 63.657 | 0.76267 | 0.37092 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L8 | 8 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 51.38 | 31.43 | 0.93933 | 50.4407 | 30.4907 | 0.40732 | 0.11207 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR49
EN649 | 20200203 | L8 | 8 | 18 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 54.18 | 33.37 | 0.93933 | 53.2407 | 32.4307 | 0.42344 | 0.12713 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR50
EN649 | 20200203 | L8 | 8 | 18 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 35.87 | 21.95 | 0.93933 | 34.9307 | 21.0107 | 0.28421 | 0.073702 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR55
EN649 | 20200203 | L8 | 8 | 18 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 37.68 | 23.16 | 0.85967 | 36.8203 | 22.3003 | 0.29646 | 0.083421 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR56
EN649 | 20200203 | L9 | 11 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 15.08 | 10.61 | 0.72767 | 14.3523 | 9.8823 | 0.092209 | 0.077875 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | lots of little hairs on filter during extraction. Maybe someones fleece?
EN649 | 20200203 | L9 | 11 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 14.81 | 10.67 | 0.72767 | 14.0823 | 9.9423 | 0.085402 | 0.085715 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | lots of little hairs on filter during extraction. Maybe someones fleece?
EN649 | 20200203 | L9 | 11 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.76 | 11.33 | 0.743 | 15.017 | 10.587 | 0.094112 | 0.09442 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L9 | 11 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 14.78 | 10.84 | 0.75667 | 14.0233 | 10.0833 | 0.081276 | 0.092267 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L9 | 11 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 16.49 | 12 | 0.75667 | 15.7333 | 11.2433 | 0.092622 | 0.10089 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L9 | 11 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 18.75 | 13.42 | 0.78817 | 17.9618 | 12.6318 | 0.10995 | 0.10746 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L9 | 11 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 35.4 | 24.4 | 0.75667 | 34.6433 | 23.6433 | 0.22691 | 0.18001 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L9 | 11 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 38.8 | 26.17 | 0.69 | 38.11 | 25.48 | 0.26054 | 0.178 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L9 | 11 | 7 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 7.566 | 5.524 | 0.80333 | 6.7627 | 4.7207 | 0.04155 | 0.038591 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR65
EN649 | 20200203 | L9 | 11 | 7 |  |  | b | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 7.486 | 5.482 | 0.80333 | 6.6827 | 4.6787 | 0.040639 | 0.038521 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR66
EN649 | 20200203 | L9 | 11 | 7 |  |  | a | 284 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.286 | 2.394 | 0.80333 | 2.4827 | 1.5907 | 0.018789 | 0.0091659 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR71; logsheet "lost 9ml during pouring" not sure if it says 9 or 1, assume says 9
EN649 | 20200203 | L9 | 11 | 7 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 2.9 | 2.158 | 0.80333 | 2.0967 | 1.3547 | 0.01515 | 0.0079267 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR72
EN649 | 20200203 | L9 | 11 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 25.5 | 17.53 | 0.78817 | 24.7118 | 16.7418 | 0.16441 | 0.12373 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L9 | 11 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 26.91 | 18.54 | 0.75667 | 26.1533 | 17.7833 | 0.17266 | 0.13341 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L9 | 11 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 68.33 | 44.6 | 0.75667 | 67.5733 | 43.8433 | 0.48951 | 0.26507 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L9 | 11 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 65.93 | 43.14 | 0.75667 | 65.1733 | 42.3833 | 0.47012 | 0.25933 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L9 | 11 | 11 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 33.9 | 21.19 | 0.80333 | 33.0967 | 20.3867 | 0.25862 | 0.087479 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR63
EN649 | 20200203 | L9 | 11 | 11 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 29.17 | 18.13 | 0.80333 | 28.3667 | 17.3267 | 0.22541 | 0.069748 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR64
EN649 | 20200203 | L9 | 11 | 11 |  |  | a | 292 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 17.43 | 10.42 | 0.80333 | 16.6267 | 9.6167 | 0.14362 | 0.020763 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR69
EN649 | 20200203 | L9 | 11 | 11 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 18.1 | 11.05 | 0.80333 | 17.2967 | 10.2467 | 0.14345 | 0.030503 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR70
EN649 | 20200203 | L9 | 11 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 145 | 92.53 | 0.75667 | 144.2433 | 91.7733 | 1.0824 | 0.49713 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L9 | 11 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 152.1 | 94.22 | 0.75667 | 151.3433 | 93.4633 | 1.194 | 0.41461 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | extra acid when doing Ra
EN649 | 20200203 | L9 | 11 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 151.1 | 94.67 | 0.78817 | 150.3118 | 93.8818 | 1.1641 | 0.45173 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L9 | 11 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 134 | 84.28 | 0.743 | 133.257 | 83.537 | 1.0563 | 0.43135 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L9 | 11 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 95.39 | 61.88 | 0.81783 | 94.5722 | 61.0622 | 0.7119 | 0.37549 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
EN649 | 20200203 | L9 | 11 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 92.34 | 59.77 | 0.69 | 91.65 | 59.08 | 0.67187 | 0.34495 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
EN649 | 20200203 | L9 | 11 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 47.98 | 29.94 | 0.80333 | 47.1767 | 29.1367 | 0.36833 | 0.12801 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR61
EN649 | 20200203 | L9 | 11 | 15 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 52.77 | 32.55 | 0.80333 | 51.9667 | 31.7467 | 0.41284 | 0.12796 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR62
EN649 | 20200203 | L9 | 11 | 15 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 28.17 | 16.98 | 0.80333 | 27.3667 | 16.1767 | 0.22847 | 0.047095 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR67
EN649 | 20200203 | L9 | 11 | 15 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 34.28 | 20.37 | 0.80333 | 33.4767 | 19.5667 | 0.28304 | 0.04914 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR68
EN649 | 20200203 | L10 | 12 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.681 | 3.842 | 0.78817 | 3.8928 | 3.0538 | 0.017307 | 0.035252 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.132 | 3.152 | 0.75667 | 3.3753 | 2.3953 | 0.020216 | 0.02101 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 14.43 | 10.42 | 0.69 | 13.74 | 9.73 | 0.08272 | 0.084742 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 14.05 | 9.948 | 0.78817 | 13.2618 | 9.1598 | 0.084618 | 0.073031 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 27.39 | 18.79 | 0.78817 | 26.6018 | 18.0018 | 0.1774 | 0.13242 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 24.87 | 17.72 | 0.72767 | 24.1423 | 16.9923 | 0.14749 | 0.14496 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 6 |  |  | a | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.477 | 3.612 | 0.80333 | 3.6737 | 2.8087 | 0.017541 | 0.029979 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR77
EN649 | 20200203 | L10 | 12 | 6 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.388 | 3.28 | 0.80333 | 3.5847 | 2.4767 | 0.022622 | 0.019567 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR78
EN649 | 20200203 | L10 | 12 | 6 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 2.214 | 2.317 | 0.80333 | 1.4107 | 1.5137 | -0.002103 | 0.027888 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR83 | was this acidified?
EN649 | 20200203 | L10 | 12 | 6 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 1.904 | 1.571 | 0.80333 | 1.1007 | 0.76767 | 0.0067758 | 0.0062567 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR84
EN649 | 20200203 | L10 | 12 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 57.5 | 39.33 | 0.78817 | 56.7118 | 38.5418 | 0.37482 | 0.28852 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 53.02 | 35.26 | 0.72767 | 52.2923 | 34.5323 | 0.36636 | 0.22797 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 45.84 | 30.99 | 0.78817 | 45.0518 | 30.2018 | 0.30633 | 0.21347 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 10 |  |  | b | 280 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 44.8 | 30.78 | 0.75667 | 44.0433 | 30.0233 | 0.29954 | 0.23564 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | hair in sample
EN649 | 20200203 | L10 | 12 | 10 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.61 | 7.986 | 0.80333 | 10.8067 | 7.1827 | 0.074246 | 0.048528 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR85
EN649 | 20200203 | L10 | 12 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.97 | 8.191 | 0.80333 | 11.1667 | 7.3877 | 0.076894 | 0.048524 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR86 | said TAR85 on vial but I guess 2 labeled 85 so one must be 86?
EN649 | 20200203 | L10 | 12 | 10 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.621 | 2.64 | 0.80333 | 2.8177 | 1.8367 | 0.020029 | 0.011258 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR87
EN649 | 20200203 | L10 | 12 | 10 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 2.922 | 2.111 | 0.80333 | 2.1187 | 1.3077 | 0.016502 | 0.0056979 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR88; filter fell onto bench top when putting in tube
EN649 | 20200203 | L10 | 12 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 65.35 | 43.08 | 0.78817 | 64.5618 | 42.2918 | 0.4594 | 0.26848 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 62.22 | 41.45 | 0.78817 | 61.4318 | 40.6618 | 0.42845 | 0.27137 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 41.38 | 27.46 | 0.72767 | 40.6523 | 26.7323 | 0.28715 | 0.17294 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | poor replication; during extraction filter seems a bit wet. Possibly diluted chl value so toss this value and high confidence in other sample
EN649 | 20200203 | L10 | 12 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 54.13 | 36.92 | 0.75667 | 53.3733 | 36.1633 | 0.35502 | 0.26739 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | poor replication
EN649 | 20200203 | L10 | 12 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 58.15 | 39.01 | 0.78817 | 57.3618 | 38.2218 | 0.39483 | 0.263 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 46.73 | 31.13 | 0.75667 | 45.9733 | 30.3733 | 0.3218 | 0.20095 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 15 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 12.88 | 8.791 | 0.80333 | 12.0767 | 7.9877 | 0.083202 | 0.052402 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR75
EN649 | 20200203 | L10 | 12 | 15 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.85 | 8.101 | 0.80333 | 11.0467 | 7.2977 | 0.076284 | 0.047607 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR76
EN649 | 20200203 | L10 | 12 | 15 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.56 | 3.06 | 0.80333 | 3.7567 | 2.2567 | 0.030626 | 0.0078157 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR81
EN649 | 20200203 | L10 | 12 | 15 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.454 | 2.314 | 0.80333 | 2.6507 | 1.5107 | 0.023276 | 0.0024581 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR82
EN649 | 20200203 | L10 | 12 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 58.46 | 39.55 | 0.78817 | 57.6718 | 38.7618 | 0.39008 | 0.27704 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 57.36 | 38.52 | 0.78817 | 56.5718 | 37.7318 | 0.38864 | 0.26076 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 43.21 | 29.95 | 0.743 | 42.467 | 29.207 | 0.2817 | 0.23841 | 20220519 | 11-95 | 20221206 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 43.48 | 28.75 | 0.69 | 42.79 | 28.06 | 0.30386 | 0.17908 | 20201113 | 10-96 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200203 | L10 | 12 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.77 | 8.166 | 0.80333 | 10.9667 | 7.3627 | 0.073584 | 0.051837 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR73
EN649 | 20200203 | L10 | 12 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.62 | 8.128 | 0.80333 | 10.8167 | 7.3247 | 0.071297 | 0.053476 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR74
EN649 | 20200203 | L10 | 12 | 19 |  |  | a | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.62 | 2.704 | 0.80333 | 2.8167 | 1.9007 | 0.018639 | 0.013629 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR79
EN649 | 20200203 | L10 | 12 | 19 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.468 | 2.393 | 0.80333 | 2.6647 | 1.5897 | 0.021949 | 0.0051309 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR80 | says 79 on vial so I guess 2 were labeled TAR79 and this one assumed to be TAR80
EN649 | 20200204 | L11 | 16 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 27.49 | 18.98 | 0.72767 | 26.7623 | 18.2523 | 0.17555 | 0.13859 | 20201113 | 10-98 | 20200821 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200204 | L11 | 16 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 27.63 | 19.29 | 0.75667 | 26.8733 | 18.5333 | 0.17204 | 0.14693 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200204 | L11 | 16 | 6 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 6.365 | 4.317 | 0.80333 | 5.5617 | 3.5137 | 0.041958 | 0.018101 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR101
EN649 | 20200204 | L11 | 16 | 6 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 6.614 | 4.824 | 0.80333 | 5.8107 | 4.0207 | 0.036423 | 0.031835 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR102; chl/phaeo ratio is high
EN649 | 20200204 | L11 | 16 | 6 |  |  | a | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 2.938 | 1.998 | 0.80333 | 2.1347 | 1.1947 | 0.019127 | 0.0011546 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR103 | somewhat poor replication; QF was 1; DF thinks its just in noise of flourometer and low values thus high variability
EN649 | 20200204 | L11 | 16 | 6 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 2.174 | 1.763 | 0.80333 | 1.3707 | 0.95967 | 0.0083915 | 0.0079561 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR104 | somewhat poor replication; QF was 1; DF thinks its just in noise of flourometer and low values thus high variability
EN649 | 20200204 | L11 | 16 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 47.02 | 32.26 | 0.75667 | 46.2633 | 31.5033 | 0.30448 | 0.23772 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200204 | L11 | 16 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 49.63 | 33.42 | 0.78817 | 48.8418 | 32.6318 | 0.33439 | 0.22724 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200204 | L11 | 16 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 39.6 | 26.67 | 0.78817 | 38.8118 | 25.8818 | 0.26673 | 0.17872 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200204 | L11 | 16 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 46.2 | 31.2 | 0.9465 | 45.2535 | 30.2535 | 0.30943 | 0.21126 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200204 | L11 | 16 | 8 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 10.77 | 7.634 | 0.80333 | 9.9667 | 6.8307 | 0.063811 | 0.052152 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR93
EN649 | 20200204 | L11 | 16 | 8 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.54 | 7.962 | 0.80333 | 10.7367 | 7.1587 | 0.073053 | 0.048893 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR94
EN649 | 20200204 | L11 | 16 | 8 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.819 | 3.401 | 0.80333 | 4.0157 | 2.5977 | 0.028952 | 0.015299 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR99
EN649 | 20200204 | L11 | 16 | 8 |  |  | b | 295 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.631 | 2.432 | 0.80333 | 2.8277 | 1.6287 | 0.024314 | 0.0032414 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR100
EN649 | 20200204 | L11 | 16 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 62.28 | 41.7 | 0.78817 | 61.4918 | 40.9118 | 0.42453 | 0.2796 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200204 | L11 | 16 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 61.21 | 40.91 | 0.75667 | 60.4533 | 40.1533 | 0.41876 | 0.27232 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200204 | L11 | 16 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 48.4 | 31.7 | 0.75667 | 47.6433 | 30.9433 | 0.3445 | 0.18807 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | assume vol filt 290ml, not entered on logsheet
EN649 | 20200204 | L11 | 16 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 55.91 | 37.15 | 0.9465 | 54.9635 | 36.2035 | 0.38699 | 0.23611 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | assume vol filt 290ml, not entered on logsheet
EN649 | 20200204 | L11 | 16 | 13 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 13.72 | 9.269 | 0.80333 | 12.9167 | 8.4657 | 0.090568 | 0.053151 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR91
EN649 | 20200204 | L11 | 16 | 13 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 15.41 | 10.07 | 0.80333 | 14.6067 | 9.2667 | 0.10866 | 0.048661 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR92
EN649 | 20200204 | L11 | 16 | 13 |  |  | a | 292 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.825 | 3.219 | 0.80333 | 4.0217 | 2.4157 | 0.032902 | 0.0083887 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR97
EN649 | 20200204 | L11 | 16 | 13 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 5.592 | 4.231 | 0.80333 | 4.7887 | 3.4277 | 0.027788 | 0.030601 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR98
EN649 | 20200204 | L11 | 16 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 76.21 | 49.18 | 0.75667 | 75.4533 | 48.4233 | 0.55759 | 0.27582 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200204 | L11 | 16 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 75.47 | 48.48 | 0.78817 | 74.6818 | 47.6918 | 0.55676 | 0.26406 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200204 | L11 | 16 | 17 |  |  | a | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN649 | 20200204 | L11 | 16 | 17 |  |  | b | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN649 | 20200204 | L11 | 16 | 21 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 132.9 | 82.53 | 0.75667 | 132.1433 | 81.7733 | 1.0391 | 0.36834 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200204 | L11 | 16 | 21 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 138.4 | 86.26 | 0.78817 | 137.6118 | 85.4718 | 1.0756 | 0.39548 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200204 | L11 | 16 | 21 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 102.8 | 65.72 | 0.75667 | 102.0433 | 64.9633 | 0.7649 | 0.35317 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200204 | L11 | 16 | 21 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 101.8 | 63.95 | 0.78817 | 101.0118 | 63.1618 | 0.78079 | 0.30629 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200204 | L11 | 16 | 21 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 29.1 | 18.13 | 0.80333 | 28.2967 | 17.3267 | 0.22398 | 0.071177 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR89
EN649 | 20200204 | L11 | 16 | 21 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 26.67 | 16.85 | 0.80333 | 25.8667 | 16.0467 | 0.2005 | 0.072853 | 20201113 | 10-86 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR90
EN649 | 20200204 | L11 | 16 | 21 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 18.27 | 11.12 | 0.80333 | 17.4667 | 10.3167 | 0.14598 | 0.029758 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR95
EN649 | 20200204 | L11 | 16 | 21 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 14.31 | 8.896 | 0.80333 | 13.5067 | 8.0927 | 0.11016 | 0.027224 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR96
EN649 | 20200205 | L6 | 19 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 59.56 | 40.82 | 0.78817 | 58.7718 | 40.0318 | 0.38658 | 0.30241 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 62.34 | 42.3 | 0.9465 | 61.3935 | 41.3535 | 0.41339 | 0.29834 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 49.95 | 34.63 | 0.9465 | 49.0035 | 33.6835 | 0.31603 | 0.2637 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 50.13 | 34.31 | 0.78817 | 49.3418 | 33.5218 | 0.32634 | 0.2506 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 80.92 | 53.83 | 0.9465 | 79.9735 | 52.8835 | 0.55883 | 0.35135 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 73.36 | 50.47 | 0.75667 | 72.6033 | 49.7133 | 0.47219 | 0.38343 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 79.03 | 51.13 | 0.78817 | 78.2418 | 50.3418 | 0.57553 | 0.29089 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 90.61 | 58.7 | 0.78817 | 89.8218 | 57.9118 | 0.65825 | 0.33846 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 4 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 49.31 | 30.82 | 0.80333 | 48.5067 | 30.0167 | 0.37881 | 0.13427 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR111
EN649 | 20200205 | L6 | 19 | 4 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 56.13 | 35.98 | 0.80333 | 55.3267 | 35.1767 | 0.41001 | 0.18718 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR112
EN649 | 20200205 | L6 | 19 | 4 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 39.72 | 24.8 | 0.80333 | 38.9167 | 23.9967 | 0.30463 | 0.10415 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR119
EN649 | 20200205 | L6 | 19 | 4 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 48.18 | 29.52 | 0.80333 | 47.3767 | 28.7167 | 0.37969 | 0.10783 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR120; barely higher than one of the 5 reps, but DF thinks just biolgoical variation
EN649 | 20200205 | L6 | 19 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 102.3 | 65.32 | 0.75667 | 101.5433 | 64.5633 | 0.76284 | 0.34835 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 106.2 | 68.38 | 0.9465 | 105.2535 | 67.4335 | 0.78017 | 0.38042 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 50.78 | 33.62 | 0.78817 | 49.9918 | 32.8318 | 0.35398 | 0.21108 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 48.06 | 31.94 | 0.78817 | 47.2718 | 31.1518 | 0.33253 | 0.20362 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 78.92 | 47.98 | 0.80333 | 78.1167 | 47.1767 | 0.63171 | 0.17193 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR109
EN649 | 20200205 | L6 | 19 | 9 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 59.7 | 37.24 | 0.80333 | 58.8967 | 36.4367 | 0.45701 | 0.16156 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR110
EN649 | 20200205 | L6 | 19 | 9 |  |  | a | 295 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 39.32 | 23.91 | 0.80333 | 38.5167 | 23.1067 | 0.3125 | 0.07845 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR117
EN649 | 20200205 | L6 | 19 | 9 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 59.67 | 37.13 | 0.80333 | 58.8667 | 36.3267 | 0.4602 | 0.15861 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR118
EN649 | 20200205 | L6 | 19 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 143.5 | 90.9 | 0.9465 | 142.5535 | 89.9535 | 1.0851 | 0.46312 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 147.8 | 93.23 | 0.9465 | 146.8535 | 92.2835 | 1.1257 | 0.46259 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 83.3 | 53.92 | 0.78817 | 82.5118 | 53.1318 | 0.60606 | 0.30838 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 84.83 | 55.23 | 0.78817 | 84.0418 | 54.4418 | 0.6106 | 0.32639 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 14 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 74.71 | 46.45 | 0.80333 | 73.9067 | 45.6467 | 0.57503 | 0.1999 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR107
EN649 | 20200205 | L6 | 19 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 81.5 | 49.12 | 0.80333 | 80.6967 | 48.3167 | 0.66111 | 0.16195 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR108
EN649 | 20200205 | L6 | 19 | 14 |  |  | a | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 70.08 | 42.85 | 0.80333 | 69.2767 | 42.0467 | 0.55407 | 0.15974 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR115
EN649 | 20200205 | L6 | 19 | 14 |  |  | b | 292 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 55.93 | 33.42 | 0.80333 | 55.1267 | 32.6167 | 0.46117 | 0.096351 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR116
EN649 | 20200205 | L6 | 19 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 127.1 | 79.57 | 0.75667 | 126.3433 | 78.8133 | 0.98047 | 0.37598 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 126.9 | 80.1 | 0.75667 | 126.1433 | 79.3433 | 0.96541 | 0.40016 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 91.57 | 59.12 | 0.9465 | 90.6235 | 58.1735 | 0.66939 | 0.33183 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 84.16 | 53.61 | 0.78817 | 83.3718 | 52.8218 | 0.6302 | 0.27891 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L6 | 19 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 78.85 | 47.7 | 0.80333 | 78.0467 | 46.8967 | 0.636 | 0.16287 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR105
EN649 | 20200205 | L6 | 19 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 73.39 | 44.6 | 0.80333 | 72.5867 | 43.7967 | 0.58781 | 0.15825 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR106
EN649 | 20200205 | L6 | 19 | 19 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 46.45 | 28.03 | 0.80333 | 45.6467 | 27.2267 | 0.37609 | 0.087712 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR113; poor replication but within biological variation
EN649 | 20200205 | L6 | 19 | 19 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 81.27 | 48.38 | 0.80333 | 80.4667 | 47.5767 | 0.66924 | 0.13846 | 20201113 | 10-87 | 20200706 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR114; poor replication but within biological variation
EN649 | 20200205 | L5 | 20 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 200.1 | 123.1 | 0.78817 | 199.3118 | 122.3118 | 1.5884 | 0.51671 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L5 | 20 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 222.2 | 134.3 | 0.78817 | 221.4118 | 133.5118 | 1.8132 | 0.48462 | 20201113 | 11-23 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L5 | 20 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 187.8 | 116.8 | 0.9465 | 186.8535 | 115.8535 | 1.4646 | 0.52932 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L5 | 20 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 191.5 | 118.8 | 0.9465 | 190.5535 | 117.8535 | 1.4997 | 0.52868 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | sample had a lot of water/ice when going to extract
EN649 | 20200205 | L5 | 20 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 207.1 | 125.6 | 0.9465 | 206.1535 | 124.6535 | 1.6812 | 0.46418 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L5 | 20 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 192.6 | 115.9 | 0.78817 | 191.8118 | 115.1118 | 1.5822 | 0.39898 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L5 | 20 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 209.7 | 126.9 | 0.9465 | 208.7535 | 125.9535 | 1.708 | 0.45974 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L5 | 20 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 199.2 | 121 | 0.9465 | 198.2535 | 120.0535 | 1.6131 | 0.45308 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L5 | 20 | 10 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 197.3 | 113.6 | 0.93933 | 196.3607 | 112.6607 | 1.7031 | 0.2095 | 20201113 | 10-78 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR125
EN649 | 20200205 | L5 | 20 | 10 |  |  | b | 292 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 219.5 | 125.2 | 0.93933 | 218.5607 | 124.2607 | 1.9319 | 0.19205 | 20201113 | 10-78 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR126 | value larger than either WSW value; QF was 1; DF thinks just most of chl signal is in >5 fraction so that's why this rep and WSW are close
EN649 | 20200205 | L5 | 20 | 10 |  |  | a | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 156.1 | 88.77 | 0.93933 | 155.1607 | 87.8307 | 1.37 | 0.12106 | 20201113 | 10-78 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR131
EN649 | 20200205 | L5 | 20 | 10 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 154.8 | 89.44 | 0.93933 | 153.8607 | 88.5007 | 1.3345 | 0.17311 | 20201113 | 10-78 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR132; spilled ~2ml when pouring
EN649 | 20200205 | L5 | 20 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 211.8 | 129.4 | 0.9465 | 210.8535 | 128.4535 | 1.6998 | 0.51102 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L5 | 20 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 280.1 | 169.7 | 0.96783 | 279.1322 | 168.7322 | 2.2774 | 0.62665 | 20201113 | 11-14 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L5 | 20 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 127.5 | 74.03 | 0.93933 | 126.5607 | 73.0907 | 1.0917 | 0.15337 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR123
EN649 | 20200205 | L5 | 20 | 12 |  |  | b | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 88.36 | 51.12 | 0.93933 | 87.4207 | 50.1807 | 0.75518 | 0.093833 | 20201113 | 10-78 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR124 | smaller than one >20 rep, was this switched with >20 repa?; QF was 1 and DF changed to 2; DF thinks this is >20 rep b
EN649 | 20200205 | L5 | 20 | 12 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 119.9 | 69.95 | 0.85967 | 119.0403 | 69.0903 | 1.0198 | 0.15709 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR129 | larger than one of the >5 reps, could this have been switched with >5 rep b?; QF was 1 and DF changed to 2; DF thinks this is >5 rep a
EN649 | 20200205 | L5 | 20 | 12 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 83.1 | 48.15 | 0.93933 | 82.1607 | 47.2107 | 0.71116 | 0.090328 | 20201113 | 10-78 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR130
EN649 | 20200205 | L5 | 20 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 184.7 | 114.3 | 0.9465 | 183.7535 | 113.3535 | 1.4522 | 0.49867 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L5 | 20 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 161.3 | 97.23 | 0.75667 | 160.5433 | 96.4733 | 1.3217 | 0.33873 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L5 | 20 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 67.54 | 43.98 | 0.78817 | 66.7518 | 43.1918 | 0.48601 | 0.25736 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L5 | 20 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 75.85 | 48.27 | 0.78817 | 75.0618 | 47.4818 | 0.56893 | 0.24827 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L5 | 20 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 111 | 64.8 | 0.93933 | 110.0607 | 63.8607 | 0.94328 | 0.14457 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR121
EN649 | 20200205 | L5 | 20 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 68.24 | 40.67 | 0.93933 | 67.3007 | 39.7307 | 0.56099 | 0.11351 | 20201113 | 10-77 |  |  |  | KM | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR122; Was surprised at the low readings, so pipetted more extraction solution into a new tube. Rb=74.74 and Ra=44.63, so first readings seem correct. Second round included particles.
EN649 | 20200205 | L5 | 20 | 14 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 100.8 | 58.61 | 0.85967 | 99.9403 | 57.7503 | 0.8614 | 0.12236 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR127
EN649 | 20200205 | L5 | 20 | 14 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 142.8 | 82.94 | 0.85967 | 141.9403 | 82.0803 | 1.2222 | 0.17604 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR128 | odd this is the highest value of all the >5 & >20 from this depth; DF agrees this is odd it's the highest one, could it have been switched with >5 rep b?; QF was 1
EN649 | 20200205 | L4 | 21 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 328.6 | 196.6 | 0.9465 | 327.6535 | 195.6535 | 2.723 | 0.64442 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 389.5 | 233.2 | 0.9465 | 388.5535 | 232.2535 | 3.2242 | 0.77306 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 212.3 | 126.4 | 0.96783 | 211.3322 | 125.4322 | 1.772 | 0.38682 | 20201113 | 11-14 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 247.7 | 150 | 0.9465 | 246.7535 | 149.0535 | 2.0154 | 0.54995 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 222.2 | 135.8 | 0.9465 | 221.2535 | 134.8535 | 1.7823 | 0.53865 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | filter may have had large gelatinous thing. Wasn't sure if fuzz or ice so did not remove
EN649 | 20200205 | L4 | 21 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 184.1 | 112.2 | 0.9465 | 183.1535 | 111.2535 | 1.4832 | 0.43159 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 52.43 | 34.56 | 0.9465 | 51.4835 | 33.6135 | 0.36863 | 0.20989 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | possibly big things on <10
EN649 | 20200205 | L4 | 21 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 72.24 | 46.08 | 0.96783 | 71.2722 | 45.1122 | 0.53964 | 0.23678 | 20201113 | 11-14 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | possibly big things on <10
EN649 | 20200205 | L4 | 21 | 5 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 152 | 89.5 | 0.80333 | 151.1967 | 88.6967 | 1.2761 | 0.23484 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR137; pump running @15psi
EN649 | 20200205 | L4 | 21 | 5 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 191.6 | 113.1 | 0.80333 | 190.7967 | 112.2967 | 1.6028 | 0.31018 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR138; pump running @15psi
EN649 | 20200205 | L4 | 21 | 5 |  |  | a | 295 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 126.9 | 74.41 | 0.85967 | 126.0403 | 73.5503 | 1.0644 | 0.17998 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR143; pump running @15psi
EN649 | 20200205 | L4 | 21 | 5 |  |  | b | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 142.6 | 82.94 | 0.85967 | 141.7403 | 82.0803 | 1.214 | 0.17951 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR144; pump running @15psi
EN649 | 20200205 | L4 | 21 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 152.9 | 94.43 | 0.96783 | 151.9322 | 93.4622 | 1.2061 | 0.40242 | 20201113 | 11-14 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 134 | 85.7 | 0.9465 | 133.0535 | 84.7535 | 0.99635 | 0.46233 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 186.2 | 117.1 | 0.96783 | 185.2322 | 116.1322 | 1.4254 | 0.57331 | 20201113 | 11-14 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 177.5 | 108.4 | 0.75667 | 176.7433 | 107.6433 | 1.4254 | 0.42721 | 20201113 | 11-25 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 61.53 | 40.36 | 0.9465 | 60.5835 | 39.4135 | 0.4367 | 0.24164 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 60.54 | 39.51 | 0.96783 | 59.5722 | 38.5422 | 0.43382 | 0.22953 | 20201113 | 11-14 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 124.1 | 75.31 | 0.80333 | 123.2967 | 74.5067 | 0.99616 | 0.27304 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR135; pump running @15psi
EN649 | 20200205 | L4 | 21 | 12 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 149.2 | 88.79 | 0.80333 | 148.3967 | 87.9867 | 1.2292 | 0.26452 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR136; pump running @15psi
EN649 | 20200205 | L4 | 21 | 12 |  |  | a | 294 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 86.94 | 51.79 | 0.85967 | 86.0803 | 50.9303 | 0.71523 | 0.14941 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR141; pump running @15psi
EN649 | 20200205 | L4 | 21 | 12 |  |  | b | 292 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 130.8 | 76.67 | 0.85967 | 129.9403 | 75.8103 | 1.109 | 0.18686 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR142; pump running @15psi
EN649 | 20200205 | L4 | 21 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 205.8 | 128 | 0.9465 | 204.8535 | 127.0535 | 1.6049 | 0.58181 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 185.7 | 119.1 | 0.9465 | 184.7535 | 118.1535 | 1.3739 | 0.65967 | 20201113 | 11-22 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 57.92 | 38.21 | 0.78817 | 57.1318 | 37.4218 | 0.40659 | 0.23748 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 54.88 | 36.11 | 0.78817 | 54.0918 | 35.3218 | 0.3872 | 0.22072 | 20201113 | 11-24 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN649 | 20200205 | L4 | 21 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 171.2 | 102.1 | 0.80333 | 170.3967 | 101.2967 | 1.4108 | 0.31472 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR133; pump running @15psi
EN649 | 20200205 | L4 | 21 | 18 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 221.2 | 134.2 | 0.80333 | 220.3967 | 133.3967 | 1.7703 | 0.49438 | 20201113 | 10-87 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR134; pump running @15psi | this value is higher than both WSW values at this depth, QF was 1; could be that most of chl signal is in >5 fraction
EN649 | 20200205 | L4 | 21 | 18 |  |  | a | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 165.9 | 98.75 | 0.85967 | 165.0403 | 97.8903 | 1.371 | 0.29651 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR139; pump running @15psi
EN649 | 20200205 | L4 | 21 | 18 |  |  | b | 293 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 131.8 | 78.51 | 0.85967 | 130.9403 | 77.6503 | 1.088 | 0.23471 | 20201113 | 10-85 |  |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR140; pump running @15psi
AR44 | 20200608 | L10 | 6 | 5 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 200.4 | 124.1 | 0.67367 | 199.7263 | 123.4263 | 2.6683 | 0.94987 | 20220519 | 13-18 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/01; Offshore
AR44 | 20200608 | L10 | 6 | 6 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 133.7 | 83.16 | 0.67367 | 133.0263 | 82.4863 | 0.29486 | 0.10854 | 20220519 | 13-18 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 |  | OOI 06/02; Offshore; Flagged bad. Should be chl max higher concentration. Likely sampled from surface Niskin. Confirmed was not diluted during analysis
AR44 | 20200608 | L10 | 6 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 133.6 | 83.74 | 0.67367 | 132.9263 | 83.0663 | 0.29089 | 0.11534 | 20220519 | 13-18 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/03; Offshore
AR44 | 20200608 | L10 | 6 | 8 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 137.8 | 84.9 | 0.67367 | 137.1263 | 84.2263 | 0.30863 | 0.10328 | 20220519 | 13-18 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/04; Offshore
AR44 | 20200609 | L12 | 8 | 5 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 277.9 | 178.2 | 0.67367 | 277.2263 | 177.5263 | 1.1622 | 0.57249 | 20220519 | 13-18 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 08/01; UpstreamOffshore
AR44 | 20200609 | L12 | 8 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 184.1 | 114.6 | 0.67367 | 183.4263 | 113.9263 | 0.40547 | 0.15167 | 20220519 | 13-18 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 08/02; UpstreamOffshore
AR44 | 20200609 | L12 | 8 | 8 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 191.6 | 118.3 | 0.67367 | 190.9263 | 117.6263 | 0.42764 | 0.1476 | 20220519 | 13-18 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 08/03; UpstreamOffshore
AR44 | 20200610 | L9 | 9 | 5 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 166.3 | 104 | 0.67367 | 165.6263 | 103.3263 | 2.1787 | 0.85025 | 20220519 | 13-18 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 09/01; CentralOffshore; poor replication | poor replication
AR44 | 20200610 | L9 | 9 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 119.8 | 75.7 | 0.67367 | 119.1263 | 75.0263 | 1.5437 | 0.65775 | 20220519 | 13-18 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 09/02; CentralOffshore; poor replication | poor replication
AR44 | 20200610 | L9 | 9 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 148.8 | 94.53 | 0.67367 | 148.1263 | 93.8563 | 0.31662 | 0.14238 | 20220519 | 13-18 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 09/03; CentralOffshore
AR44 | 20200610 | L9 | 9 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 93.35 | 58.01 | 0.67367 | 92.6763 | 57.3363 | 0.41236 | 0.14844 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 09/04; CentralOffshore
AR44 | 20200612 | L7 | 23 | 3 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.34 | 42.69 | 0.67367 | 51.6663 | 42.0163 | 0.056246 | 0.14904 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 23/01; CentralInshore | phae>chl; depth 74m
AR44 | 20200612 | L7 | 23 | 4 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 54.44 | 43.93 | 0.67367 | 53.7663 | 43.2563 | 0.061317 | 0.15022 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 23/02; CentralInshore | phae>chl; depth 74m
AR44 | 20200612 | L7 | 23 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 208.4 | 134.2 | 0.67367 | 207.7263 | 133.5263 | 2.5974 | 1.3206 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 23/03; CentralInshore
AR44 | 20200612 | L7 | 23 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 204.4 | 129.1 | 0.67367 | 203.7263 | 128.4263 | 2.6359 | 1.1325 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 23/04; CentralInshore
AR44 | 20200612 | L7 | 23 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 140.7 | 87.87 | 0.67367 | 140.0263 | 87.1963 | 0.30793 | 0.1181 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 23/05; CentralInshore
AR44 | 20200612 | L7 | 23 | 8 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 149.1 | 92.2 | 0.67367 | 148.4263 | 91.5263 | 0.33196 | 0.11564 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 23/06; CentralInshore
AR44 | 20200613 | L13 | 24 | 3 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 170.1 | 124.5 | 0.67367 | 169.4263 | 123.8263 | 0.26579 | 0.3392 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 24/01; UpstreamInshore | phae>chl; depth 57m
AR44 | 20200613 | L13 | 24 | 4 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 166.1 | 128.6 | 0.67367 | 165.4263 | 127.9263 | 0.21878 | 0.40683 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 24/02; UpstreamInshore | phae>chl; depth 57m
AR44 | 20200613 | L13 | 24 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 219.4 | 138.9 | 0.67367 | 218.7263 | 138.2263 | 2.8179 | 1.238 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 24/03; UpstreamInshore
AR44 | 20200613 | L13 | 24 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 211.3 | 136.1 | 0.67367 | 210.6263 | 135.4263 | 2.6324 | 1.3414 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 24/04; UpstreamInshore
AR44 | 20200613 | L13 | 24 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 161 | 101.6 | 0.67367 | 160.3263 | 100.9263 | 0.34622 | 0.14688 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 24/05; UpstreamInshore
AR44 | 20200613 | L13 | 24 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 88.12 | 54.77 | 0.67367 | 87.4463 | 54.0963 | 0.38914 | 0.13997 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 24/06; UpstreamInshore
AR44 | 20200613 | L6 | 25 | 1 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 85.98 | 70.97 | 0.67367 | 85.3063 | 70.2963 | 0.087488 | 0.25597 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 25/01; Inshore | phae>chl; depth 90m
AR44 | 20200613 | L6 | 25 | 3 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 142.6 | 106.2 | 0.67367 | 141.9263 | 105.5263 | 0.21236 | 0.30371 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 25/02; Inshore | phae>chl; depth 55m
AR44 | 20200613 | L6 | 25 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 248 | 154.3 | 0.67367 | 247.3263 | 153.6263 | 3.28 | 1.2278 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 25/03; Inshore
AR44 | 20200613 | L6 | 25 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 238.5 | 151.2 | 0.67367 | 237.8263 | 150.5263 | 3.0559 | 1.3609 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 25/04; Inshore
AR44 | 20200613 | L6 | 25 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 186.4 | 116.5 | 0.67367 | 185.7263 | 115.8263 | 0.40742 | 0.15848 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 25/05; Inshore
AR44 | 20200614 | L8 | 26 | 5 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 166.7 | 105.4 | 0.67367 | 166.0263 | 104.7263 | 2.1438 | 0.92626 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 26/01; Central
AR44 | 20200614 | L8 | 26 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 210 | 132.8 | 0.67367 | 209.3263 | 132.1263 | 2.7024 | 1.1745 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 26/02; Central
AR44 | 20200614 | L8 | 26 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76.23 | 46.48 | 0.67367 | 75.5563 | 45.8063 | 0.34713 | 0.10089 | 20220519 | 13-19 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 26/03; Central
AR44 | 20200614 | L8 | 26 | 8 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 151.8 | 93.23 | 0.67367 | 151.1263 | 92.5563 | 0.34171 | 0.11093 | 20220519 | 13-20 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 26/04; Central
AR44 | 20200615 |  | 32 | 3 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 30.26 | 24.56 | 0.67367 | 29.5863 | 23.8863 | 0.033255 | 0.08356 | 20220519 | 13-20 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 32/01; Gliders 339/514/583 | phae>chl; depth 86m
AR44 | 20200615 |  | 32 | 4 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.64 | 45.62 | 0.67367 | 62.9663 | 44.9463 | 0.10513 | 0.11468 | 20220519 | 13-20 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 32/02; Gliders 339/514/583 | phae>chl; depth 62m
AR44 | 20200615 |  | 32 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 176.4 | 108 | 0.67367 | 175.7263 | 107.3263 | 2.3943 | 0.75489 | 20220519 | 13-20 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 32/03; Gliders 339/514/583
AR44 | 20200615 |  | 32 | 5 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 164.4 | 101.8 | 0.67367 | 163.7263 | 101.1263 | 2.1913 | 0.776 | 20220519 | 13-20 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 32/04; Gliders 339/514/583
AR44 | 20200615 |  | 32 | 6 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 157.4 | 98.76 | 0.67367 | 156.7263 | 98.0863 | 0.34179 | 0.13744 | 20220519 | 13-20 | 20230322 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 32/05; Gliders 339/514/583
EN655 | 20200725 | L1 | 1 | 5 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 204.5 | 125.1 | 0.6235 | 203.8765 | 124.4765 | 3.3736 | 1.0597 | 20220519 | 11-70 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | NS? Filter might not have been on all the wayl;hear sucking
EN655 | 20200725 | L1 | 1 | 5 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 190.4 | 118.9 | 0.6235 | 189.7765 | 118.2765 | 3.0379 | 1.1746 | 20220519 | 11-70 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L1 | 1 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 135.2 | 86.28 | 0.76967 | 134.4303 | 85.5103 | 1.0393 | 0.48348 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200725 | L1 | 1 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 136.7 | 89.77 | 0.76967 | 135.9303 | 89.0003 | 0.997 | 0.58791 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200725 | L1 | 1 | 5 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 275.7 | 163.6 | 0.51667 | 275.1833 | 163.0833 | 2.281 | 0.48764 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR5
EN655 | 20200725 | L1 | 1 | 5 |  |  | b | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 263.8 | 158.7 | 0.51667 | 263.2833 | 158.1833 | 2.1313 | 0.54503 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR6
EN655 | 20200725 | L1 | 1 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_111320b_12mm | 1.8343 | 0.5442 | 144.1 | 84.77 | 0.51667 | 143.5833 | 84.2533 | 1.9325 | 0.35714 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR11
EN655 | 20200725 | L1 | 1 | 5 |  |  | b | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_111320b_12mm | 1.8343 | 0.5442 | 147.8 | 87.09 | 0.51667 | 147.2833 | 86.5733 | 1.9774 | 0.37524 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR12
EN655 | 20200725 | L1 | 1 | 7 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 190.2 | 117.2 | 0.6235 | 189.5765 | 116.5765 | 3.1017 | 1.0503 | 20220519 | 11-70 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L1 | 1 | 7 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 194.6 | 119.5 | 0.58 | 194.02 | 118.92 | 3.1909 | 1.0445 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L1 | 1 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 135 | 86.51 | 0.76967 | 134.2303 | 85.7403 | 1.0301 | 0.49671 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200725 | L1 | 1 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 137.9 | 89.4 | 0.76967 | 137.1303 | 88.6303 | 1.0303 | 0.54797 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200725 | L1 | 1 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 236.3 | 140.1 | 0.51667 | 235.7833 | 139.5833 | 1.9641 | 0.41362 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR3
EN655 | 20200725 | L1 | 1 | 7 |  |  | b | 292 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 231.6 | 137.6 | 0.51667 | 231.0833 | 137.0833 | 1.9258 | 0.41737 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR4
EN655 | 20200725 | L1 | 1 | 7 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 277.9 | 162.8 | 0.51667 | 277.3833 | 162.2833 | 1.2519 | 0.22077 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR9
EN655 | 20200725 | L1 | 1 | 7 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 298.1 | 170.7 | 0.51667 | 297.5833 | 170.1833 | 1.3857 | 0.15868 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR10
EN655 | 20200725 | L1 | 1 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 99.02 | 60.8 | 0.6235 | 98.3965 | 60.1765 | 0.81196 | 0.25966 | 20220519 | 11-70 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L1 | 1 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 105.3 | 65.13 | 0.6235 | 104.6765 | 64.5065 | 0.85338 | 0.29534 | 20220519 | 11-70 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L1 | 1 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 51.49 | 31.45 | 0.6235 | 50.8665 | 30.8265 | 0.42574 | 0.12322 | 20220519 | 11-70 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L1 | 1 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.55 | 30.98 | 0.6235 | 48.9265 | 30.3565 | 0.39451 | 0.14608 | 20220519 | 11-70 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L1 | 1 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 35.41 | 21.71 | 0.76967 | 34.6403 | 20.9403 | 0.29105 | 0.081855 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200725 | L1 | 1 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 41.06 | 24.97 | 0.76967 | 40.2903 | 24.2003 | 0.34182 | 0.089135 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200725 | L1 | 1 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 14.73 | 8.981 | 0.51667 | 14.2133 | 8.4643 | 0.11738 | 0.026808 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR1
EN655 | 20200725 | L1 | 1 | 16 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 13.51 | 8.1 | 0.51667 | 12.9933 | 7.5833 | 0.11046 | 0.018722 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR2
EN655 | 20200725 | L1 | 1 | 16 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 7.599 | 4.318 | 0.51667 | 7.0823 | 3.8013 | 0.035622 | -0.0011882 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR7; neg phae
EN655 | 20200725 | L1 | 1 | 16 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 6.229 | 3.738 | 0.51667 | 5.7123 | 3.2213 | 0.027045 | 0.002135 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR8
EN655 | 20200725 | L2 | 2 | 1 |  |  | a | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 349.5 | 207.2 | 0.6235 | 348.8765 | 206.5765 | 9.0692 | 1.9669 | 20220519 | 11-70 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | reps slightly off? But percent range is only 8%...is it b/c values are so high? Therefore within bio variation?
EN655 | 20200725 | L2 | 2 | 1 |  |  | b | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 321.6 | 190.4 | 0.6235 | 320.9765 | 189.7765 | 8.3618 | 1.7768 | 20220519 | 11-70 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | reps slightly off? But percent range is only 8%...is it b/c values are so high? Therefore within bio variation?
EN655 | 20200725 | L2 | 2 | 4 |  |  | a | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 352.4 | 213.6 | 0.6235 | 351.7765 | 212.9765 | 8.8461 | 2.5318 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | reps slightly off? But percent range is only 12%...is it b/c values are so high? Therefore within bio variation?
EN655 | 20200725 | L2 | 2 | 4 |  |  | b | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 301.7 | 178.8 | 0.6235 | 301.0765 | 178.1765 | 7.8328 | 1.686 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | reps slightly off? But percent range is only 12%...is it b/c values are so high? Therefore within bio variation?
EN655 | 20200725 | L2 | 2 | 4 |  |  | a | 170 | 5 | 5 | 1 | 0 | 0.33333 | Cal_111320b_12mm | 1.8343 | 0.5442 | 237.6 | 138.4 | 0.51667 | 237.0833 | 137.8833 | 10.4725 | 1.6722 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR17; poor replication (27% percent range); this value is higher than both WSW values so toss
EN655 | 20200725 | L2 | 2 | 4 |  |  | b | 295 | 5 | 5 | 1 | 0 | 0.33333 | Cal_111320b_12mm | 1.8343 | 0.5442 | 322.9 | 192.1 | 0.51667 | 322.3833 | 191.5833 | 7.9574 | 1.7669 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR18; reps off?
EN655 | 20200725 | L2 | 2 | 4 |  |  | a | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_111320b_12mm | 1.8343 | 0.5442 | 148.1 | 86.01 | 0.51667 | 147.5833 | 85.4933 | 2.0223 | 0.30094 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR23; reps are clearly off (107% percent range); given size fractionation of this cast, this sample is clearly low so toss
EN655 | 20200725 | L2 | 2 | 4 |  |  | b | 551 | 20 | 5 | 1 | 0 | 0.2 | Cal_111320b_12mm | 1.8343 | 0.5442 | 307.6 | 184 | 0.51667 | 307.0833 | 183.4833 | 6.7097 | 1.6006 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR24; reps off; lab notebook also says: what's on filter? Very concentration
EN655 | 20200725 | L2 | 2 | 7 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 574.3 | 342.6 | 0.76967 | 573.5303 | 341.8303 | 9.8446 | 2.3299 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | have undiluted and diluted fluor values but undiluted are more in line with replicate so use that one
EN655 | 20200725 | L2 | 2 | 7 |  |  | b | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 357.7 | 208.3 | 0.6235 | 357.0765 | 207.6765 | 9.5217 | 1.5731 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200725 | L2 | 2 | 7 |  |  | a | 292 | 5 | 5 | 1 | 0 | 0.33333 | Cal_111320b_12mm | 1.8343 | 0.5442 | 329.5 | 199.6 | 0.51667 | 328.9833 | 199.0833 | 7.9839 | 2.225 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR15; lab notebook says "what's on filter? Very concentrated, maybe also zoop or folds of filter? ; reps off but 23% range, DF thinks just bio variation since chl is so large
EN655 | 20200725 | L2 | 2 | 7 |  |  | b | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_111320b_12mm | 1.8343 | 0.5442 | 387.5 | 238.3 | 0.51667 | 386.9833 | 237.7833 | 9.1388 | 3.0129 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR16; reps off but 23% range, DF thinks just bio variation since chl is so large
EN655 | 20200725 | L2 | 2 | 7 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.2 | Cal_111320b_12mm | 1.8343 | 0.5442 | 387 | 231.9 | 0.51667 | 386.4833 | 231.3833 | 8.435 | 2.0639 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR21; lab notebook says: what is on filter? Very concentrated, maybe also zooplankton or folds of filter?; somewhat poor replication (23% range) -- also large chl concentrations, so DF thinks this is just within bio variation
EN655 | 20200725 | L2 | 2 | 7 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.2 | Cal_111320b_12mm | 1.8343 | 0.5442 | 314.1 | 191.9 | 0.51667 | 313.5833 | 191.3833 | 6.6457 | 2.0381 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR22; lab notebook says: what is on filter? Very concentrated; somewhat poor replication (23% range) -- also large chl concentrations, so DF thinks this is just within bio variation
EN655 | 20200725 | L2 | 2 | 13 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 392.6 | 232.4 | 0.6235 | 391.9765 | 231.7765 | 6.8067 | 1.4482 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L2 | 2 | 13 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 418.9 | 259.2 | 0.6235 | 418.2765 | 258.5765 | 6.7854 | 2.424 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L2 | 2 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.82 | 32.32 | 0.6235 | 52.1965 | 31.6965 | 0.43551 | 0.12894 | 20220519 | 11-70 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L2 | 2 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.05 | 32.16 | 0.6235 | 51.4265 | 31.5365 | 0.42255 | 0.13905 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L2 | 2 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 41.63 | 25.71 | 0.76967 | 40.8603 | 24.9403 | 0.33821 | 0.10592 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200725 | L2 | 2 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 42.18 | 26.3 | 0.76967 | 41.4103 | 25.5303 | 0.33736 | 0.11728 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200725 | L2 | 2 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 10.35 | 6.223 | 0.51667 | 9.8333 | 5.7063 | 0.084262 | 0.012944 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR13
EN655 | 20200725 | L2 | 2 | 15 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 10.86 | 6.237 | 0.51667 | 10.3433 | 5.7203 | 0.094389 | 0.0030551 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR14
EN655 | 20200725 | L2 | 2 | 15 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 5.656 | 3.212 | 0.51667 | 5.1393 | 2.6953 | 0.026583 | -0.0021232 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR19; neg phaeo but very minimally neg
EN655 | 20200725 | L2 | 2 | 15 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 5.283 | 3.086 | 0.51667 | 4.7663 | 2.5693 | 0.02381 | -0.00057799 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR20; neg phaeo but very minimally neg
EN655 | 20200725 | L3 | 3 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 60.15 | 42.25 | 0.6235 | 59.5265 | 41.6265 | 0.38027 | 0.36101 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L3 | 3 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 58.76 | 41.16 | 0.58 | 58.18 | 40.58 | 0.3739 | 0.34874 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | Cast# missing from tissue capsule but process of elimination it is clearly C3N1b
EN655 | 20200725 | L3 | 3 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 56.03 | 38.86 | 0.6235 | 55.4065 | 38.2365 | 0.36477 | 0.31615 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L3 | 3 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.4 | 34.55 | 0.6235 | 48.7765 | 33.9265 | 0.31548 | 0.28868 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L3 | 3 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 192.8 | 126.1 | 0.58 | 192.22 | 125.52 | 1.417 | 0.81825 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L3 | 3 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 185.6 | 119.4 | 0.6235 | 184.9765 | 118.7765 | 1.4064 | 0.70878 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L3 | 3 | 5 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 61.24 | 38.86 | 0.73183 | 60.5082 | 38.1282 | 0.45694 | 0.19256 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR28
EN655 | 20200725 | L3 | 3 | 5 |  |  | b | 292 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 46.31 | 29.98 | 0.51667 | 45.7933 | 29.4633 | 0.33456 | 0.16906 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR29
EN655 | 20200725 | L3 | 3 | 5 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 43.23 | 27.03 | 0.51667 | 42.7133 | 26.5133 | 0.1762 | 0.0644 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR34
EN655 | 20200725 | L3 | 3 | 5 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 34.57 | 21.73 | 0.51667 | 34.0533 | 21.2133 | 0.13966 | 0.052849 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR35
EN655 | 20200725 | L3 | 3 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 98.02 | 63.76 | 0.6235 | 97.3965 | 63.1365 | 0.72783 | 0.3965 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L3 | 3 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 99.01 | 63 | 0.6235 | 98.3865 | 62.3765 | 0.76501 | 0.34578 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L3 | 3 | 10 |  |  | a | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 78.47 | 46.01 | 0.51667 | 77.9533 | 45.4933 | 0.65825 | 0.11146 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR30; very poor replication (106% range); given this size structure of this cast (ratio of 5/WSW), DF thinks there's something wrong with this sample s toss
EN655 | 20200725 | L3 | 3 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 26.65 | 16.8 | 0.51667 | 26.1333 | 16.2833 | 0.20043 | 0.076012 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR31;very poor replication (106% range)
EN655 | 20200725 | L3 | 3 | 10 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 32.69 | 19.88 | 0.51667 | 32.1733 | 19.3633 | 0.13908 | 0.036321 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR36; poor replication (58% range) but could just be biological variation?
EN655 | 20200725 | L3 | 3 | 10 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 18.3 | 11.31 | 0.51667 | 17.7833 | 10.7933 | 0.075891 | 0.021879 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR37; poor replication (58% range) but could just be biological variation?
EN655 | 20200725 | L3 | 3 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 81.47 | 51.76 | 0.4865 | 80.9835 | 51.2735 | 0.63117 | 0.2819 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L3 | 3 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 90.02 | 57.17 | 0.58 | 89.44 | 56.59 | 0.69788 | 0.30987 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200725 | L3 | 3 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 56.09 | 34.55 | 0.76967 | 55.3203 | 33.7803 | 0.4576 | 0.14395 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200725 | L3 | 3 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.29 | 31.03 | 0.6235 | 48.6665 | 30.4065 | 0.38792 | 0.15355 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200725 | L3 | 3 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.49 | 25.33 | 0.76967 | 39.7203 | 24.5603 | 0.32206 | 0.1153 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200725 | L3 | 3 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 48.37 | 29.41 | 0.76967 | 47.6003 | 28.6403 | 0.40279 | 0.10723 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | chl higher than WSW value and poor replication though phaeo is lower and similar to <10a rep
EN655 | 20200725 | L3 | 3 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 10.06 | 6.157 | 0.51667 | 9.5433 | 5.6403 | 0.079689 | 0.016393 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR26
EN655 | 20200725 | L3 | 3 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 9.271 | 5.549 | 0.51667 | 8.7543 | 5.0323 | 0.075993 | 0.0097312 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR27
EN655 | 20200725 | L3 | 3 | 19 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.185 | 2.206 | 0.51667 | 2.6683 | 1.6893 | 0.01061 | 0.0046651 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR32; poor replication (60% range) but values are so low, within noise of flourometer
EN655 | 20200725 | L3 | 3 | 19 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.391 | 2.558 | 0.51667 | 3.8743 | 2.0413 | 0.019937 | -0.0014124 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR33; poor replication (60% range) but values are so low, within noise of flourometer
EN655 | 20200726 | L4 | 7 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 37.09 | 27.91 | 0.6235 | 36.4665 | 27.2865 | 0.19502 | 0.29089 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L4 | 7 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.86 | 29.33 | 0.58 | 36.28 | 28.75 | 0.15997 | 0.35201 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L4 | 7 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 46.41 | 32.56 | 0.6235 | 45.7865 | 31.9365 | 0.29423 | 0.27449 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L4 | 7 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 44.28 | 30.71 | 0.4865 | 43.7935 | 30.2235 | 0.28829 | 0.24993 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L4 | 7 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 169.4 | 115.8 | 0.58 | 168.82 | 115.22 | 1.1387 | 0.91313 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L4 | 7 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 169.1 | 112.2 | 0.6235 | 168.4765 | 111.5765 | 1.2088 | 0.77814 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L4 | 7 | 5 |  |  | a | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 60.59 | 39.15 | 0.51667 | 60.0733 | 38.6333 | 0.43478 | 0.21887 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR41
EN655 | 20200726 | L4 | 7 | 5 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 57.75 | 36.47 | 0.51667 | 57.2333 | 35.9533 | 0.433 | 0.17737 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR42
EN655 | 20200726 | L4 | 7 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 29.91 | 18.2 | 0.51667 | 29.3933 | 17.6833 | 0.12714 | 0.033046 | 20201113 | 11-27 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR51
EN655 | 20200726 | L4 | 7 | 5 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 36.23 | 23.52 | 0.73183 | 35.4982 | 22.7882 | 0.13799 | 0.06843 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR52
EN655 | 20200726 | L4 | 7 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 200.4 | 137.7 | 0.58 | 199.82 | 137.12 | 1.332 | 1.1098 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L4 | 7 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 194.7 | 135.6 | 0.6235 | 194.0765 | 134.9765 | 1.2555 | 1.1481 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L4 | 7 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 145.8 | 92.05 | 0.6235 | 145.1765 | 91.4265 | 1.1419 | 0.48623 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L4 | 7 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 146.3 | 93.25 | 0.58 | 145.72 | 92.67 | 1.127 | 0.52324 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L4 | 7 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 121.9 | 76.18 | 0.76967 | 121.1303 | 75.4103 | 0.97129 | 0.37161 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200726 | L4 | 7 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 131 | 83.02 | 0.76967 | 130.2303 | 82.2503 | 1.0193 | 0.4454 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200726 | L4 | 7 | 9 |  |  | a | 282 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 20.99 | 13.95 | 0.51667 | 20.4733 | 13.4333 | 0.14934 | 0.088415 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR39;
EN655 | 20200726 | L4 | 7 | 9 |  |  | b | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 22.39 | 14.81 | 0.51667 | 21.8733 | 14.2933 | 0.15371 | 0.088118 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR40
EN655 | 20200726 | L4 | 7 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 10.07 | 6.702 | 0.51667 | 9.5533 | 6.1853 | 0.036633 | 0.019498 | 20201113 | 11-27 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR49
EN655 | 20200726 | L4 | 7 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.66 | 7.992 | 0.51667 | 11.1433 | 7.4753 | 0.039896 | 0.027941 | 20201113 | 11-27 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR50
EN655 | 20200726 | L4 | 7 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 153.5 | 101.6 | 0.6235 | 152.8765 | 100.9765 | 1.1026 | 0.6956 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L4 | 7 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 153.7 | 99.61 | 0.6235 | 153.0765 | 98.9865 | 1.1491 | 0.61363 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L4 | 7 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 119.3 | 77.14 | 0.76967 | 118.5303 | 76.3703 | 0.89566 | 0.46433 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200726 | L4 | 7 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 120.7 | 78.51 | 0.76967 | 119.9303 | 77.7403 | 0.8963 | 0.48809 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200726 | L4 | 7 | 13 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 19.3 | 12.81 | 0.51667 | 18.7833 | 12.2933 | 0.13206 | 0.076644 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR43
EN655 | 20200726 | L4 | 7 | 13 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 19.59 | 12.8 | 0.51667 | 19.0733 | 12.2833 | 0.13816 | 0.070369 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR44
EN655 | 20200726 | L4 | 7 | 13 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 9.518 | 6.629 | 0.51667 | 9.0013 | 6.1123 | 0.031423 | 0.024045 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR45
EN655 | 20200726 | L4 | 7 | 13 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 12.25 | 8.085 | 0.51667 | 11.7333 | 7.5683 | 0.045138 | 0.023295 | 20201113 | 11-27 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR46
EN655 | 20200726 | L4 | 7 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 59.17 | 37.39 | 0.58 | 58.59 | 36.81 | 0.4627 | 0.19281 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L4 | 7 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 60.83 | 38.77 | 0.58 | 60.25 | 38.19 | 0.46865 | 0.21143 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L4 | 7 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 41.03 | 26.19 | 0.76967 | 40.2603 | 25.4203 | 0.31527 | 0.13742 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200726 | L4 | 7 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 41.47 | 26.23 | 0.76967 | 40.7003 | 25.4603 | 0.32376 | 0.12963 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200726 | L4 | 7 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 17.44 | 10.55 | 0.51667 | 16.9233 | 10.0333 | 0.14067 | 0.03024 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR25
EN655 | 20200726 | L4 | 7 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 16.76 | 9.928 | 0.51667 | 16.2433 | 9.4113 | 0.13949 | 0.020828 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR38
EN655 | 20200726 | L4 | 7 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 13.62 | 8.134 | 0.51667 | 13.1033 | 7.6173 | 0.05967 | 0.0094558 | 20201113 | 11-27 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR47
EN655 | 20200726 | L4 | 7 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 14.19 | 8.444 | 0.51667 | 13.6733 | 7.9273 | 0.062272 | 0.0094068 | 20201113 | 11-27 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR48
EN655 | 20200726 | L5 | 10 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.77 | 14.67 | 0.58 | 38.19 | 14.09 | 0.51199 | -0.26108 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | This is not Rb value, it is the Ra value from the previous sample run on the fluorometer. Assume other rep is totally fine.
EN655 | 20200726 | L5 | 10 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 18.87 | 15.66 | 0.58 | 18.29 | 15.08 | 0.068194 | 0.20035 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | This chl value consistent with CTD fluor trace
EN655 | 20200726 | L5 | 10 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.6 | 15.18 | 0.6235 | 19.9765 | 14.5565 | 0.11514 | 0.14408 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L5 | 10 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.9 | 14.42 | 0.6235 | 19.2765 | 13.7965 | 0.11642 | 0.12927 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L5 | 10 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 69.97 | 47.26 | 0.58 | 69.39 | 46.68 | 0.48246 | 0.34881 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L5 | 10 | 5 |  |  | b | 280 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.26 | 47.24 | 0.58 | 69.68 | 46.66 | 0.50651 | 0.35408 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | WSW-b spilled some while pouring
EN655 | 20200726 | L5 | 10 | 7 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 123.2 | 81.48 | 0.58 | 122.62 | 80.9 | 1.7726 | 1.1087 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L5 | 10 | 7 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 124.5 | 81.86 | 0.58 | 123.92 | 81.28 | 1.8117 | 1.0831 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L5 | 10 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 29.14 | 18.48 | 0.51667 | 28.6233 | 17.9633 | 0.21765 | 0.088352 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR55
EN655 | 20200726 | L5 | 10 | 7 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 27.14 | 17.5 | 0.51667 | 26.6233 | 16.9833 | 0.19682 | 0.092483 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR56
EN655 | 20200726 | L5 | 10 | 7 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 16.91 | 10.95 | 0.51667 | 16.3933 | 10.4333 | 0.064826 | 0.029855 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR61
EN655 | 20200726 | L5 | 10 | 7 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 13.03 | 8.587 | 0.51667 | 12.5133 | 8.0703 | 0.048326 | 0.024911 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR62
EN655 | 20200726 | L5 | 10 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 95.2 | 60.35 | 0.58 | 94.62 | 59.77 | 0.7189 | 0.30979 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | WSWa collected later
EN655 | 20200726 | L5 | 10 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 95.93 | 61.41 | 0.6235 | 95.3065 | 60.7865 | 0.71209 | 0.3341 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L5 | 10 | 10 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.46 | 7.393 | 0.51667 | 10.9433 | 6.8763 | 0.082755 | 0.033983 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR57
EN655 | 20200726 | L5 | 10 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.84 | 7.764 | 0.51667 | 11.3233 | 7.2473 | 0.082938 | 0.040098 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR58
EN655 | 20200726 | L5 | 10 | 10 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 5.666 | 4.214 | 0.51667 | 5.1493 | 3.6973 | 0.015764 | 0.017727 | 20201113 | 11-27 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR63
EN655 | 20200726 | L5 | 10 | 10 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.495 | 3.181 | 0.51667 | 3.9783 | 2.6643 | 0.014266 | 0.0098683 | 20201113 | 11-27 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR64
EN655 | 20200726 | L5 | 10 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 84.68 | 52.32 | 0.58 | 84.1 | 51.74 | 0.68747 | 0.23391 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L5 | 10 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 83.43 | 52.37 | 0.58 | 82.85 | 51.79 | 0.65985 | 0.26242 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L5 | 10 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.55 | 20.15 | 0.6235 | 44.9265 | 19.5265 | 0.53961 | -0.19188 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | WSWa droped filter side down; reps slightly off? Very neg phaeo – assume sample is compromised and toss
EN655 | 20200726 | L5 | 10 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.89 | 28.43 | 0.58 | 45.31 | 27.85 | 0.37093 | 0.12502 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L5 | 10 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 32.86 | 20.66 | 0.76967 | 32.0903 | 19.8903 | 0.25918 | 0.095024 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200726 | L5 | 10 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.74 | 21.19 | 0.76967 | 32.9703 | 20.4203 | 0.26662 | 0.097026 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200726 | L5 | 10 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 10.43 | 6.568 | 0.51667 | 9.9133 | 6.0513 | 0.078851 | 0.024231 | 20201113 | 11-27 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR53
EN655 | 20200726 | L5 | 10 | 16 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 11.31 | 6.746 | 0.51667 | 10.7933 | 6.2293 | 0.093184 | 0.01293 | 20201113 | 11-27 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR54
EN655 | 20200726 | L5 | 10 | 16 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 7.358 | 4.382 | 0.51667 | 6.8413 | 3.8653 | 0.032252 | 0.002698 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR59
EN655 | 20200726 | L5 | 10 | 16 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 6.174 | 3.661 | 0.51667 | 5.6573 | 3.1443 | 0.027333 | 0.0012009 | 20201113 | 11-26 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR60
EN655 | 20200726 | L7 | 11 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.004 | 5.817 | 0.58 | 7.424 | 5.237 | 0.046461 | 0.046798 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication(52% range) but also low chl values so this could be why or could there have been a switch with one of these samples with N3 WSW?
EN655 | 20200726 | L7 | 11 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.065 | 5.794 | 0.6235 | 6.4415 | 5.1705 | 0.027002 | 0.065074 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication(52% range) but also low chl values so this could be why or could there have been a switch with one of these samples with N3 WSW?
EN655 | 20200726 | L7 | 11 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.699 | 5.976 | 0.58 | 7.119 | 5.396 | 0.036604 | 0.059487 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication(48% range) but also low chl values so this could be why or could there have been a switch with one of these samples with N1 WSW?
EN655 | 20200726 | L7 | 11 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.838 | 5.792 | 0.58 | 6.258 | 5.212 | 0.022222 | 0.070593 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication(48% range) but also low chl values so this could be why or could there have been a switch with one of these samples with N1 WSW?
EN655 | 20200726 | L7 | 11 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.79 | 8.693 | 0.58 | 10.21 | 8.113 | 0.044549 | 0.099926 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L7 | 11 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.14 | 7.987 | 0.58 | 9.56 | 7.407 | 0.045739 | 0.086164 | 20220519 | 11-68 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L7 | 11 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 110.1 | 72.71 | 0.58 | 109.52 | 72.13 | 0.79433 | 0.49016 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L7 | 11 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 112.9 | 74.75 | 0.58 | 112.32 | 74.17 | 0.81047 | 0.51034 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L7 | 11 | 7 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 65.52 | 42.6 | 0.73183 | 64.7882 | 41.8682 | 0.46637 | 0.24441 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR69
EN655 | 20200726 | L7 | 11 | 7 |  |  | b | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 74.44 | 49.65 | 0.73183 | 73.7082 | 48.9182 | 0.50271 | 0.32494 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR70
EN655 | 20200726 | L7 | 11 | 7 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 59.13 | 37.59 | 0.73183 | 58.3982 | 36.8582 | 0.23386 | 0.10001 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR75
EN655 | 20200726 | L7 | 11 | 7 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 76.69 | 48.87 | 0.73183 | 75.9582 | 48.1382 | 0.30204 | 0.13401 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR76
EN655 | 20200726 | L7 | 11 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 132.6 | 64.19 | 0.58 | 132.02 | 63.61 | 1.4533 | -0.32057 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | reps off and neg phaeo? Why is Ra so low? Clearly somethin wrong with Ra reading
EN655 | 20200726 | L7 | 11 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 133 | 86.03 | 0.58 | 132.42 | 85.45 | 0.99785 | 0.52384 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L7 | 11 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 222.9 | 140.8 | 0.58 | 222.32 | 140.22 | 1.7442 | 0.75286 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L7 | 11 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 220.8 | 138.4 | 0.58 | 220.22 | 137.82 | 1.7505 | 0.70375 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L7 | 11 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 186.1 | 117.1 | 0.76967 | 185.3303 | 116.3303 | 1.4659 | 0.60574 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200726 | L7 | 11 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 188.4 | 119.6 | 0.76967 | 187.6303 | 118.8303 | 1.4616 | 0.6545 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200726 | L7 | 11 | 11 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 39.83 | 26.3 | 0.73183 | 39.0982 | 25.5682 | 0.27625 | 0.1593 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR67
EN655 | 20200726 | L7 | 11 | 11 |  |  | b | 292 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 40.25 | 24.91 | 0.73183 | 39.5182 | 24.1782 | 0.31427 | 0.099005 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR68
EN655 | 20200726 | L7 | 11 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 14.15 | 10.43 | 0.73183 | 13.4182 | 9.6982 | 0.040462 | 0.047548 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR73
EN655 | 20200726 | L7 | 11 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 13.84 | 10.42 | 0.73183 | 13.1082 | 9.6882 | 0.037199 | 0.05072 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR74
EN655 | 20200726 | L7 | 11 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.53 | 7.602 | 0.58 | 10.95 | 7.022 | 0.083448 | 0.041599 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L7 | 11 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.62 | 7.795 | 0.58 | 11.04 | 7.215 | 0.08126 | 0.047224 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L7 | 11 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.53 | 7.707 | 0.58 | 10.95 | 7.127 | 0.081217 | 0.0457 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L7 | 11 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.67 | 7.967 | 0.58 | 12.09 | 7.387 | 0.099912 | 0.031635 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200726 | L7 | 11 | 17 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.37 | 6.711 | 0.76967 | 9.6003 | 5.9413 | 0.077733 | 0.028069 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200726 | L7 | 11 | 17 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 9.632 | 6.525 | 0.76967 | 8.8623 | 5.7553 | 0.066006 | 0.036484 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200726 | L7 | 11 | 17 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 2.525 | 2.115 | 0.73183 | 1.7932 | 1.3832 | 0.0083711 | 0.015191 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR65; poor replication (45% range) but also low chl value and could be within bio variation
EN655 | 20200726 | L7 | 11 | 17 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 2.609 | 1.959 | 0.73183 | 1.8772 | 1.2272 | 0.013271 | 0.0076332 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR66; poor replication (45% range) but also low chl value and could be within bio variation
EN655 | 20200726 | L7 | 11 | 17 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 1.763 | 1.434 | 0.73183 | 1.0312 | 0.70217 | 0.0035655 | 0.0027835 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR71
EN655 | 20200726 | L7 | 11 | 17 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 1.825 | 1.396 | 0.73183 | 1.0932 | 0.66417 | 0.0046662 | 0.001361 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR72
EN655 | 20200727 | L10 | 12 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.03 | 7.625 | 0.58 | 9.45 | 7.045 | 0.051093 | 0.074364 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L10 | 12 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.13 | 8.095 | 0.4865 | 10.6435 | 7.6085 | 0.064477 | 0.071015 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L10 | 12 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.09 | 45.01 | 0.4865 | 64.6035 | 44.5235 | 0.42659 | 0.36628 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L10 | 12 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.97 | 44.85 | 0.58 | 63.39 | 44.27 | 0.40619 | 0.38216 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L10 | 12 | 5 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 45.05 | 30.94 | 0.73183 | 44.3182 | 30.2082 | 0.28711 | 0.22573 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR83
EN655 | 20200727 | L10 | 12 | 5 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 34.69 | 23.28 | 0.73183 | 33.9582 | 22.5482 | 0.23217 | 0.15063 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR84
EN655 | 20200727 | L10 | 12 | 5 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 46.15 | 30.28 | 0.73183 | 45.4182 | 29.5482 | 0.17262 | 0.09553 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR85
EN655 | 20200727 | L10 | 12 | 5 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 52.13 | 34.11 | 0.73183 | 51.3982 | 33.3782 | 0.19529 | 0.10651 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR86
EN655 | 20200727 | L10 | 12 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 149.3 | 100.5 | 0.58 | 148.72 | 99.92 | 1.0367 | 0.74264 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L10 | 12 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 150.2 | 101 | 0.58 | 149.62 | 100.42 | 1.0452 | 0.74304 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L10 | 12 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 71.42 | 48.26 | 0.76967 | 70.6503 | 47.4903 | 0.49202 | 0.35368 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200727 | L10 | 12 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76.43 | 51.63 | 0.76967 | 75.6603 | 50.8603 | 0.52686 | 0.37885 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200727 | L10 | 12 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 51.03 | 32.81 | 0.73183 | 50.2982 | 32.0782 | 0.372 | 0.17444 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR79; 5 values are smaller than 20s – possible that they're switched? Or is it just that >5 and >20 chl is basically the same?
EN655 | 20200727 | L10 | 12 | 8 |  |  | b | 292 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 53.66 | 33.32 | 0.73183 | 52.9282 | 32.5882 | 0.41671 | 0.14032 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR80; 5 values are smaller than 20s – possible that they're switched? Or is it just that >5 and >20 chl is basically the same?
EN655 | 20200727 | L10 | 12 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 111.9 | 71.2 | 0.73183 | 111.1682 | 70.4682 | 0.44269 | 0.1968 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR89; 5 values are smaller than 20s – possible that they're switched? Or is it just that >5 and >20 chl is basically the same?
EN655 | 20200727 | L10 | 12 | 8 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 105.6 | 67.38 | 0.73183 | 104.8682 | 66.6482 | 0.41571 | 0.18911 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR90; 5 values are smaller than 20s – possible that they're switched? Or is it just that >5 and >20 chl is basically the same?
EN655 | 20200727 | L10 | 12 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 32.66 | 20.97 | 0.58 | 32.08 | 20.39 | 0.24835 | 0.11476 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | Niskin 13 leaky stopcock
EN655 | 20200727 | L10 | 12 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 32.41 | 21.2 | 0.4865 | 31.9235 | 20.7135 | 0.23815 | 0.13071 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L10 | 12 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 26.02 | 16.79 | 0.76967 | 25.2503 | 16.0203 | 0.19609 | 0.089203 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200727 | L10 | 12 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 26.59 | 17.01 | 0.76967 | 25.8203 | 16.2403 | 0.20352 | 0.085685 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200727 | L10 | 12 | 13 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.894 | 2.725 | 0.73183 | 3.1622 | 1.9932 | 0.023787 | 0.010051 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR81
EN655 | 20200727 | L10 | 12 | 13 |  |  | b | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.174 | 2.298 | 0.73183 | 2.4422 | 1.5662 | 0.017764 | 0.008734 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR82
EN655 | 20200727 | L10 | 12 | 13 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.004 | 3.096 | 0.73183 | 3.2722 | 2.3642 | 0.0098582 | 0.011557 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR91
EN655 | 20200727 | L10 | 12 | 13 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.812 | 2.619 | 0.73183 | 3.0802 | 1.8872 | 0.012953 | 0.0041422 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR92
EN655 | 20200727 | L10 | 12 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.37 | 13.9 | 0.58 | 20.79 | 13.32 | 0.1587 | 0.078506 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L10 | 12 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.31 | 13.1 | 0.58 | 19.73 | 12.52 | 0.15317 | 0.069783 | 20220519 | 11-69 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L10 | 12 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.45 | 10.68 | 0.58 | 15.87 | 10.1 | 0.12258 | 0.05728 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L10 | 12 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.2 | 10.23 | 0.58 | 15.62 | 9.65 | 0.12683 | 0.045017 | 20220519 | 11-69 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L10 | 12 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.03 | 7.801 | 0.76967 | 11.2603 | 7.0313 | 0.089842 | 0.035371 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200727 | L10 | 12 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.55 | 8.527 | 0.76967 | 12.7803 | 7.7573 | 0.10671 | 0.031431 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200727 | L10 | 12 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 2.343 | 1.724 | 0.73183 | 1.6112 | 0.99217 | 0.012638 | 0.004263 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR77; reps off
EN655 | 20200727 | L10 | 12 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 0.668 | 0.51 | 0.73183 | -0.063833 | -0.22183 | 0.0032259 | -0.0070048 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR78; reps off, very low RA/RB values and neg phaeo; Rb read twice on fluor before acidifying so it was noticed in lab, seems more likely a field sampling issue
EN655 | 20200727 | L10 | 12 | 18 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 4.017 | 2.688 | 0.73183 | 3.2852 | 1.9562 | 0.014403 | 0.0032847 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR87; 20s Rb values greater than 5s but calculated chl and phaeo about the same; leave as is; seems more likely there was an issue with 5s
EN655 | 20200727 | L10 | 12 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 5.264 | 3.558 | 0.73183 | 4.5322 | 2.8262 | 0.018556 | 0.0070912 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR88; 20s Rb values greater than 5s but calculated chl and phaeo about the same; leave as is; seems more likely there was an issue with 5s
EN655 | 20200727 | L9 | 15 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.6 | 15.32 | 0.4865 | 20.1135 | 14.8335 | 0.11217 | 0.15198 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L9 | 15 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.74 | 13.1 | 0.4865 | 17.2535 | 12.6135 | 0.098574 | 0.12605 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L9 | 15 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 219.7 | 148.3 | 0.58 | 219.12 | 147.72 | 1.5168 | 1.1137 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L9 | 15 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 226.4 | 148.7 | 0.4865 | 225.9135 | 148.2135 | 1.6507 | 0.98868 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L9 | 15 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 168.1 | 108.5 | 0.73183 | 167.3682 | 107.7682 | 1.2169 | 0.61893 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR95
EN655 | 20200727 | L9 | 15 | 9 |  |  | b | 292 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 142.5 | 90.7 | 0.73183 | 141.7682 | 89.9682 | 1.0612 | 0.47659 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR96
EN655 | 20200727 | L9 | 15 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 243.4 | 155.1 | 0.32729 | 243.0727 | 154.7727 | 0.96042 | 0.44411 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR101
EN655 | 20200727 | L9 | 15 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 216 | 140.5 | 0.32729 | 215.6727 | 140.1727 | 0.8212 | 0.45084 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR102
EN655 | 20200727 | L9 | 15 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 140.9 | 88.79 | 0.4865 | 140.4135 | 88.3035 | 1.107 | 0.46545 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L9 | 15 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 137.3 | 85.5 | 0.4865 | 136.8135 | 85.0135 | 1.1005 | 0.41345 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L9 | 15 | 11 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 77.36 | 49.39 | 0.73183 | 76.6282 | 48.6582 | 0.56913 | 0.25693 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR97
EN655 | 20200727 | L9 | 15 | 11 |  |  | b | 295 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 62.81 | 39.5 | 0.73183 | 62.0782 | 38.7682 | 0.4727 | 0.18323 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR98
EN655 | 20200727 | L9 | 15 | 11 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 93.14 | 58.74 | 0.32729 | 92.8127 | 58.4127 | 0.37348 | 0.15564 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR103
EN655 | 20200727 | L9 | 15 | 11 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 109.6 | 67.7 | 0.32729 | 109.2727 | 67.3727 | 0.45491 | 0.15538 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR104
EN655 | 20200727 | L9 | 15 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 35.69 | 23.12 | 0.58 | 35.11 | 22.54 | 0.26704 | 0.13435 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L9 | 15 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.07 | 23.89 | 0.4865 | 35.5835 | 23.4035 | 0.25876 | 0.15801 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L9 | 15 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23.85 | 15.43 | 0.58 | 23.27 | 14.85 | 0.17888 | 0.08557 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L9 | 15 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.91 | 13.01 | 0.4865 | 20.4235 | 12.5235 | 0.16783 | 0.055187 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L9 | 15 | 22 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.19 | 6.528 | 0.58 | 9.61 | 5.948 | 0.077797 | 0.028124 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L9 | 15 | 22 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.1 | 6.424 | 0.4865 | 9.6135 | 5.9375 | 0.078094 | 0.02764 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200727 | L9 | 15 | 22 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.461 | 5.419 | 0.76967 | 7.6913 | 4.6493 | 0.064625 | 0.018169 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200727 | L9 | 15 | 22 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.735 | 5.592 | 0.76967 | 7.9653 | 4.8223 | 0.066771 | 0.019105 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200727 | L9 | 15 | 22 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 1.434 | 0.972 | 0.32729 | 1.1067 | 0.64471 | 0.0094328 | 0.0015497 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR99
EN655 | 20200727 | L9 | 15 | 22 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 1.121 | 0.905 | 0.32729 | 0.79371 | 0.57771 | 0.0044101 | 0.0054311 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR100
EN655 | 20200727 | L9 | 15 | 22 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 2.9 | 1.896 | 0.73183 | 2.1682 | 1.1642 | 0.01092 | -0.00035571 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR93
EN655 | 20200727 | L9 | 15 | 22 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_111320b_12mm | 1.8343 | 0.5442 | 3.253 | 2.329 | 0.73183 | 2.5212 | 1.5972 | 0.010014 | 0.0044277 | 20201113 | 10-99 | 20200825 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR94
EN655 | 20200728 | L6 | 16 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.74 | 12.62 | 0.4865 | 14.2535 | 12.1335 | 0.045038 | 0.17103 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200728 | L6 | 16 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.86 | 12.68 | 0.4865 | 14.3735 | 12.1935 | 0.046313 | 0.17083 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200728 | L6 | 16 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 37.97 | 23.12 | 0.4865 | 37.4835 | 22.6335 | 0.31548 | 0.087576 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (56% range)
EN655 | 20200728 | L6 | 16 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 30.37 | 22.03 | 0.4865 | 29.8835 | 21.5435 | 0.17718 | 0.20647 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (56% range) ; was Rb recorded properly from flourometer?
EN655 | 20200728 | L6 | 16 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.19 | 43.85 | 0.4865 | 65.7035 | 43.3635 | 0.4746 | 0.29761 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200728 | L6 | 16 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.25 | 42.76 | 0.4865 | 63.7635 | 42.2735 | 0.45654 | 0.29626 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200728 | L6 | 16 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 221 | 141 | 0.4865 | 220.5135 | 140.5135 | 1.6995 | 0.8027 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200728 | L6 | 16 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 215.9 | 139.4 | 0.4865 | 215.4135 | 138.9135 | 1.6252 | 0.84856 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200728 | L6 | 16 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 183.5 | 116.3 | 0.76967 | 182.7303 | 115.5303 | 1.4276 | 0.62973 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200728 | L6 | 16 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 175.3 | 113.3 | 0.76967 | 174.5303 | 112.5303 | 1.3171 | 0.68678 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200728 | L6 | 16 | 7 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.08 | 31.49 | 0.32729 | 51.7527 | 31.1627 | 0.43442 | 0.11672 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR107
EN655 | 20200728 | L6 | 16 | 7 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.9 | 32.04 | 0.32729 | 49.5727 | 31.7127 | 0.37554 | 0.18341 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR108
EN655 | 20200728 | L6 | 16 | 7 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.94 | 12.83 | 0.32729 | 19.6127 | 12.5027 | 0.079643 | 0.037752 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR113
EN655 | 20200728 | L6 | 16 | 7 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.14 | 12.82 | 0.32729 | 19.8127 | 12.4927 | 0.081995 | 0.035306 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR114
EN655 | 20200728 | L6 | 16 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 134 | 86.97 | 0.58 | 133.42 | 86.39 | 0.99912 | 0.5393 | 20220519 | 11-69 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200728 | L6 | 16 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 135.2 | 87.66 | 0.4865 | 134.7135 | 87.1735 | 1.01 | 0.54242 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200728 | L6 | 16 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 114.2 | 72.31 | 0.76967 | 113.4303 | 71.5403 | 0.88992 | 0.38406 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200728 | L6 | 16 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 118.3 | 74.24 | 0.76967 | 117.5303 | 73.4703 | 0.93603 | 0.37232 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200728 | L6 | 16 | 11 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 24.86 | 15.69 | 0.32729 | 24.5327 | 15.3627 | 0.19216 | 0.077695 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR109
EN655 | 20200728 | L6 | 16 | 11 |  |  | b | 295 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.28 | 15.56 | 0.32729 | 24.9527 | 15.2327 | 0.20299 | 0.06367 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR110
EN655 | 20200728 | L6 | 16 | 11 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.7 | 4.635 | 0.32729 | 6.3727 | 4.3077 | 0.023089 | 0.017285 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR115
EN655 | 20200728 | L6 | 16 | 11 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.564 | 4.165 | 0.32729 | 5.2367 | 3.8377 | 0.015643 | 0.020327 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR116
EN655 | 20200728 | L6 | 16 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 57.34 | 36.29 | 0.4865 | 56.8535 | 35.8035 | 0.44719 | 0.19039 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | Switch of <10 funnels A and B between surface and chl max
EN655 | 20200728 | L6 | 16 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 58.43 | 36.74 | 0.4865 | 57.9435 | 36.2535 | 0.46079 | 0.18481 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200728 | L6 | 16 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 37.47 | 23.35 | 0.4865 | 36.9835 | 22.8635 | 0.29997 | 0.10718 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200728 | L6 | 16 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 35.98 | 22.84 | 0.4865 | 35.4935 | 22.3535 | 0.27915 | 0.11892 | 20220519 | 11-67 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN655 | 20200728 | L6 | 16 | 17 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.96 | 15.94 | 0.76967 | 25.1903 | 15.1703 | 0.21287 | 0.057283 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200728 | L6 | 16 | 17 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 26.04 | 15.75 | 0.76967 | 25.2703 | 14.9803 | 0.2186 | 0.048164 | 20220519 | 11-98 | 20230210 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN655 | 20200728 | L6 | 16 | 17 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.43 | 6.323 | 0.32729 | 10.1027 | 5.9957 | 0.086357 | 0.019321 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR105
EN655 | 20200728 | L6 | 16 | 17 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 9.616 | 5.811 | 0.32729 | 9.2887 | 5.4837 | 0.080007 | 0.016646 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR106
EN655 | 20200728 | L6 | 16 | 17 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.044 | 3.65 | 0.32729 | 5.7167 | 3.3227 | 0.026719 | 0.0043666 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR111
EN655 | 20200728 | L6 | 16 | 17 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.699 | 3.788 | 0.32729 | 5.3717 | 3.4607 | 0.021406 | 0.011089 | 20201113 | 10-100 | 20200826 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR112
EN657 | 20201013 | L1 | 1 | 2 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 186.8 | 117.4 | 0.719 | 186.081 | 116.681 | 2.9487 | 1.207 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201013 | L1 | 1 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 178.7 | 114.8 | 0.59583 | 178.1042 | 114.2042 | 2.715 | 1.3524 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201013 | L1 | 1 | 2 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 172.3 | 113.5 | 0.6525 | 171.6475 | 112.8475 | 1.2492 | 0.76041 | 20220519 | 11-96 | 20230208 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201013 | L1 | 1 | 2 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 213.4 | 138.5 | 0.6525 | 212.7475 | 137.8475 | 1.5912 | 0.86357 | 20220519 | 11-96 | 20230208 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201013 | L1 | 1 | 2 |  |  | a | 295 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 202.1 | 126.2 | 0.66167 | 201.4383 | 125.5383 | 1.5851 | 0.61256 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR5 | Says Niskin 18 on tar sheet, but depth is 18. seems like a mistake and says Niskin 2 in HS logsheet
EN657 | 20201013 | L1 | 1 | 2 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 183.2 | 113 | 0.66167 | 182.5383 | 112.3383 | 1.4711 | 0.50223 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR6 | Says Niskin 18 on tar sheet, but depth is 18. seems like a mistake and says Niskin 2 in HS logsheet
EN657 | 20201013 | L1 | 1 | 2 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 176.9 | 105 | 0.66167 | 176.2383 | 104.3383 | 0.80393 | 0.17399 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR11 | Says Niskin 18 on tar sheet, but depth is 18. seems like a mistake and says Niskin 2 in HS logsheet
EN657 | 20201013 | L1 | 1 | 2 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 153.8 | 89.53 | 0.66167 | 153.1383 | 88.8683 | 0.71862 | 0.11431 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR12 | Says Niskin 18 on tar sheet, but depth is 18. seems like a mistake and says Niskin 2 in HS logsheet
EN657 | 20201013 | L1 | 1 | 8 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 165.6 | 109.5 | 0.59583 | 165.0042 | 108.9042 | 2.3836 | 1.4951 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201013 | L1 | 1 | 8 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 175.2 | 113.7 | 0.59583 | 174.6042 | 113.1042 | 2.6131 | 1.4152 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | filter fell on floor particle side down and saw fuzz on filter
EN657 | 20201013 | L1 | 1 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 207.1 | 136.9 | 0.6525 | 206.4475 | 136.2475 | 1.4914 | 0.93493 | 20220519 | 11-96 | 20230208 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201013 | L1 | 1 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 189.6 | 125.3 | 0.6525 | 188.9475 | 124.6475 | 1.366 | 0.8537 | 20220519 | 11-96 | 20230208 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201013 | L1 | 1 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 235.8 | 146 | 0.66167 | 235.1383 | 145.3383 | 1.8882 | 0.67346 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR3
EN657 | 20201013 | L1 | 1 | 8 |  |  | b | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 212.2 | 130.1 | 0.66167 | 211.5383 | 129.4383 | 1.7322 | 0.55702 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR4
EN657 | 20201013 | L1 | 1 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 246.4 | 145.3 | 0.66167 | 245.7383 | 144.6383 | 1.1325 | 0.22562 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR9
EN657 | 20201013 | L1 | 1 | 8 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 245.7 | 144 | 0.66167 | 245.0383 | 143.3383 | 1.1392 | 0.20669 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR10
EN657 | 20201013 | L1 | 1 | 22 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 189.9 | 124.7 | 0.59583 | 189.3042 | 124.1042 | 2.7703 | 1.6498 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201013 | L1 | 1 | 22 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 174 | 112.9 | 0.59583 | 173.4042 | 112.3042 | 2.5961 | 1.4037 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201013 | L1 | 1 | 22 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 186.1 | 120.8 | 0.6525 | 185.4475 | 120.1475 | 1.3873 | 0.75232 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201013 | L1 | 1 | 22 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 201.5 | 134.4 | 0.6525 | 200.8475 | 133.7475 | 1.4255 | 0.95626 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201013 | L1 | 1 | 22 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 165.6 | 102.2 | 0.66167 | 164.9383 | 101.5383 | 1.3331 | 0.45657 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR1
EN657 | 20201013 | L1 | 1 | 22 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 157.7 | 96.96 | 0.66167 | 157.0383 | 96.2983 | 1.2772 | 0.42014 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR2
EN657 | 20201013 | L1 | 1 | 22 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 163.5 | 96 | 0.66167 | 162.8383 | 95.3383 | 0.75337 | 0.13858 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR7
EN657 | 20201013 | L1 | 1 | 22 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 165.5 | 97.6 | 0.66167 | 164.8383 | 96.9383 | 0.76059 | 0.14963 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR8
EN657 | 20201014 | L3 | 2 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.98 | 47.26 | 0.719 | 63.261 | 46.541 | 0.35521 | 0.47359 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L3 | 2 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 62.38 | 44.97 | 0.719 | 61.661 | 44.251 | 0.36986 | 0.41815 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L3 | 2 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 143.8 | 94.07 | 0.72867 | 143.0713 | 93.3413 | 1.0565 | 0.60573 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L3 | 2 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 142.4 | 97.23 | 0.59583 | 141.8042 | 96.6342 | 0.95961 | 0.76124 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L3 | 2 | 6 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 61.49 | 42.02 | 0.66167 | 60.8283 | 41.3583 | 0.408 | 0.31848 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR17
EN657 | 20201014 | L3 | 2 | 6 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 60.93 | 42.74 | 0.66167 | 60.2683 | 42.0783 | 0.38118 | 0.35795 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR18
EN657 | 20201014 | L3 | 2 | 6 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 79.2 | 49.92 | 0.66167 | 78.5383 | 49.2583 | 0.32739 | 0.13429 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR23
EN657 | 20201014 | L3 | 2 | 6 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 78.12 | 51 | 0.66167 | 77.4583 | 50.3383 | 0.30323 | 0.16856 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR24
EN657 | 20201014 | L3 | 2 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 113.1 | 76.55 | 0.59583 | 112.5042 | 75.9542 | 0.77648 | 0.5761 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L3 | 2 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 122.5 | 83.6 | 0.59583 | 121.9042 | 83.0042 | 0.8264 | 0.65172 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L3 | 2 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.81 | 31.07 | 0.66167 | 45.1483 | 30.4083 | 0.30994 | 0.22603 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR15
EN657 | 20201014 | L3 | 2 | 8 |  |  | b | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 39.75 | 28.02 | 0.66167 | 39.0883 | 27.3583 | 0.24749 | 0.23637 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR16
EN657 | 20201014 | L3 | 2 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 44.74 | 29.19 | 0.66167 | 44.0783 | 28.5283 | 0.17418 | 0.093686 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR21; poor replication (44% range) but DF thinks just within bio variation
EN657 | 20201014 | L3 | 2 | 8 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 68.23 | 43.69 | 0.66167 | 67.5683 | 43.0283 | 0.27489 | 0.12913 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR22; poor replication (44% range) but DF thinks just within bio variation
EN657 | 20201014 | L3 | 2 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 102.9 | 71.3 | 0.59583 | 102.3042 | 70.7042 | 0.67132 | 0.58777 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L3 | 2 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 104.8 | 71.8 | 0.59583 | 104.2042 | 71.2042 | 0.70106 | 0.56693 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L3 | 2 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.34 | 53.63 | 0.6525 | 79.6875 | 52.9775 | 0.56744 | 0.37598 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L3 | 2 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 81.63 | 53.79 | 0.6525 | 80.9775 | 53.1375 | 0.59144 | 0.35482 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L3 | 2 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 30.77 | 20.46 | 0.66167 | 30.1083 | 19.7983 | 0.21679 | 0.13217 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR13
EN657 | 20201014 | L3 | 2 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 29.14 | 20.01 | 0.66167 | 28.4783 | 19.3483 | 0.19197 | 0.14905 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR14
EN657 | 20201014 | L3 | 2 | 12 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.23 | 21.37 | 0.66167 | 32.5683 | 20.7083 | 0.13285 | 0.061593 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR19
EN657 | 20201014 | L3 | 2 | 12 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 39.55 | 25.22 | 0.66167 | 38.8883 | 24.5583 | 0.15994 | 0.069821 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR20
EN657 | 20201014 | L2 | 5 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 224.4 | 149.6 | 0.59583 | 223.8042 | 149.0042 | 1.5891 | 1.0644 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (46% range)
EN657 | 20201014 | L2 | 5 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 159.3 | 99.09 | 0.719 | 158.581 | 98.371 | 2.5582 | 0.94531 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (46% range) but maybe acceptable range given repb was diluted? Maybe toss rep b?
EN657 | 20201014 | L2 | 5 | 5 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 152.4 | 95.5 | 0.72867 | 151.6713 | 94.7713 | 2.4176 | 0.95775 | 20220519 | 11-90 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L2 | 5 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 283.7 | 184 | 0.59583 | 283.1042 | 183.4042 | 2.1181 | 1.148 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L2 | 5 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 69.4 | 45.24 | 0.6525 | 68.7475 | 44.5875 | 0.51326 | 0.28075 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L2 | 5 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 73.85 | 49.9 | 0.6525 | 73.1975 | 49.2475 | 0.5088 | 0.36819 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L2 | 5 | 5 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 231.3 | 154.3 | 0.66167 | 230.6383 | 153.6383 | 1.6136 | 1.0852 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR29
EN657 | 20201014 | L2 | 5 | 5 |  |  | b | 295 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 200.3 | 130.2 | 0.66167 | 199.6383 | 129.5383 | 1.464 | 0.80372 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR30
EN657 | 20201014 | L2 | 5 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 355.1 | 227.8 | 0.66167 | 354.4383 | 227.1383 | 1.4234 | 0.7055 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR35
EN657 | 20201014 | L2 | 5 | 5 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 363.4 | 233.5 | 0.66167 | 362.7383 | 232.8383 | 1.4524 | 0.72985 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR36
EN657 | 20201014 | L2 | 5 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 234 | 152 | 0.59583 | 233.4042 | 151.4042 | 1.742 | 0.95415 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | somewhat poor replication (21% range)
EN657 | 20201014 | L2 | 5 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 289.8 | 188 | 0.59583 | 289.2042 | 187.4042 | 2.1627 | 1.1746 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | somewhat poor replication (21% range)
EN657 | 20201014 | L2 | 5 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 69.12 | 46.29 | 0.6525 | 68.4675 | 45.6375 | 0.48501 | 0.3277 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L2 | 5 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 88.34 | 60.07 | 0.6525 | 87.6875 | 59.4175 | 0.60058 | 0.45752 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L2 | 5 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 216 | 140.9 | 0.66167 | 215.3383 | 140.2383 | 1.5791 | 0.89266 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR27
EN657 | 20201014 | L2 | 5 | 9 |  |  | b | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 216 | 142.2 | 0.66167 | 215.3383 | 141.5383 | 1.5571 | 0.94614 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR28
EN657 | 20201014 | L2 | 5 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 345.3 | 220.6 | 0.66167 | 344.6383 | 219.9383 | 1.3968 | 0.6683 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR33
EN657 | 20201014 | L2 | 5 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 350.2 | 226.1 | 0.66167 | 349.5383 | 225.4383 | 1.3901 | 0.72666 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR34
EN657 | 20201014 | L2 | 5 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 201.4 | 137.6 | 0.59583 | 200.8042 | 137.0042 | 1.3554 | 1.0844 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (33% range)
EN657 | 20201014 | L2 | 5 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 255.8 | 166.1 | 0.59583 | 255.2042 | 165.5042 | 1.9056 | 1.0417 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (33% range)
EN657 | 20201014 | L2 | 5 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.9 | 41.99 | 0.6525 | 63.2475 | 41.3375 | 0.46546 | 0.27067 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L2 | 5 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 72.24 | 48.34 | 0.6525 | 71.5875 | 47.6875 | 0.50774 | 0.34147 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L2 | 5 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 209.6 | 140.6 | 0.66167 | 208.9383 | 139.9383 | 1.4508 | 1.0156 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR25
EN657 | 20201014 | L2 | 5 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 173.3 | 113 | 0.66167 | 172.6383 | 112.3383 | 1.2679 | 0.71211 | 20220519 | 11-60 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR26
EN657 | 20201014 | L2 | 5 | 14 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 338 | 220.2 | 0.66167 | 337.3383 | 219.5383 | 1.3148 | 0.73915 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR31; higher than one 5 rep, but within bio variation
EN657 | 20201014 | L2 | 5 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 319.4 | 205.5 | 0.66167 | 318.7383 | 204.8383 | 1.2759 | 0.64749 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR32
EN657 | 20201014 | L4 | 6 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.17 | 26.73 | 0.719 | 33.451 | 26.011 | 0.15806 | 0.30514 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L4 | 6 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.7 | 26.98 | 0.59583 | 34.1042 | 26.3842 | 0.16401 | 0.30584 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L4 | 6 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.95 | 43.38 | 0.59583 | 64.3542 | 42.7842 | 0.45824 | 0.30365 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L4 | 6 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.29 | 42.31 | 0.59583 | 62.6942 | 41.7142 | 0.44571 | 0.29713 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L4 | 6 | 4 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22.89 | 15.32 | 0.66167 | 22.2283 | 14.6583 | 0.15863 | 0.098851 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR43
EN657 | 20201014 | L4 | 6 | 4 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.63 | 16.96 | 0.66167 | 24.9683 | 16.2983 | 0.18168 | 0.10461 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR44
EN657 | 20201014 | L4 | 6 | 4 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.15 | 19.36 | 0.66167 | 30.4883 | 18.6983 | 0.13207 | 0.043504 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR51
EN657 | 20201014 | L4 | 6 | 4 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 30.84 | 18.63 | 0.66167 | 30.1783 | 17.9683 | 0.13628 | 0.031829 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR52
EN657 | 20201014 | L4 | 6 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 137.9 | 89.15 | 0.719 | 137.181 | 88.431 | 1.0357 | 0.53911 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (29% range)
EN657 | 20201014 | L4 | 6 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 108.3 | 72.25 | 0.59583 | 107.7042 | 71.6542 | 0.76586 | 0.51015 | 20220519 | 11-82 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (29% range)
EN657 | 20201014 | L4 | 6 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 79.55 | 52.85 | 0.6525 | 78.8975 | 52.1975 | 0.56722 | 0.3623 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L4 | 6 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.86 | 53.52 | 0.6525 | 80.2075 | 52.8675 | 0.58082 | 0.36064 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L4 | 6 | 8 |  |  | a | 295 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 27.25 | 17.33 | 0.66167 | 26.5883 | 16.6683 | 0.20717 | 0.084625 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR41 | Are 5s/20s switched? Checked logsheet and no indication that they're switched but data tell otherwise; DF thinks they're switched
EN657 | 20201014 | L4 | 6 | 8 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 24.94 | 16.43 | 0.66167 | 24.2783 | 15.7683 | 0.17833 | 0.09865 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR42 | Are 5s/20s switched? Checked logsheet and no indication that they're switched but data tell otherwise; DF thinks they're switched
EN657 | 20201014 | L4 | 6 | 8 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 48.37 | 28.58 | 0.66167 | 47.7083 | 27.9183 | 0.22128 | 0.04039 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR49 | Are 5s/20s switched? Checked logsheet and no indication that they're switched but data tell otherwise; DF thinks they're switched
EN657 | 20201014 | L4 | 6 | 8 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 51.64 | 30.88 | 0.66167 | 50.9783 | 30.2183 | 0.23212 | 0.051101 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR50 | Are 5s/20s switched? Checked logsheet and no indication that they're switched but data tell otherwise; DF thinks they're switched
EN657 | 20201014 | L4 | 6 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 72.25 | 62.78 | 0.59583 | 71.6542 | 62.1842 | 0.20118 | 0.90618 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | reps off; RA isn't much lower than RB. I think I might have forgotten to record RB… so toss and assume other rep is fine
EN657 | 20201014 | L4 | 6 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 105.2 | 69.48 | 0.59583 | 104.6042 | 68.8842 | 0.75885 | 0.46783 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L4 | 6 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 72.27 | 46.96 | 0.6525 | 71.6175 | 46.3075 | 0.53769 | 0.28694 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L4 | 6 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 74.92 | 49.06 | 0.6525 | 74.2675 | 48.4075 | 0.54938 | 0.31266 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L4 | 6 | 13 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.49 | 16.85 | 0.66167 | 24.8283 | 16.1883 | 0.18167 | 0.10366 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR39 | Are 5s/20s switched? Checked logsheet and no indication that they're switched but data tell otherwise; DF thinks they're switched
EN657 | 20201014 | L4 | 6 | 13 |  |  | b | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23.68 | 15.41 | 0.66167 | 23.0183 | 14.7483 | 0.17449 | 0.08635 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR40 | Are 5s/20s switched? Checked logsheet and no indication that they're switched but data tell otherwise; DF thinks they're switched
EN657 | 20201014 | L4 | 6 | 13 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 37.42 | 22.25 | 0.66167 | 36.7583 | 21.5883 | 0.16993 | 0.032778 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR47 | Are 5s/20s switched? Checked logsheet and no indication that they're switched but data tell otherwise; DF thinks they're switched
EN657 | 20201014 | L4 | 6 | 13 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 47.83 | 29.09 | 0.66167 | 47.1683 | 28.4283 | 0.20992 | 0.057014 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR48 | Are 5s/20s switched? Checked logsheet and no indication that they're switched but data tell otherwise; DF thinks they're switched
EN657 | 20201014 | L4 | 6 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 116.7 | 76.21 | 0.72867 | 115.9713 | 75.4813 | 0.86018 | 0.48398 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L4 | 6 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 112.2 | 73.12 | 0.59583 | 111.6042 | 72.5242 | 0.83023 | 0.46127 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L4 | 6 | 17 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 73.08 | 47.17 | 0.6525 | 72.4275 | 46.5175 | 0.55044 | 0.27794 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L4 | 6 | 17 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 71.34 | 47.51 | 0.6525 | 70.6875 | 46.8575 | 0.50625 | 0.32818 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L4 | 6 | 17 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.28 | 10.4 | 0.66167 | 14.6183 | 9.7383 | 0.10261 | 0.069032 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR37 | Are 5s/20s switched? Checked logsheet and no indication that they're switched but data tell otherwise; DF thinks they're switched
EN657 | 20201014 | L4 | 6 | 17 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22.69 | 14.95 | 0.66167 | 22.0283 | 14.2883 | 0.16275 | 0.089092 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR38 | Are 5s/20s switched? Checked logsheet and no indication that they're switched but data tell otherwise; DF thinks they're switched
EN657 | 20201014 | L4 | 6 | 17 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.07 | 23.17 | 0.66167 | 37.4083 | 22.5083 | 0.1669 | 0.044441 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR45; poor replication (45% range) but DF thinks within bio variation | Are 5s/20s switched? Checked logsheet and no indication that they're switched but data tell otherwise; DF thinks they're switched
EN657 | 20201014 | L4 | 6 | 17 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 46.69 | 28.01 | 0.66167 | 46.0283 | 27.3483 | 0.20849 | 0.047373 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR46; poor replication (45% range) but DF thinks within bio variation | Are 5s/20s switched? Checked logsheet and no indication that they're switched but data tell otherwise; DF thinks they're switched
EN657 | 20201014 | L5 | 7 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 24.03 | 19.44 | 0.59583 | 23.4342 | 18.8442 | 0.097511 | 0.23806 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L5 | 7 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22.84 | 16.89 | 0.59583 | 22.2442 | 16.2942 | 0.1264 | 0.16376 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L5 | 7 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 44.23 | 29.31 | 0.59583 | 43.6342 | 28.7142 | 0.31697 | 0.19437 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L5 | 7 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 41.68 | 27.3 | 0.59583 | 41.0842 | 26.7042 | 0.30549 | 0.17005 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L5 | 7 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 77.45 | 47.93 | 0.719 | 76.731 | 47.211 | 0.62713 | 0.21359 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L5 | 7 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 77.32 | 50.1 | 0.719 | 76.601 | 49.381 | 0.57827 | 0.3011 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L5 | 7 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 29.52 | 19.01 | 0.66167 | 28.8583 | 18.3483 | 0.22099 | 0.10241 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR55
EN657 | 20201014 | L5 | 7 | 6 |  |  | b | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22 | 14.3 | 0.66167 | 21.3383 | 13.6383 | 0.16246 | 0.078745 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR56
EN657 | 20201014 | L5 | 7 | 6 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 30.94 | 19.2 | 0.66167 | 30.2783 | 18.5383 | 0.13151 | 0.042561 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR61
EN657 | 20201014 | L5 | 7 | 6 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 29.48 | 18.31 | 0.66167 | 28.8183 | 17.6483 | 0.12512 | 0.040589 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR62
EN657 | 20201014 | L5 | 7 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76.79 | 49.21 | 0.719 | 76.071 | 48.491 | 0.58592 | 0.2776 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L5 | 7 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.54 | 45.55 | 0.59583 | 69.9442 | 44.9542 | 0.5309 | 0.26964 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201014 | L5 | 7 | 9 |  |  | a | 295 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22.64 | 14.97 | 0.66167 | 21.9783 | 14.3083 | 0.16018 | 0.0903 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR57
EN657 | 20201014 | L5 | 7 | 9 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 26.85 | 17.82 | 0.66167 | 26.1883 | 17.1583 | 0.18923 | 0.11217 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR58
EN657 | 20201014 | L5 | 7 | 9 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23.64 | 15.23 | 0.66167 | 22.9783 | 14.5683 | 0.094034 | 0.042508 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR63; poor replication (40% range) but DF just thinks within bio variation
EN657 | 20201014 | L5 | 7 | 9 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.62 | 20.98 | 0.66167 | 32.9583 | 20.3183 | 0.14133 | 0.049104 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR64; poor replication (40% range) but DF just thinks within bio variation
EN657 | 20201014 | L5 | 7 | 12 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN657 | 20201014 | L5 | 7 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 85.99 | 55.75 | 0.59583 | 85.3942 | 55.1542 | 0.64243 | 0.33975 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L5 | 7 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 57.45 | 38.03 | 0.6525 | 56.7975 | 37.3775 | 0.41256 | 0.25305 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L5 | 7 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 57.89 | 38.97 | 0.6525 | 57.2375 | 38.3175 | 0.40194 | 0.28041 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201014 | L5 | 7 | 13 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.06 | 20.46 | 0.66167 | 32.3983 | 19.7983 | 0.26494 | 0.084018 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR53
EN657 | 20201014 | L5 | 7 | 13 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.72 | 23.42 | 0.66167 | 36.0583 | 22.7583 | 0.27966 | 0.12147 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR54
EN657 | 20201014 | L5 | 7 | 13 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 37.99 | 23.32 | 0.66167 | 37.3283 | 22.6583 | 0.16403 | 0.048338 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR59
EN657 | 20201014 | L5 | 7 | 13 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.72 | 28.25 | 0.66167 | 45.0583 | 27.5883 | 0.19534 | 0.063237 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR60
EN657 | 20201015 | L10 | 9 | 4 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.501 | 5.327 | 0.59583 | 5.9052 | 4.7312 | 0.025378 | 0.060352 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (41 % range) but low values to begin with
EN657 | 20201015 | L10 | 9 | 4 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.228 | 5.441 | 0.719 | 6.509 | 4.722 | 0.03863 | 0.046934 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (41 % range) but low values to begin with
EN657 | 20201015 | L10 | 9 | 6 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.01 | 15.23 | 0.59583 | 19.4142 | 14.6342 | 0.10333 | 0.16185 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L10 | 9 | 6 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.88 | 16.01 | 0.59583 | 20.2842 | 15.4142 | 0.10527 | 0.17403 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L10 | 9 | 6 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.595 | 5.213 | 0.66167 | 5.9333 | 4.5513 | 0.02896 | 0.050987 | 20220519 | 11-62 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR71
EN657 | 20201015 | L10 | 9 | 6 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.454 | 4.965 | 0.66167 | 5.7923 | 4.3033 | 0.031202 | 0.044388 | 20220519 | 11-62 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR72
EN657 | 20201015 | L10 | 9 | 6 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.797 | 3.138 | 0.66167 | 3.1353 | 2.4763 | 0.0073818 | 0.01587 | 20220519 | 11-62 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR79; poor replication (36% range) but low values and just within noise of flourometer?
EN657 | 20201015 | L10 | 9 | 6 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.899 | 2.945 | 0.66167 | 3.2373 | 2.2833 | 0.010648 | 0.010714 | 20220519 | 11-62 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR79; poor replication (36% range) but low values and just within noise of flourometer?
EN657 | 20201015 | L10 | 9 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.13 | 21.94 | 0.59583 | 30.5342 | 21.3442 | 0.19524 | 0.18486 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L10 | 9 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.57 | 25.25 | 0.72867 | 35.8413 | 24.5213 | 0.24049 | 0.19619 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L10 | 9 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 24.36 | 16.64 | 0.6525 | 23.7075 | 15.9875 | 0.16401 | 0.1207 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201015 | L10 | 9 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23.94 | 16.76 | 0.6525 | 23.2875 | 16.1075 | 0.15253 | 0.13431 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201015 | L10 | 9 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.73 | 41.35 | 0.719 | 64.011 | 40.631 | 0.49669 | 0.22686 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L10 | 9 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.88 | 42.02 | 0.72867 | 65.1513 | 41.2913 | 0.50689 | 0.22842 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L10 | 9 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 50.67 | 32.31 | 0.6525 | 50.0175 | 31.6575 | 0.39005 | 0.17371 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201015 | L10 | 9 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 48.63 | 30.92 | 0.6525 | 47.9775 | 30.2675 | 0.37624 | 0.16276 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201015 | L10 | 9 | 11 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 18.25 | 11.75 | 0.66167 | 17.5883 | 11.0883 | 0.13714 | 0.058965 | 20220519 | 11-62 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR69
EN657 | 20201015 | L10 | 9 | 11 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.5 | 10.78 | 0.66167 | 15.8383 | 10.1183 | 0.12027 | 0.058068 | 20220519 | 11-62 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR70
EN657 | 20201015 | L10 | 9 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.07 | 11.63 | 0.66167 | 15.4083 | 10.9683 | 0.049735 | 0.053253 | 20220519 | 11-62 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR77
EN657 | 20201015 | L10 | 9 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.47 | 14.57 | 0.66167 | 20.8083 | 13.9083 | 0.077291 | 0.053303 | 20220519 | 11-62 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR78
EN657 | 20201015 | L10 | 9 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 60.6 | 39.28 | 0.59583 | 60.0042 | 38.6842 | 0.45293 | 0.23595 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L10 | 9 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 59.1 | 39.3 | 0.59583 | 58.5042 | 38.7042 | 0.42064 | 0.2686 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L10 | 9 | 14 |  |  | a | 295 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.47 | 9.718 | 0.66167 | 13.8083 | 9.0563 | 0.099242 | 0.059299 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR67
EN657 | 20201015 | L10 | 9 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.35 | 11.65 | 0.66167 | 16.6883 | 10.9883 | 0.11945 | 0.073571 | 20220519 | 11-62 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR68
EN657 | 20201015 | L10 | 9 | 14 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 18.86 | 12.5 | 0.66167 | 18.1983 | 11.8383 | 0.071113 | 0.039843 | 20220519 | 11-62 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR75
EN657 | 20201015 | L10 | 9 | 14 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.47 | 10.63 | 0.66167 | 14.8083 | 9.9683 | 0.054117 | 0.039312 | 20220519 | 11-62 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR76
EN657 | 20201015 | L10 | 9 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.18 | 40.3 | 0.59583 | 62.5842 | 39.7042 | 0.48607 | 0.22098 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L10 | 9 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.55 | 42.39 | 0.719 | 64.831 | 41.671 | 0.49202 | 0.25005 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L10 | 9 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 47.57 | 30.77 | 0.6525 | 46.9175 | 30.1175 | 0.3569 | 0.17942 | 20220519 | 11-96 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201015 | L10 | 9 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.24 | 32.41 | 0.6525 | 48.5875 | 31.7575 | 0.35754 | 0.20799 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201015 | L10 | 9 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.07 | 13.18 | 0.66167 | 18.4083 | 12.5183 | 0.12385 | 0.096794 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR65
EN657 | 20201015 | L10 | 9 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.72 | 10.24 | 0.66167 | 15.0583 | 9.5783 | 0.11523 | 0.053596 | 20220519 | 11-61 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR66
EN657 | 20201015 | L10 | 9 | 18 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.61 | 12.11 | 0.66167 | 16.9483 | 11.4483 | 0.061385 | 0.04572 | 20220519 | 11-62 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR73
EN657 | 20201015 | L10 | 9 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.06 | 11.25 | 0.66167 | 15.3983 | 10.5883 | 0.053879 | 0.045541 | 20220519 | 11-62 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR74
EN657 | 20201015 | L11 | 12 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.82 | 12.21 | 0.59583 | 15.2242 | 11.6142 | 0.076692 | 0.13013 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L11 | 12 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.39 | 11.59 | 0.59583 | 14.7942 | 10.9942 | 0.080728 | 0.11505 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L11 | 12 | 4 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.916 | 3.826 | 0.74067 | 4.1753 | 3.0853 | 0.022841 | 0.031354 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR87
EN657 | 20201015 | L11 | 12 | 4 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.824 | 3.748 | 0.74067 | 4.0833 | 3.0073 | 0.022548 | 0.030278 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR88
EN657 | 20201015 | L11 | 12 | 4 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.896 | 2.379 | 0.74067 | 2.1553 | 1.6383 | 0.0057912 | 0.0095921 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR95
EN657 | 20201015 | L11 | 12 | 4 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.535 | 2.25 | 0.74067 | 1.7943 | 1.5093 | 0.0031809 | 0.01094 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR96
EN657 | 20201015 | L11 | 12 | 6 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN657 | 20201015 | L11 | 12 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 42.13 | 30.25 | 0.59583 | 41.5342 | 29.6542 | 0.25238 | 0.27569 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201015 | L11 | 12 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.45 | 45.67 | 0.59583 | 65.8542 | 45.0742 | 0.44146 | 0.36122 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L11 | 12 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.5 | 44.11 | 0.59583 | 64.9042 | 43.5142 | 0.45442 | 0.32048 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L11 | 12 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 58.78 | 38.68 | 0.6525 | 58.1275 | 38.0275 | 0.42701 | 0.25018 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | value close to WSW, seems suspicious
EN657 | 20201015 | L11 | 12 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 55.34 | 37.91 | 0.6525 | 54.6875 | 37.2575 | 0.37029 | 0.29319 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201015 | L11 | 12 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.84 | 9.291 | 0.74067 | 13.0993 | 8.5503 | 0.095651 | 0.055053 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR85
EN657 | 20201015 | L11 | 12 | 8 |  |  | b | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.34 | 8.567 | 0.74067 | 11.5993 | 7.8263 | 0.079606 | 0.05881 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR86
EN657 | 20201015 | L11 | 12 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.37 | 7.795 | 0.74067 | 11.6293 | 7.0543 | 0.051247 | 0.01499 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR93
EN657 | 20201015 | L11 | 12 | 8 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.19 | 8.667 | 0.74067 | 12.4493 | 7.9263 | 0.050665 | 0.02376 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR94
EN657 | 20201015 | L11 | 12 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 51.93 | 34.46 | 0.59583 | 51.3342 | 33.8642 | 0.37114 | 0.23191 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L11 | 12 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 56.33 | 37.74 | 0.59583 | 55.7342 | 37.1442 | 0.39493 | 0.26653 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201015 | L11 | 12 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 42.45 | 28.63 | 0.6525 | 41.7975 | 27.9775 | 0.2936 | 0.20462 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201015 | L11 | 12 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.36 | 30.81 | 0.6525 | 44.7075 | 30.1575 | 0.3091 | 0.22794 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201015 | L11 | 12 | 13 |  |  | a | 295 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.39 | 8.498 | 0.74067 | 11.6493 | 7.7573 | 0.081282 | 0.054519 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR83
EN657 | 20201015 | L11 | 12 | 13 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.08 | 8.209 | 0.74067 | 11.3393 | 7.4683 | 0.081118 | 0.050068 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR84
EN657 | 20201015 | L11 | 12 | 13 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.52 | 7.138 | 0.74067 | 9.7793 | 6.3973 | 0.037815 | 0.022144 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR91
EN657 | 20201015 | L11 | 12 | 13 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 9.407 | 6.269 | 0.74067 | 8.6663 | 5.5283 | 0.035087 | 0.016728 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR92
EN657 | 20201015 | L11 | 12 | 18 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN657 | 20201015 | L11 | 12 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.56 | 25.55 | 0.59583 | 37.9642 | 24.9542 | 0.27639 | 0.16799 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201015 | L11 | 12 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.01 | 21.79 | 0.6525 | 32.3575 | 21.1375 | 0.23836 | 0.13805 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201015 | L11 | 12 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.24 | 21.95 | 0.6525 | 32.5875 | 21.2975 | 0.23985 | 0.13941 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201015 | L11 | 12 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.515 | 5.066 | 0.74067 | 6.7743 | 4.3253 | 0.051495 | 0.024742 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR81
EN657 | 20201015 | L11 | 12 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.775 | 5.045 | 0.74067 | 6.0343 | 4.3043 | 0.036376 | 0.03949 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR82
EN657 | 20201015 | L11 | 12 | 18 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.888 | 4.668 | 0.74067 | 6.1473 | 3.9273 | 0.024777 | 0.011965 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR89
EN657 | 20201015 | L11 | 12 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.484 | 4.997 | 0.74067 | 6.7433 | 4.2563 | 0.027858 | 0.012107 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR90
EN657 | 20201016 | L9 | 14 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 9.366 | 7.953 | 0.72867 | 8.6373 | 7.2243 | 0.030018 | 0.098632 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L9 | 14 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 9.957 | 8.223 | 0.59583 | 9.3612 | 7.6272 | 0.036838 | 0.098986 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L9 | 14 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.03 | 10.76 | 0.59583 | 13.4342 | 10.1642 | 0.069469 | 0.11153 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L9 | 14 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.86 | 11.27 | 0.59583 | 14.2642 | 10.6742 | 0.076267 | 0.11382 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L9 | 14 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 46.48 | 32.85 | 0.719 | 45.761 | 32.131 | 0.28956 | 0.28262 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (100% range) which one to keep?
EN657 | 20201016 | L9 | 14 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 18.24 | 13.71 | 0.59583 | 17.6442 | 13.1142 | 0.096237 | 0.1373 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (100% range) which one to keep?
EN657 | 20201016 | L9 | 14 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.05 | 27.01 | 0.72867 | 39.3213 | 26.2813 | 0.27703 | 0.19099 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L9 | 14 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 32.53 | 22.09 | 0.59583 | 31.9342 | 21.4942 | 0.22179 | 0.16097 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L9 | 14 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.35 | 43.12 | 0.59583 | 63.7542 | 42.5242 | 0.45102 | 0.30625 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L9 | 14 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 69.08 | 45.73 | 0.72867 | 68.3513 | 45.0013 | 0.49606 | 0.30532 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L9 | 14 | 11 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14 | 9.924 | 0.74067 | 13.2593 | 9.1833 | 0.085999 | 0.076417 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR101
EN657 | 20201016 | L9 | 14 | 11 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.34 | 9.277 | 0.74067 | 12.5993 | 8.5363 | 0.085432 | 0.065025 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR102
EN657 | 20201016 | L9 | 14 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.89 | 8.058 | 0.74067 | 11.1493 | 7.3173 | 0.042924 | 0.025782 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR107
EN657 | 20201016 | L9 | 14 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.69 | 8.371 | 0.74067 | 11.9493 | 7.6303 | 0.04838 | 0.023266 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR108
EN657 | 20201016 | L9 | 14 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 67.52 | 44.96 | 0.59583 | 66.9242 | 44.3642 | 0.47927 | 0.31076 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L9 | 14 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.76 | 47.03 | 0.59583 | 70.1642 | 46.4342 | 0.50413 | 0.32277 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L9 | 14 | 13 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.7 | 9.399 | 0.74067 | 12.9593 | 8.6583 | 0.090129 | 0.06196 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR99
EN657 | 20201016 | L9 | 14 | 13 |  |  | b | 295 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.07 | 8.921 | 0.74067 | 12.3293 | 8.1803 | 0.086649 | 0.056556 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR100
EN657 | 20201016 | L9 | 14 | 13 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.54 | 7.125 | 0.74067 | 9.7993 | 6.3843 | 0.038184 | 0.021654 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR105
EN657 | 20201016 | L9 | 14 | 13 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.85 | 9.923 | 0.74067 | 14.1093 | 9.1823 | 0.05509 | 0.030972 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR106
EN657 | 20201016 | L9 | 14 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.15 | 46.06 | 0.59583 | 69.5542 | 45.4642 | 0.51178 | 0.29784 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L9 | 14 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.5 | 46.44 | 0.719 | 69.781 | 45.721 | 0.51114 | 0.30305 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L9 | 14 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 54.78 | 35.7 | 0.6525 | 54.1275 | 35.0475 | 0.40534 | 0.21878 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L9 | 14 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 57.06 | 37.23 | 0.6525 | 56.4075 | 36.5775 | 0.42128 | 0.23009 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L9 | 14 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.1 | 7.743 | 0.74067 | 10.3593 | 7.0023 | 0.070587 | 0.052833 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR97
EN657 | 20201016 | L9 | 14 | 16 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.1 | 8.241 | 0.74067 | 11.3593 | 7.5003 | 0.081142 | 0.051055 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR98
EN657 | 20201016 | L9 | 14 | 16 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.09 | 8.882 | 0.74067 | 12.3493 | 8.1413 | 0.046965 | 0.029202 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR103
EN657 | 20201016 | L9 | 14 | 16 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.2 | 6.897 | 0.74067 | 9.4593 | 6.1563 | 0.036999 | 0.020807 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR104
EN657 | 20201016 | L8 | 15 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.2 | 9.974 | 0.59583 | 11.6042 | 9.3782 | 0.04729 | 0.11972 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L8 | 15 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.82 | 9.77 | 0.59583 | 11.2242 | 9.1742 | 0.043551 | 0.11982 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L8 | 15 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.88 | 13.19 | 0.59583 | 16.2842 | 12.5942 | 0.078392 | 0.14588 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L8 | 15 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 18.91 | 14.7 | 0.72867 | 18.1813 | 13.9713 | 0.089439 | 0.15936 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L8 | 15 | 5 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.521 | 6.63 | 0.74067 | 7.7803 | 5.8893 | 0.039626 | 0.063823 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR115
EN657 | 20201016 | L8 | 15 | 5 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.029 | 6.342 | 0.74067 | 7.2883 | 5.6013 | 0.035352 | 0.063039 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR116
EN657 | 20201016 | L8 | 15 | 5 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.792 | 4.204 | 0.74067 | 5.0513 | 3.4633 | 0.017724 | 0.014678 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR123
EN657 | 20201016 | L8 | 15 | 5 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.998 | 3.672 | 0.74067 | 4.2573 | 2.9313 | 0.014853 | 0.012671 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR124
EN657 | 20201016 | L8 | 15 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 24.22 | 17.16 | 0.59583 | 23.6242 | 16.5642 | 0.14998 | 0.14499 | 20220519 | 11-83 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L8 | 15 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.16 | 18.11 | 0.72867 | 24.4313 | 17.3813 | 0.14977 | 0.15975 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L8 | 15 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.94 | 43.67 | 0.719 | 66.221 | 42.951 | 0.49436 | 0.27051 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L8 | 15 | 11 |  |  | b | 283 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.67 | 45.2 | 0.719 | 65.951 | 44.481 | 0.4674 | 0.34431 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L8 | 15 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.76 | 32.99 | 0.6525 | 49.1075 | 32.3375 | 0.35627 | 0.21959 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L8 | 15 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.35 | 29.33 | 0.6525 | 44.6975 | 28.6775 | 0.34033 | 0.17035 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L8 | 15 | 11 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.71 | 11 | 0.74067 | 14.9693 | 10.2593 | 0.099375 | 0.08207 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR113
EN657 | 20201016 | L8 | 15 | 11 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.86 | 11.57 | 0.74067 | 15.1193 | 10.8293 | 0.090205 | 0.10067 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR114
EN657 | 20201016 | L8 | 15 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 18.27 | 12.35 | 0.74067 | 17.5293 | 11.6093 | 0.066313 | 0.042694 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR121
EN657 | 20201016 | L8 | 15 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.1 | 11.85 | 0.74067 | 16.3593 | 11.1093 | 0.058808 | 0.045504 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR122
EN657 | 20201016 | L8 | 15 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 68.91 | 44.9 | 0.72867 | 68.1813 | 44.1713 | 0.51008 | 0.27652 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L8 | 15 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.73 | 43.63 | 0.59583 | 66.1342 | 43.0342 | 0.49074 | 0.2756 | 20220519 | 11-84 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L8 | 15 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 44.62 | 29.11 | 0.6525 | 43.9675 | 28.4575 | 0.3295 | 0.17727 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L8 | 15 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.95 | 30.6 | 0.6525 | 45.2975 | 29.9475 | 0.3261 | 0.2072 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L8 | 15 | 14 |  |  | a | 295 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.1 | 9.838 | 0.74067 | 13.3593 | 9.0973 | 0.089009 | 0.07025 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR111
EN657 | 20201016 | L8 | 15 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.36 | 9.807 | 0.74067 | 13.6193 | 9.0663 | 0.095409 | 0.063846 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR112
EN657 | 20201016 | L8 | 15 | 14 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.73 | 11.69 | 0.74067 | 15.9893 | 10.9493 | 0.056353 | 0.04627 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR119
EN657 | 20201016 | L8 | 15 | 14 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.26 | 13.89 | 0.74067 | 18.5193 | 13.1493 | 0.060043 | 0.0632 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR120
EN657 | 20201016 | L8 | 15 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.36 | 43.4 | 0.59583 | 63.7642 | 42.8042 | 0.44528 | 0.31697 | 20220519 | 11-84 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L8 | 15 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.2 | 44.43 | 0.72867 | 65.4713 | 43.7013 | 0.46249 | 0.31574 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L8 | 15 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.33 | 29.59 | 0.6525 | 44.6775 | 28.9375 | 0.33439 | 0.18093 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L8 | 15 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.68 | 29.83 | 0.6525 | 45.0275 | 29.1775 | 0.33672 | 0.18287 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L8 | 15 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.15 | 10.17 | 0.74067 | 14.4093 | 9.4293 | 0.10471 | 0.061484 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR109
EN657 | 20201016 | L8 | 15 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.49 | 10.06 | 0.74067 | 13.7493 | 9.3193 | 0.093149 | 0.071109 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR110
EN657 | 20201016 | L8 | 15 | 18 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.37 | 9.406 | 0.74067 | 12.6293 | 8.6653 | 0.044242 | 0.036827 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR117; poor replication (61% range)
EN657 | 20201016 | L8 | 15 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23.19 | 15.74 | 0.74067 | 22.4493 | 14.9993 | 0.083452 | 0.057386 | 20220519 | 11-52 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR118; poor replication (61% range)
EN657 | 20201016 | L6 | 17 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 29.33 | 23.91 | 0.72867 | 28.6013 | 23.1813 | 0.11514 | 0.29767 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | TC change after DF2 | reps fine but phaeo larger than chl?
EN657 | 20201016 | L6 | 17 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.66 | 23.85 | 0.72867 | 27.9313 | 23.1213 | 0.10219 | 0.30956 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | TC change after DF2 | reps fine but phaeo larger than chl?
EN657 | 20201016 | L6 | 17 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 24.63 | 18.46 | 0.72867 | 23.9013 | 17.7313 | 0.13108 | 0.18468 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | TC change after DF2 | reps fine but phaeo larger than chl?
EN657 | 20201016 | L6 | 17 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 24.58 | 19.26 | 0.59583 | 23.9842 | 18.6642 | 0.11302 | 0.21935 | 20220519 | 11-84 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | TC change after DF2 | reps fine but phaeo larger than chl?
EN657 | 20201016 | L6 | 17 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 46.62 | 31.65 | 0.72867 | 45.8913 | 30.9213 | 0.31803 | 0.23262 | 20220519 | 11-89 | 20221203 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L6 | 17 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 47.36 | 31.41 | 0.72867 | 46.6313 | 30.6813 | 0.33885 | 0.20752 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L6 | 17 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 26.44 | 17.74 | 0.6525 | 25.7875 | 17.0875 | 0.18483 | 0.11947 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L6 | 17 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.15 | 17.09 | 0.6525 | 24.4975 | 16.4375 | 0.17123 | 0.12149 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L6 | 17 | 6 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.97 | 7.918 | 0.74067 | 11.2293 | 7.1773 | 0.084911 | 0.041163 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR131
EN657 | 20201016 | L6 | 17 | 6 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.62 | 8.677 | 0.74067 | 11.8793 | 7.9363 | 0.082627 | 0.05678 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR132
EN657 | 20201016 | L6 | 17 | 6 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.43 | 10.53 | 0.74067 | 14.6893 | 9.7893 | 0.054888 | 0.03703 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR139
EN657 | 20201016 | L6 | 17 | 6 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.25 | 10.05 | 0.74067 | 14.5093 | 9.3093 | 0.058037 | 0.029057 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR140
EN657 | 20201016 | L6 | 17 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 82.79 | 53.54 | 0.72867 | 82.0613 | 52.8113 | 0.6214 | 0.31906 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L6 | 17 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 86.74 | 55.45 | 0.719 | 86.021 | 54.731 | 0.66474 | 0.30991 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L6 | 17 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 91.36 | 58.92 | 0.72867 | 90.6313 | 58.1913 | 0.68917 | 0.3471 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L6 | 17 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 94.44 | 61.09 | 0.72867 | 93.7113 | 60.3613 | 0.7085 | 0.36641 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L6 | 17 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 54.65 | 34.84 | 0.6525 | 53.9975 | 34.1875 | 0.42085 | 0.18796 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L6 | 17 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 61.71 | 38.8 | 0.6525 | 61.0575 | 38.1475 | 0.48671 | 0.19262 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L6 | 17 | 13 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.65 | 13.8 | 0.74067 | 19.9093 | 13.0593 | 0.14403 | 0.086144 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR129
EN657 | 20201016 | L6 | 17 | 13 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22.12 | 15.16 | 0.74067 | 21.3793 | 14.4193 | 0.14635 | 0.1078 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR130
EN657 | 20201016 | L6 | 17 | 13 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.48 | 23.44 | 0.74067 | 33.7393 | 22.6993 | 0.12322 | 0.089148 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR137
EN657 | 20201016 | L6 | 17 | 13 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.58 | 23.75 | 0.74067 | 35.8393 | 23.0093 | 0.14372 | 0.072333 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR138
EN657 | 20201016 | L6 | 17 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 72.55 | 46.93 | 0.72867 | 71.8213 | 46.2013 | 0.54428 | 0.27847 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L6 | 17 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 72.54 | 47.57 | 0.72867 | 71.8113 | 46.8413 | 0.53047 | 0.30367 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L6 | 17 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 53.28 | 34.38 | 0.6525 | 52.6275 | 33.7275 | 0.40152 | 0.1991 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L6 | 17 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 50.8 | 33.04 | 0.6525 | 50.1475 | 32.3875 | 0.3773 | 0.19945 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L6 | 17 | 15 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.95 | 9.303 | 0.74067 | 13.2093 | 8.5623 | 0.097379 | 0.053023 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR127
EN657 | 20201016 | L6 | 17 | 15 |  |  | b | 295 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.83 | 8.556 | 0.74067 | 12.0893 | 7.8153 | 0.089259 | 0.047556 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR128
EN657 | 20201016 | L6 | 17 | 15 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.68 | 11.3 | 0.74067 | 14.9393 | 10.5593 | 0.048974 | 0.049994 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR135
EN657 | 20201016 | L6 | 17 | 15 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.34 | 9.772 | 0.74067 | 13.5993 | 9.0313 | 0.051076 | 0.033571 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR136
EN657 | 20201016 | L6 | 17 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.67 | 46.49 | 0.72867 | 69.9413 | 45.7613 | 0.51369 | 0.30122 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L6 | 17 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.49 | 46.68 | 0.72867 | 69.7613 | 45.9513 | 0.50583 | 0.31247 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201016 | L6 | 17 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.93 | 32.19 | 0.6525 | 49.2775 | 31.5375 | 0.37687 | 0.18474 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L6 | 17 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.17 | 31.9 | 0.6525 | 48.5175 | 31.2475 | 0.36689 | 0.18956 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201016 | L6 | 17 | 18 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.48 | 15.4 | 0.74067 | 20.7393 | 14.6593 | 0.12828 | 0.13098 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR125
EN657 | 20201016 | L6 | 17 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 24.72 | 16.87 | 0.74067 | 23.9793 | 16.1293 | 0.16506 | 0.11923 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR126
EN657 | 20201016 | L6 | 17 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23.12 | 15.16 | 0.74067 | 22.3793 | 14.4193 | 0.089164 | 0.046227 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR133; poor replication (33 %range) but DF thinks within bio variation
EN657 | 20201016 | L6 | 17 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 37.51 | 26.32 | 0.74067 | 36.7693 | 25.5793 | 0.12535 | 0.11483 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR134; poor replication (33 %range) but DF thinks within bio variation
EN657 | 20201017 | L7 | 19 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.42 | 11.29 | 0.72867 | 13.6913 | 10.5613 | 0.066495 | 0.12158 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201017 | L7 | 19 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.72 | 10.81 | 0.719 | 13.001 | 10.091 | 0.061821 | 0.11788 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201017 | L7 | 19 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.34 | 10.14 | 0.72867 | 11.6113 | 9.4113 | 0.046738 | 0.12086 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201017 | L7 | 19 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.1 | 9.016 | 0.72867 | 10.3713 | 8.2873 | 0.044273 | 0.10331 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201017 | L7 | 19 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.49 | 11.36 | 0.719 | 13.771 | 10.641 | 0.066495 | 0.123 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201017 | L7 | 19 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.5 | 12.46 | 0.719 | 14.781 | 11.741 | 0.064583 | 0.1445 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201017 | L7 | 19 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.48 | 19.22 | 0.72867 | 27.7513 | 18.4913 | 0.19672 | 0.13257 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201017 | L7 | 19 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 30.86 | 21.51 | 0.72867 | 30.1313 | 20.7813 | 0.19863 | 0.17144 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201017 | L7 | 19 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.21 | 45.33 | 0.72867 | 69.4813 | 44.6013 | 0.52856 | 0.2657 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201017 | L7 | 19 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76.56 | 49.23 | 0.72867 | 75.8313 | 48.5013 | 0.58061 | 0.2831 | 20220519 | 11-89 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201017 | L7 | 19 | 9 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.31 | 10.8 | 0.74067 | 15.5693 | 10.0593 | 0.11546 | 0.061234 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR145
EN657 | 20201017 | L7 | 19 | 9 |  |  | b | 295 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.83 | 10.24 | 0.74067 | 15.0893 | 9.4993 | 0.11674 | 0.049553 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR146
EN657 | 20201017 | L7 | 19 | 9 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.21 | 10.96 | 0.74067 | 16.4693 | 10.2193 | 0.069883 | 0.025899 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR151
EN657 | 20201017 | L7 | 19 | 9 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.3 | 9.928 | 0.74067 | 14.5593 | 9.1873 | 0.060066 | 0.026043 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR152
EN657 | 20201017 | L7 | 19 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 73.95 | 47.43 | 0.72867 | 73.2213 | 46.7013 | 0.5634 | 0.26825 | 20220519 | 11-90 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201017 | L7 | 19 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 69.38 | 45.35 | 0.72867 | 68.6513 | 44.6213 | 0.5105 | 0.28411 | 20220519 | 11-90 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201017 | L7 | 19 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.41 | 33.9 | 0.6525 | 51.7575 | 33.2475 | 0.39323 | 0.19883 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | added N11 <10 rows. not in sosik logsheet. Assume vol filt 290ml
EN657 | 20201017 | L7 | 19 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 50.62 | 32.48 | 0.6525 | 49.9675 | 31.8275 | 0.38537 | 0.18141 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | added N11 <10 rows. not in sosik logsheet. Assume vol filt 290ml
EN657 | 20201017 | L7 | 19 | 11 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.91 | 9.743 | 0.74067 | 14.1693 | 9.0023 | 0.10902 | 0.050197 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR143
EN657 | 20201017 | L7 | 19 | 11 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.96 | 9.967 | 0.74067 | 14.2193 | 9.2263 | 0.10499 | 0.057632 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR144
EN657 | 20201017 | L7 | 19 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.44 | 7.98 | 0.74067 | 10.6993 | 7.2393 | 0.038757 | 0.029217 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR149; poor replication (60% range) but low values and within noise of flourometer?
EN657 | 20201017 | L7 | 19 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 18.56 | 12.09 | 0.74067 | 17.8193 | 11.3493 | 0.072474 | 0.034092 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR150; poor replication (60% range) but low values and within noise of flourometer?
EN657 | 20201017 | L7 | 19 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.98 | 51.62 | 0.719 | 80.261 | 50.901 | 0.62373 | 0.2827 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201017 | L7 | 19 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 74.08 | 48.1 | 0.72867 | 73.3513 | 47.3713 | 0.55193 | 0.29165 | 20220519 | 11-90 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201017 | L7 | 19 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.18 | 31.29 | 0.6525 | 48.5275 | 30.6375 | 0.38006 | 0.16553 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201017 | L7 | 19 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 51.59 | 33.02 | 0.6525 | 50.9375 | 32.3675 | 0.39451 | 0.18189 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201017 | L7 | 19 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.75 | 11.55 | 0.74067 | 16.0093 | 10.8093 | 0.10934 | 0.081181 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR141
EN657 | 20201017 | L7 | 19 | 15 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.66 | 11.2 | 0.74067 | 15.9193 | 10.4593 | 0.11481 | 0.069545 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR142
EN657 | 20201017 | L7 | 19 | 15 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.19 | 12.72 | 0.74067 | 16.4493 | 11.9793 | 0.04989 | 0.062184 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR147
EN657 | 20201017 | L7 | 19 | 15 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.71 | 8.761 | 0.74067 | 11.9693 | 8.0203 | 0.044235 | 0.031073 | 20220519 | 11-53 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR148
EN657 | 20201018 | MVCO | 20 | 2 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 174 | 111.3 | 0.72867 | 173.2713 | 110.5713 | 2.664 | 1.274 | 20220519 | 11-90 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201018 | MVCO | 20 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 168.9 | 111.1 | 0.72867 | 168.1713 | 110.3713 | 2.4558 | 1.4751 | 20220519 | 11-90 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201018 | MVCO | 20 | 4 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 175.6 | 113 | 0.72867 | 174.8713 | 112.2713 | 2.6598 | 1.3388 | 20220519 | 11-90 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201018 | MVCO | 20 | 4 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 171.3 | 111.6 | 0.719 | 170.581 | 110.881 | 2.5366 | 1.4125 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN657 | 20201018 | MVCO | 20 | 7 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 180.3 | 117.5 | 0.72867 | 179.5713 | 116.7713 | 2.6683 | 1.4906 | 20220519 | 11-90 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201018 | MVCO | 20 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 289.9 | 190.9 | 0.6525 | 289.2475 | 190.2475 | 2.1032 | 1.2847 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3
EN657 | 20201018 | MVCO | 20 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 159.3 | 106.3 | 0.6525 | 158.6475 | 105.6475 | 1.1259 | 0.75541 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN657 | 20201018 | MVCO | 20 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 193.5 | 127.2 | 0.6525 | 192.8475 | 126.5475 | 1.4085 | 0.84504 | 20220519 | 11-97 | 20230209 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR48A | 20201029 | L8 | 4 | 4 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 75.71 | 59.67 | 0.6765 | 75.0335 | 58.9935 | 0.093491 | 0.19474 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/01; CentralMooring and Glider583recovery | phae>chl; depth 52m
AR48A | 20201029 | L8 | 4 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 234.3 | 155.8 | 0.6765 | 233.6235 | 155.1235 | 0.91596 | 0.60128 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/02; CentralMooring and Glider583recovery
AR48A | 20201029 | L8 | 4 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 238.4 | 158 | 0.6765 | 237.7235 | 157.3235 | 0.93813 | 0.60063 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/03; CentralMooring and Glider583recovery
AR48A | 20201029 | L8 | 4 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 235.2 | 156.5 | 0.6765 | 234.5235 | 155.8235 | 0.91829 | 0.60579 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 |  | OOI 04/04; fail integrity check-spigot pushed in when come up on deck; CentralMooring and Glider583recovery
AR48A | 20201029 | L8 | 4 | 9 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 236.9 | 156.8 | 0.6765 | 236.2235 | 156.1235 | 0.93375 | 0.59183 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/05; CentralMooring and Glider583recovery
AR48A | 20201031 | L6 | 5 | 3 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 155.4 | 108.7 | 0.6765 | 154.7235 | 108.0235 | 0.2722 | 0.25558 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/01; InshoreMooring
AR48A | 20201031 | L6 | 5 | 4 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 130.7 | 87.66 | 0.6765 | 130.0235 | 86.9835 | 0.5022 | 0.34857 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/02; InshoreMooring
AR48A | 20201031 | L6 | 5 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 130.5 | 87.98 | 0.6765 | 129.8235 | 87.3035 | 0.49614 | 0.35777 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/03; InshoreMooring
AR48A | 20201031 | L6 | 5 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 129.5 | 87.78 | 0.6765 | 128.8235 | 87.1035 | 0.4868 | 0.36514 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/04; InshoreMooring
AR48A | 20201031 | L6 | 5 | 9 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 124.7 | 84.18 | 0.6765 | 124.0235 | 83.5035 | 0.47235 | 0.34361 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/05; InshoreMooring
AR48B | 20201104 | L10 | 1 | 8 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 42.6 | 33.77 | 0.62 | 41.98 | 33.15 | 0.051467 | 0.1105 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 01/01; not counted as OOI sample but still has ID; OffshoreMooring | phae>chl; depth 60m
AR48B | 20201104 | L10 | 1 | 9 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 215.8 | 143.4 | 0.62 | 215.18 | 142.78 | 0.84478 | 0.55172 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 01/02; not counted as OOI sample but still has ID; OffshoreMooring
AR48B | 20201104 | L10 | 1 | 9 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 214.8 | 143.3 | 0.62 | 214.18 | 142.68 | 0.83428 | 0.56125 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 01/03; not counted as OOI sample but still has ID; OffshoreMooring
AR48B | 20201104 | L10 | 1 | 10 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 198.5 | 130.3 | 0.62 | 197.88 | 129.68 | 0.79578 | 0.4726 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 01/04; Offshore Mooring
AR48B | 20201104 | L10 | 1 | 11 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 208.8 | 137.2 | 0.62 | 208.18 | 136.58 | 0.83466 | 0.49994 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 01/05; Offshore Mooring
AR48B | 20201105 | L12 | 2 | 8 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 32.92 | 27.02 | 0.62 | 32.3 | 26.4 | 0.034389 | 0.094596 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 02/01; not counted as OOI sample but still has ID; UpstreamOffshore | phae>chl; depth 60m
AR48B | 20201105 | L12 | 2 | 9 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 235.2 | 153.3 | 0.62 | 234.58 | 152.68 | 0.95563 | 0.53771 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 02/02; not counted as OOI sample but still has ID; UpstreamOffshore
AR48B | 20201105 | L12 | 2 | 9 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 218.1 | 144.6 | 0.62 | 217.48 | 143.98 | 0.85762 | 0.55063 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 02/03; not counted as OOI sample but still has ID; UpstreamOffshore
AR48B | 20201105 | L12 | 2 | 10 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 189.4 | 123.7 | 0.62 | 188.78 | 123.08 | 0.76661 | 0.43722 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 02/04; UpstreamOffshore
AR48B | 20201105 | L12 | 2 | 11 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 192.3 | 125.6 | 0.62 | 191.68 | 124.98 | 0.77754 | 0.44371 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 02/05; UpstreamOffshore
EN661 | 20210203 | L1 | 1 | 2 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 269.3 | 163.2 | 0.736 | 268.564 | 162.464 | 4.508 | 1.2782 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210203 | L1 | 1 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 265.5 | 156.4 | 0.736 | 264.764 | 155.664 | 4.6355 | 0.90857 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210203 | L1 | 1 | 2 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 125.7 | 83.85 | 0.62083 | 125.0792 | 83.2292 | 0.88908 | 0.59306 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN661 | 20210203 | L1 | 1 | 2 |  |  | b | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [6 empty] | Turner Designs Aquafluor Handheld Serial Number 800446 | 9 |  | missed Rb, reading lost
EN661 | 20210203 | L1 | 1 | 2 |  |  | a | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 211.8 | 127.6 | 0.76217 | 211.0378 | 126.8378 | 5.3114 | 1.3954 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR5; poor replication (but only 11% range) could this be a 20 rep b? seems likely | 5 um slow to filter at this station
EN661 | 20210203 | L1 | 1 | 2 |  |  | b | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 240 | 145.7 | 0.76217 | 239.2378 | 144.9378 | 5.9485 | 1.7153 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR6; poor replication (but only 11% range)
EN661 | 20210203 | L1 | 1 | 2 |  |  | a | 512 | 20 | 5 | 1 | 0 | 0.33 | Cal_051922a | 1.8382 | 0.56187 | 343.6 | 208.8 | 0.778 | 342.822 | 208.022 | 4.9153 | 1.4429 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR11; somewhat poor replication (17% range)
EN661 | 20210203 | L1 | 1 | 2 |  |  | b | 503 | 20 | 5 | 1 | 0 | 0.33 | Cal_051922a | 1.8382 | 0.56187 | 407.3 | 249 | 0.778 | 406.522 | 248.222 | 5.8754 | 1.8472 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR12; somewhat poor replication (17% range) but could this be a rep 5 a? seems likely
EN661 | 20210203 | L1 | 1 | 7 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 318.4 | 198.2 | 0.736 | 317.664 | 197.464 | 5.1071 | 1.9257 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication-- possible that these filters got switched between niskins? Seems likely based on values
EN661 | 20210203 | L1 | 1 | 7 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 297.2 | 186.8 | 0.736 | 296.464 | 186.064 | 4.6907 | 1.9361 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication-- possible that these filters got switched between niskins? Seems likely based on values
EN661 | 20210203 | L1 | 1 | 9 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 308.7 | 190.9 | 0.736 | 307.964 | 190.164 | 5.0052 | 1.7677 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication-- possible that these filters got switched between niskins? Seems likely based on values
EN661 | 20210203 | L1 | 1 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33 | Cal_051922a | 1.8382 | 0.56187 | 206.3 | 132.3 | 0.70817 | 205.5918 | 131.5918 | 4.7639 | 2.3372 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication-- possible that these filters got switched between niskins? Seems likely based on values
EN661 | 20210203 | L1 | 1 | 9 |  |  | a | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20220519 | [5 empty] | Turner Designs Aquafluor Handheld Serial Number 800446 | 9
EN661 | 20210203 | L1 | 1 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 121.4 | 82.35 | 0.70817 | 120.6918 | 81.6418 | 0.82959 | 0.62428 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN661 | 20210203 | L1 | 1 | 9 |  |  | a | 294 | 5 | 5 | 1 | 0 | 0.33 | Cal_051922a | 1.8382 | 0.56187 | 194.3 | 119 | 0.778 | 193.522 | 118.222 | 4.7816 | 1.5112 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR3
EN661 | 20210203 | L1 | 1 | 9 |  |  | b | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 203.1 | 127.9 | 0.76217 | 202.3378 | 127.1378 | 4.7436 | 1.979 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR4
EN661 | 20210203 | L1 | 1 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 134.3 | 78.22 | 0.76217 | 133.5378 | 77.4578 | 1.8845 | 0.29735 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR9; poor replication (87% range) DF thinks this sample should be tossed based on other >20 values in this cast, something is funky with this sample
EN661 | 20210203 | L1 | 1 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33 | Cal_051922a | 1.8382 | 0.56187 | 355.7 | 214.6 | 0.778 | 354.922 | 213.822 | 4.7895 | 1.2944 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR10; reps are off; poor replication (87% range)
EN661 | 20210203 | L1 | 1 | 15 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 308.7 | 194.2 | 0.736 | 307.964 | 193.464 | 4.865 | 2.0254 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210203 | L1 | 1 | 15 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 280.6 | 175.2 | 0.736 | 279.864 | 174.464 | 4.4783 | 1.7354 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210203 | L1 | 1 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 99.19 | 65.28 | 0.62083 | 98.5692 | 64.6592 | 0.7204 | 0.43105 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210203 | L1 | 1 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 111.4 | 75.17 | 0.70817 | 110.6918 | 74.4618 | 0.76968 | 0.55632 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210203 | L1 | 1 | 15 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 389.7 | 235 | 0.778 | 388.922 | 234.222 | 3.2418 | 0.87246 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR1; should this be a 20?
EN661 | 20210203 | L1 | 1 | 15 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 194.8 | 131.7 | 0.74067 | 194.0593 | 130.9593 | 1.3268 | 0.98144 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR2; note in lab notebook says DF isn't sure if this is a leftover sample from last extration time? it was underneath the lip of the hood -- toss
EN661 | 20210203 | L1 | 1 | 15 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 284.6 | 159.3 | 0.76217 | 283.8378 | 158.5378 | 4.2107 | 0.25516 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR7; somewhat poor replication (13%) and also value is larger than 5 rep -- should this be a 5?
EN661 | 20210203 | L1 | 1 | 15 |  |  | b | 552 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 263.3 | 153.5 | 0.76217 | 262.5378 | 152.7378 | 3.6764 | 0.61043 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR8; somewhat poor replication (13%) and also value is larger than 5 rep -- should this be a 5?
EN661 | 20210204 | L2 | 2 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 168.8 | 107 | 0.736 | 168.064 | 106.264 | 1.3129 | 0.57944 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L2 | 2 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 185 | 116.5 | 0.736 | 184.264 | 115.764 | 1.4552 | 0.60627 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L2 | 2 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 204 | 131.8 | 0.736 | 203.264 | 131.064 | 1.5338 | 0.80013 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L2 | 2 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 213.5 | 135.9 | 0.736 | 212.764 | 135.164 | 1.6486 | 0.75842 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L2 | 2 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 211.3 | 131.2 | 0.736 | 210.564 | 130.464 | 1.7017 | 0.62161 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L2 | 2 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 211.5 | 132.6 | 0.736 | 210.764 | 131.864 | 1.6762 | 0.67204 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L2 | 2 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 92.01 | 61.54 | 0.62083 | 91.3892 | 60.9192 | 0.64731 | 0.43753 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (50% range)
EN661 | 20210204 | L2 | 2 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 59.78 | 41.6 | 0.62083 | 59.1592 | 40.9792 | 0.38622 | 0.34353 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (50% range)
EN661 | 20210204 | L2 | 2 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 125.8 | 76.94 | 0.76217 | 125.0378 | 76.1778 | 1.0274 | 0.3153 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR17
EN661 | 20210204 | L2 | 2 | 7 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 148.7 | 91.36 | 0.76217 | 147.9378 | 90.5978 | 1.2016 | 0.38983 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR18
EN661 | 20210204 | L2 | 2 | 7 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 171.1 | 98.7 | 0.778 | 170.322 | 97.922 | 0.80952 | 0.10826 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR23
EN661 | 20210204 | L2 | 2 | 7 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 131.2 | 73.41 | 0.778 | 130.422 | 72.632 | 0.64616 | 0.034585 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR24
EN661 | 20210204 | L2 | 2 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 68.63 | 47.13 | 0.4865 | 68.1435 | 46.6435 | 0.45675 | 0.37387 | 20220519 | 11-65 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L2 | 2 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 72.67 | 50.04 | 0.736 | 71.934 | 49.304 | 0.48076 | 0.39724 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L2 | 2 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 249.5 | 164.2 | 0.70817 | 248.7918 | 163.4918 | 1.8121 | 1.0993 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L2 | 2 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 224.7 | 149.2 | 0.62083 | 224.0792 | 148.5792 | 1.6039 | 1.0419 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L2 | 2 | 11 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 154.6 | 93.08 | 0.778 | 153.822 | 92.302 | 1.2936 | 0.3333 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR15
EN661 | 20210204 | L2 | 2 | 11 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 145.9 | 88.72 | 0.778 | 145.122 | 87.942 | 1.2023 | 0.34771 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR16
EN661 | 20210204 | L2 | 2 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 180.8 | 102.2 | 0.76217 | 180.0378 | 101.4378 | 0.88044 | 0.072019 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR21; poor replication (48%) DF thinks within bio variation
EN661 | 20210204 | L2 | 2 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 111.7 | 63.89 | 0.778 | 110.922 | 63.112 | 0.53555 | 0.05705 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR22; poor replication (48%) DF thinks within bio variation
EN661 | 20210204 | L2 | 2 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 182.3 | 116 | 0.70817 | 181.5918 | 115.2918 | 1.4085 | 0.6446 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L2 | 2 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 170.7 | 110.3 | 0.70817 | 169.9918 | 109.5918 | 1.2832 | 0.66844 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | Tissue capsule unclear if N16 or N10. no N10 so def N16
EN661 | 20210204 | L2 | 2 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 72 | 50.52 | 0.62083 | 71.3792 | 49.8992 | 0.45633 | 0.43227 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L2 | 2 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 68.84 | 47.87 | 0.62083 | 68.2192 | 47.2492 | 0.44549 | 0.39591 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L2 | 2 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 187.8 | 114.3 | 0.76217 | 187.0378 | 113.5378 | 1.5455 | 0.45569 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR13
EN661 | 20210204 | L2 | 2 | 16 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 164.4 | 97.57 | 0.76217 | 163.6378 | 96.8078 | 1.4004 | 0.30004 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR14
EN661 | 20210204 | L2 | 2 | 16 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 246.6 | 140.7 | 0.76217 | 245.8378 | 139.9378 | 1.1862 | 0.12772 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR19 poor replication (33%) DF thinks within bio variation
EN661 | 20210204 | L2 | 2 | 16 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 179 | 103.5 | 0.76217 | 178.2378 | 102.7378 | 0.84265 | 0.11852 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR20; poor replication (33%) DF thinks within bio variation
EN661 | 20210204 | L4 | 5 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 183.3 | 113 | 0.70817 | 182.5918 | 112.2918 | 4.4804 | 1.5186 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | poor replication (14% range); note in Lab notebook that extraction volume was about 4.75 mls (spilled some), assume extraction was compromised and toss and assume other rep is fine
EN661 | 20210204 | L4 | 5 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 304.2 | 181.8 | 0.736 | 303.464 | 181.064 | 5.2006 | 1.2481 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | poor replication (14% range);
EN661 | 20210204 | L4 | 5 | 4 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 348.8 | 208.8 | 0.736 | 348.064 | 208.064 | 5.9484 | 1.4619 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (23%)
EN661 | 20210204 | L4 | 5 | 4 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 280 | 168.9 | 0.736 | 279.264 | 168.164 | 4.7205 | 1.2688 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (23%)
EN661 | 20210204 | L4 | 5 | 6 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 267.6 | 160.3 | 0.7505 | 266.8495 | 159.5495 | 4.559 | 1.1234 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L4 | 5 | 6 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 251.2 | 151.2 | 0.7505 | 250.4495 | 150.4495 | 4.2489 | 1.1095 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L4 | 5 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 43.38 | 33.52 | 0.70817 | 42.6718 | 32.8118 | 0.20947 | 0.37484 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L4 | 5 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 42.79 | 32.66 | 0.62083 | 42.1692 | 32.0392 | 0.21521 | 0.35534 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L4 | 5 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 183.5 | 103.2 | 0.76217 | 182.7378 | 102.4378 | 5.0654 | 0.35121 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR29
EN661 | 20210204 | L4 | 5 | 6 |  |  | b | 294 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 203.5 | 116 | 0.76217 | 202.7378 | 115.2378 | 5.5008 | 0.5719 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR30
EN661 | 20210204 | L4 | 5 | 6 |  |  | a | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 298.9 | 170 | 0.76217 | 298.1378 | 169.2378 | 4.3238 | 0.4348 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR35
EN661 | 20210204 | L4 | 5 | 6 |  |  | b | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 270.2 | 154.5 | 0.76217 | 269.4378 | 153.7378 | 3.881 | 0.44175 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR36
EN661 | 20210204 | L4 | 5 | 11 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 289.1 | 173.1 | 0.736 | 288.364 | 172.364 | 4.9287 | 1.2102 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L4 | 5 | 11 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 292.4 | 173.7 | 0.736 | 291.664 | 172.964 | 5.0434 | 1.1168 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L4 | 5 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 61.98 | 43.48 | 0.62083 | 61.3592 | 42.8592 | 0.39302 | 0.37021 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | several large cells visible, picked off a few from a, not from b; | poor replication (35% range)
EN661 | 20210204 | L4 | 5 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.97 | 37.05 | 0.62083 | 49.3492 | 36.4292 | 0.27448 | 0.37425 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | several large cells visible, picked off a few from a, not from b | poor replication (35% range)
EN661 | 20210204 | L4 | 5 | 11 |  |  | a | 292 | 5 | 5 | 1 | 0 | 0.33 | Cal_051922a | 1.8382 | 0.56187 | 153.1 | 86.27 | 0.778 | 152.322 | 85.492 | 4.2728 | 0.309 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR27
EN661 | 20210204 | L4 | 5 | 11 |  |  | b | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 180.2 | 104.9 | 0.76217 | 179.4378 | 104.1378 | 4.75 | 0.7565 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR28
EN661 | 20210204 | L4 | 5 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 339.7 | 193.1 | 0.76217 | 338.9378 | 192.3378 | 4.9264 | 0.49148 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR33
EN661 | 20210204 | L4 | 5 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 353.2 | 198.3 | 0.76217 | 352.4378 | 197.5378 | 5.2054 | 0.35904 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR34
EN661 | 20210204 | L4 | 5 | 23 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 262.5 | 159.5 | 0.736 | 261.764 | 158.764 | 4.3763 | 1.2782 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L4 | 5 | 23 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 280.6 | 165.6 | 0.736 | 279.864 | 164.864 | 4.8862 | 0.98556 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L4 | 5 | 23 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 48.79 | 35 | 0.62083 | 48.1692 | 34.3792 | 0.29296 | 0.31926 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN661 | 20210204 | L4 | 5 | 23 |  |  | b | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20220519 | [5 empty] | Turner Designs Aquafluor Handheld Serial Number 800446 | 9 |  | Funnel leaked, sample no good, crossed off on data sheet
EN661 | 20210204 | L4 | 5 | 23 |  |  | a | 294 | 5 | 5 | 1 | 0 | 0.33 | Cal_051922a | 1.8382 | 0.56187 | 174.5 | 100.4 | 0.778 | 173.722 | 99.622 | 4.7054 | 0.59737 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR25; poor replication (21% range ) but high chl …just in range of bio variation?
EN661 | 20210204 | L4 | 5 | 23 |  |  | b | 293 | 5 | 5 | 1 | 0 | 0.33 | Cal_051922a | 1.8382 | 0.56187 | 211.5 | 120.2 | 0.778 | 210.722 | 119.422 | 5.8174 | 0.561 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR26; poor replication (21% range ) but high chl …just in range of bio variation?
EN661 | 20210204 | L4 | 5 | 23 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 319.8 | 182.5 | 0.76217 | 319.0378 | 181.7378 | 4.6139 | 0.50542 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR31
EN661 | 20210204 | L4 | 5 | 23 |  |  | b | 552 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 286.1 | 158.3 | 0.76217 | 285.3378 | 157.5378 | 4.2791 | 0.14246 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR32
EN661 | 20210204 | L5 | 6 | 1 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 151.8 | 91.8 | 0.736 | 151.064 | 91.064 | 2.5493 | 0.69399 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L5 | 6 | 1 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 156.2 | 94.68 | 0.736 | 155.464 | 93.944 | 2.6139 | 0.73198 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L5 | 6 | 4 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 165.4 | 101.3 | 0.736 | 164.664 | 100.564 | 2.7235 | 0.85814 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L5 | 6 | 4 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 153.9 | 91.26 | 0.736 | 153.164 | 90.524 | 2.6615 | 0.56259 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L5 | 6 | 4 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 253.8 | 145.2 | 0.778 | 253.022 | 144.422 | 2.2913 | 0.2629 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR41
EN661 | 20210204 | L5 | 6 | 4 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 264.2 | 150.3 | 0.778 | 263.422 | 149.522 | 2.395 | 0.24045 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR42
EN661 | 20210204 | L5 | 6 | 4 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 387 | 217.5 | 0.778 | 386.222 | 216.722 | 1.8952 | 0.13602 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR47; poor replication (36 % range) but DF thinks within bio variation
EN661 | 20210204 | L5 | 6 | 4 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 271.4 | 154.3 | 0.778 | 270.622 | 153.522 | 1.3093 | 0.12957 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR48; poor replication (36 % range) but DF thinks within bio variation
EN661 | 20210204 | L5 | 6 | 6 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 159.9 | 96.44 | 0.736 | 159.164 | 95.704 | 2.6963 | 0.71224 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L5 | 6 | 6 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 142.7 | 86.5 | 0.736 | 141.964 | 85.764 | 2.3879 | 0.66669 | 20220519 | 11-79 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L5 | 6 | 6 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 249.4 | 145.8 | 0.76217 | 248.6378 | 145.0378 | 2.1858 | 0.37929 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR39
EN661 | 20210204 | L5 | 6 | 6 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 281.5 | 159.2 | 0.778 | 280.722 | 158.422 | 2.5716 | 0.22069 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR40
EN661 | 20210204 | L5 | 6 | 6 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 316.8 | 176.2 | 0.778 | 316.022 | 175.422 | 1.5749 | 0.072204 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR45
EN661 | 20210204 | L5 | 6 | 6 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 391.4 | 219.6 | 0.778 | 390.622 | 218.822 | 1.9244 | 0.13022 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR46
EN661 | 20210204 | L5 | 6 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 263.4 | 159.9 | 0.7505 | 262.6495 | 159.1495 | 2.1988 | 0.63532 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L5 | 6 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 252.9 | 153.5 | 0.736 | 252.164 | 152.764 | 2.1117 | 0.60871 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L5 | 6 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 311.1 | 188.8 | 0.736 | 310.364 | 188.064 | 2.5982 | 0.75084 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L5 | 6 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 277.3 | 168.5 | 0.736 | 276.564 | 167.764 | 2.3114 | 0.67614 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210204 | L5 | 6 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 128.1 | 80.12 | 0.62083 | 127.4792 | 79.4992 | 1.0193 | 0.39641 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | question, how is this <10? Many cells visible; poor replication (30% range)
EN661 | 20210204 | L5 | 6 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 96.73 | 61.54 | 0.62083 | 96.1092 | 60.9192 | 0.74759 | 0.33725 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | several large cells around edge but not in middle like in rep a; poor replication (30% range)
EN661 | 20210204 | L5 | 6 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 252.8 | 147.4 | 0.76217 | 252.0378 | 146.6378 | 2.2162 | 0.36834 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR37
EN661 | 20210204 | L5 | 6 | 12 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 261.7 | 152.6 | 0.76217 | 260.9378 | 151.8378 | 2.2862 | 0.3809 | 20220519 | 11-48 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR38; Combo of 2 surface niskins (other one is 14)
EN661 | 20210204 | L5 | 6 | 12 |  |  | a | 552 | 20 | 5 | 1 | 0 | 0.33 | Cal_051922a | 1.8382 | 0.56187 | 152.6 | 87.18 | 0.778 | 151.822 | 86.402 | 2.2126 | 0.23694 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR43; poor replication (45% range). This value is really close to the 5 values
EN661 | 20210204 | L5 | 6 | 12 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33 | Cal_051922a | 1.8382 | 0.56187 | 96.07 | 54.92 | 0.778 | 95.292 | 54.142 | 1.3968 | 0.14372 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR44; poor replication (45% range)
EN661 | 20210205 | L10 | 7 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 134.2 | 85.48 | 0.70817 | 133.4918 | 84.7718 | 1.035 | 0.47458 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | TC and DF2 | chl could not be this high at 430m; checked tissue capsule and label match logbook so was it mislabeled on the cruise; Must have sampled wrong water. Discard
EN661 | 20210205 | L10 | 7 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 147.9 | 93.24 | 0.7505 | 147.1495 | 92.4895 | 1.1612 | 0.48583 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | TC and DF2 | chl could not be this high at 430m; checked tissue capsule and label match logbook so was it mislabeled on the cruise; Must have sampled wrong water. Discard
EN661 | 20210205 | L10 | 7 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.643 | 4.423 | 0.736 | 4.907 | 3.687 | 0.025918 | 0.03974 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | poor replication (38% range) but values are low and within noise of flourometer
EN661 | 20210205 | L10 | 7 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.73 | 4.968 | 0.736 | 5.994 | 4.232 | 0.037433 | 0.03793 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | poor replication (38% range) but values are low and within noise of flourometer
EN661 | 20210205 | L10 | 7 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.06 | 10.02 | 0.736 | 13.324 | 9.284 | 0.085827 | 0.079501 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (39% range) but values are low and within noise of flourometer
EN661 | 20210205 | L10 | 7 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 9.844 | 7.149 | 0.7505 | 9.0935 | 6.3985 | 0.057253 | 0.05669 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (39% range) but values are low and within noise of flourometer
EN661 | 20210205 | L10 | 7 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 35.84 | 24.71 | 0.736 | 35.104 | 23.974 | 0.23645 | 0.19048 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L10 | 7 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.82 | 24.77 | 0.70817 | 33.1118 | 24.0618 | 0.19226 | 0.23623 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | phaeo higher than chl but 100m depth offshore at L10 so within reason
EN661 | 20210205 | L10 | 7 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 103 | 67.15 | 0.736 | 102.264 | 66.414 | 0.76161 | 0.42108 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L10 | 7 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 94.94 | 61.83 | 0.7505 | 94.1895 | 61.0795 | 0.7034 | 0.3843 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L10 | 7 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.6 | 25.89 | 0.62083 | 37.9792 | 25.2692 | 0.27002 | 0.17997 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (41%)
EN661 | 20210205 | L10 | 7 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 57.45 | 38.11 | 0.70817 | 56.7418 | 37.4018 | 0.41087 | 0.25518 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (41%)
EN661 | 20210205 | L10 | 7 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 75.59 | 48.53 | 0.778 | 74.812 | 47.752 | 0.56899 | 0.27267 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR53
EN661 | 20210205 | L10 | 7 | 8 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.13 | 41.3 | 0.778 | 64.352 | 40.522 | 0.49936 | 0.21243 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR54
EN661 | 20210205 | L10 | 7 | 8 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76.72 | 48.64 | 0.76217 | 75.9578 | 47.8778 | 0.31397 | 0.13477 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR59
EN661 | 20210205 | L10 | 7 | 8 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 82.8 | 51.59 | 0.76217 | 82.0378 | 50.8278 | 0.34897 | 0.12742 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR60
EN661 | 20210205 | L10 | 7 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 138.6 | 86.93 | 0.736 | 137.864 | 86.194 | 1.0977 | 0.43724 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L10 | 7 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 140 | 88.14 | 0.70817 | 139.2918 | 87.4318 | 1.1017 | 0.45524 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L10 | 7 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 157.9 | 99.65 | 0.736 | 157.164 | 98.914 | 1.2375 | 0.52397 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L10 | 7 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 137.4 | 86.66 | 0.736 | 136.664 | 85.924 | 1.0779 | 0.45219 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L10 | 7 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 57.6 | 37.05 | 0.62083 | 56.9792 | 36.4292 | 0.43657 | 0.21215 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | second RB value in lab notebook is 155.96? DF confirmed this was a mistake. Value recorded here is correct | poor replication (42% range)
EN661 | 20210205 | L10 | 7 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 86.97 | 55.45 | 0.62083 | 86.3492 | 54.8292 | 0.66962 | 0.30677 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 |  | poor replication (42% range)
EN661 | 20210205 | L10 | 7 | 15 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 117.4 | 72.05 | 0.76217 | 116.6378 | 71.2878 | 0.95683 | 0.30396 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR51
EN661 | 20210205 | L10 | 7 | 15 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 99.69 | 60.75 | 0.778 | 98.912 | 59.972 | 0.81878 | 0.23825 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR52
EN661 | 20210205 | L10 | 7 | 15 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 165 | 97.9 | 0.778 | 164.222 | 97.122 | 0.75162 | 0.16031 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR57
EN661 | 20210205 | L10 | 7 | 15 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 123.7 | 74.06 | 0.76217 | 122.9378 | 73.2978 | 0.55605 | 0.13219 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR58
EN661 | 20210205 | L10 | 7 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 134.1 | 87.89 | 0.736 | 133.364 | 87.154 | 0.9817 | 0.57033 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L10 | 7 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 148.9 | 93.61 | 0.736 | 148.164 | 92.874 | 1.1746 | 0.47929 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L10 | 7 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 61.72 | 40 | 0.62083 | 61.0992 | 39.3792 | 0.46143 | 0.23983 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L10 | 7 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 71 | 44.81 | 0.62083 | 70.3792 | 44.1892 | 0.55639 | 0.23053 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L10 | 7 | 18 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 89.74 | 53.87 | 0.76217 | 88.9778 | 53.1078 | 0.75167 | 0.1812 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR49
EN661 | 20210205 | L10 | 7 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 85.26 | 52.5 | 0.76217 | 84.4978 | 51.7378 | 0.68884 | 0.22307 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR50
EN661 | 20210205 | L10 | 7 | 18 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 147.5 | 89.19 | 0.778 | 146.722 | 88.412 | 0.6508 | 0.17635 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR55
EN661 | 20210205 | L10 | 7 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 167.3 | 100.8 | 0.778 | 166.522 | 100.022 | 0.7449 | 0.19426 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR56
EN661 | 20210205 | L11 | 10 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 39.6 | 27.7 | 0.70817 | 38.8918 | 26.9918 | 0.25281 | 0.22786 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | HS: possibly N1 and N2 sampled for chl, possibly N3 was not
EN661 | 20210205 | L11 | 10 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.31 | 27.28 | 0.4865 | 37.8235 | 26.7935 | 0.23432 | 0.24281 | 20220519 | 11-65 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | HS: possibly N1 and N2 sampled for chl, possibly N3 was not
EN661 | 20210205 | L11 | 10 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 56.37 | 39.5 | 0.70817 | 55.6618 | 38.7918 | 0.35839 | 0.33241 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | tissue capsule didn't have sample type but no 10 sample in this niskin so put as WSW 0
EN661 | 20210205 | L11 | 10 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 57.82 | 38.52 | 0.736 | 57.084 | 37.784 | 0.41002 | 0.26284 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L11 | 10 | 5 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.9 | 19.94 | 0.76217 | 28.1378 | 19.1778 | 0.18776 | 0.14911 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR67
EN661 | 20210205 | L11 | 10 | 5 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 30.02 | 20.17 | 0.76217 | 29.2578 | 19.4078 | 0.20641 | 0.1345 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR68
EN661 | 20210205 | L11 | 10 | 5 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.21 | 16.6 | 0.778 | 24.432 | 15.822 | 0.096096 | 0.051928 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR75
EN661 | 20210205 | L11 | 10 | 5 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.48 | 22.43 | 0.778 | 32.702 | 21.652 | 0.12378 | 0.079526 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR76
EN661 | 20210205 | L11 | 10 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 71.96 | 46.99 | 0.736 | 71.224 | 46.254 | 0.53047 | 0.29321 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L11 | 10 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 77.48 | 50.23 | 0.736 | 76.744 | 49.494 | 0.57891 | 0.30247 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L11 | 10 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 47.32 | 30.81 | 0.70817 | 46.6118 | 30.1018 | 0.35074 | 0.18531 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (32%)
EN661 | 20210205 | L11 | 10 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.52 | 24.57 | 0.70817 | 35.8118 | 23.8618 | 0.25387 | 0.17106 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (32%)
EN661 | 20210205 | L11 | 10 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 44.75 | 29.14 | 0.76217 | 43.9878 | 28.3778 | 0.32823 | 0.17195 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR65
EN661 | 20210205 | L11 | 10 | 8 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 39.17 | 25.05 | 0.76217 | 38.4078 | 24.2878 | 0.29589 | 0.13074 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR66
EN661 | 20210205 | L11 | 10 | 8 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.55 | 41.66 | 0.778 | 64.772 | 40.882 | 0.26712 | 0.11605 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR73
EN661 | 20210205 | L11 | 10 | 8 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 71.01 | 46.35 | 0.778 | 70.232 | 45.572 | 0.27573 | 0.1514 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR74
EN661 | 20210205 | L11 | 10 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 112.7 | 70.07 | 0.70817 | 111.9918 | 69.3618 | 0.90565 | 0.32954 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L11 | 10 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 114 | 74.97 | 0.70817 | 113.2918 | 74.2618 | 0.82917 | 0.49328 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L11 | 10 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.99 | 25.74 | 0.70817 | 38.2818 | 25.0318 | 0.28149 | 0.16428 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L11 | 10 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.61 | 27.24 | 0.62083 | 39.9892 | 26.6192 | 0.28404 | 0.18999 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L11 | 10 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 62.77 | 38.98 | 0.778 | 61.992 | 38.202 | 0.50023 | 0.1731 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR63
EN661 | 20210205 | L11 | 10 | 12 |  |  | b | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 57.53 | 36.08 | 0.76217 | 56.7678 | 35.3178 | 0.45257 | 0.17206 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR64
EN661 | 20210205 | L11 | 10 | 12 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 110.6 | 66.34 | 0.76217 | 109.8378 | 65.5778 | 0.49578 | 0.11997 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR71
EN661 | 20210205 | L11 | 10 | 12 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76.6 | 46.49 | 0.76217 | 75.8378 | 45.7278 | 0.33728 | 0.092087 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR72
EN661 | 20210205 | L11 | 10 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 95.1 | 61.26 | 0.70817 | 94.3918 | 60.5518 | 0.71891 | 0.35939 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L11 | 10 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 103.9 | 66.02 | 0.736 | 103.164 | 65.284 | 0.80474 | 0.35783 | 20220519 | 11-79 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L11 | 10 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 53.3 | 35.12 | 0.70817 | 52.5918 | 34.4118 | 0.38622 | 0.22658 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L11 | 10 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 43.65 | 28.98 | 0.70817 | 42.9418 | 28.2718 | 0.31165 | 0.19181 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210205 | L11 | 10 | 18 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 54.61 | 34.88 | 0.778 | 53.832 | 34.102 | 0.41345 | 0.18557 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR61
EN661 | 20210205 | L11 | 10 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.52 | 40.21 | 0.778 | 64.742 | 39.432 | 0.53219 | 0.16282 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR62
EN661 | 20210205 | L11 | 10 | 18 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 99.32 | 60.63 | 0.76217 | 98.5578 | 59.8678 | 0.43182 | 0.12828 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR69
EN661 | 20210205 | L11 | 10 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 94 | 56.2 | 0.76217 | 93.2378 | 55.4378 | 0.42342 | 0.09712 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR70
EN661 | 20210206 | L3 | 12 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 173 | 108.3 | 0.70817 | 172.2918 | 107.5918 | 1.3745 | 0.54147 | 20220519 | 11-58 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L3 | 12 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 195.1 | 120 | 0.70817 | 194.3918 | 119.2918 | 1.5954 | 0.52888 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L3 | 12 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 180.9 | 114.3 | 0.70817 | 180.1918 | 113.5918 | 1.4149 | 0.60796 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L3 | 12 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 193 | 123.2 | 0.736 | 192.264 | 122.464 | 1.4829 | 0.69797 | 20220519 | 11-79 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L3 | 12 | 3 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 182.8 | 110.8 | 0.778 | 182.022 | 110.022 | 1.5088 | 0.42382 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR81
EN661 | 20210206 | L3 | 12 | 3 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 145.4 | 89.04 | 0.778 | 144.622 | 88.262 | 1.1851 | 0.37059 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR82
EN661 | 20210206 | L3 | 12 | 3 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 216.9 | 126.5 | 0.74067 | 216.1593 | 125.7593 | 1.0108 | 0.1679 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR87
EN661 | 20210206 | L3 | 12 | 3 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 279.2 | 161.5 | 0.74067 | 278.4593 | 160.7593 | 1.316 | 0.1907 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR88
EN661 | 20210206 | L3 | 12 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 211.8 | 131.7 | 0.70817 | 211.0918 | 130.9918 | 1.7017 | 0.63101 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L3 | 12 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 191.8 | 119 | 0.70817 | 191.0918 | 118.2918 | 1.5466 | 0.55994 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L3 | 12 | 6 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 170.7 | 105.3 | 0.74067 | 169.9593 | 104.5593 | 1.3799 | 0.46936 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR79
EN661 | 20210206 | L3 | 12 | 6 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 137.9 | 84.22 | 0.778 | 137.122 | 83.442 | 1.1287 | 0.34199 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR80
EN661 | 20210206 | L3 | 12 | 6 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 221.9 | 130.2 | 0.778 | 221.122 | 129.422 | 1.0272 | 0.18804 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR85
EN661 | 20210206 | L3 | 12 | 6 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 261.7 | 153.6 | 0.74067 | 260.9593 | 152.8593 | 1.2109 | 0.2244 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR86
EN661 | 20210206 | L3 | 12 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 247.4 | 151.2 | 0.4865 | 246.9135 | 150.7135 | 2.0437 | 0.64018 | 20220519 | 11-65 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L3 | 12 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 229.5 | 139.3 | 0.70817 | 228.7918 | 138.5918 | 1.9162 | 0.55179 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L3 | 12 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 244.5 | 150.6 | 0.70817 | 243.7918 | 149.8918 | 1.9948 | 0.67441 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L3 | 12 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 234.2 | 144.2 | 0.4865 | 233.7135 | 143.7135 | 1.912 | 0.64724 | 20220519 | 11-65 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L3 | 12 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.27 | 52.2 | 0.70817 | 79.5618 | 51.4918 | 0.59633 | 0.32063 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | very poor replication (112% range); was nitex placed properly?
EN661 | 20210206 | L3 | 12 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.28 | 23.4 | 0.62083 | 30.6592 | 22.7792 | 0.16741 | 0.23824 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | very poor replication (112% range); was nitex placed properly?
EN661 | 20210206 | L3 | 12 | 13 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 236.7 | 143.2 | 0.74067 | 235.9593 | 142.4593 | 1.9593 | 0.54306 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR77
EN661 | 20210206 | L3 | 12 | 13 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 200.8 | 121.9 | 0.74067 | 200.0593 | 121.1593 | 1.659 | 0.47648 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR78
EN661 | 20210206 | L3 | 12 | 13 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 243.6 | 144.9 | 0.778 | 242.822 | 144.122 | 1.1016 | 0.24676 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR83
EN661 | 20210206 | L3 | 12 | 13 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 242.5 | 143.6 | 0.778 | 241.722 | 142.822 | 1.1078 | 0.23321 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR84
EN661 | 20210206 | L1.2 | 13 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 170.4 | 103.5 | 0.70817 | 169.6918 | 102.7918 | 4.2637 | 1.2278 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L1.2 | 13 | 1 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 253.2 | 157.4 | 0.736 | 252.464 | 156.664 | 4.0704 | 1.5093 | 20220519 | 11-79 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L1.2 | 13 | 3 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 235.5 | 143.3 | 0.736 | 234.764 | 142.564 | 3.9175 | 1.1601 | 20220519 | 11-79 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L1.2 | 13 | 3 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 240.7 | 148.3 | 0.736 | 239.964 | 147.564 | 3.926 | 1.3296 | 20220519 | 11-79 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L1.2 | 13 | 6 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 250.2 | 153.4 | 0.736 | 249.464 | 152.664 | 4.1129 | 1.3243 | 20220519 | 11-79 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L1.2 | 13 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 155.3 | 96.38 | 0.70817 | 154.5918 | 95.6718 | 3.7551 | 1.356 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L1.2 | 13 | 12 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 287 | 177.5 | 0.7505 | 286.2495 | 176.7495 | 4.6525 | 1.6426 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L1.2 | 13 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 176.4 | 106.9 | 0.70817 | 175.6918 | 106.1918 | 4.4294 | 1.2437 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210206 | L1.2 | 13 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 197.3 | 122.9 | 0.70817 | 196.5918 | 122.1918 | 1.5806 | 0.5954 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (66%) was nitex placed properly?
EN661 | 20210206 | L1.2 | 13 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 108.9 | 71.71 | 0.70817 | 108.1918 | 71.0018 | 0.79008 | 0.47431 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (66%) was nitex placed properly?
EN661 | 20210206 | L1.2 | 13 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 182 | 108.8 | 0.74067 | 181.2593 | 108.0593 | 4.6175 | 1.0963 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR91; larger than both WSW values… | No WSW or <10 in log sheet -- DF is confused by this comment, seems like there are WSW and <20 values recorded above?
EN661 | 20210206 | L1.2 | 13 | 6 |  |  | b | 292 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 188 | 112.9 | 0.74067 | 187.2593 | 112.1593 | 4.7536 | 1.1974 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR92; larger than both WSW values…
EN661 | 20210206 | L1.2 | 13 | 6 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 261.2 | 150.7 | 0.778 | 260.422 | 149.922 | 3.7133 | 0.50981 | 20220519 | 11-47 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR95
EN661 | 20210206 | L1.2 | 13 | 6 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 231.1 | 133.4 | 0.778 | 230.322 | 132.622 | 3.2832 | 0.45263 | 20220519 | 11-48 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR96
EN661 | 20210206 | L1.2 | 13 | 12 |  |  | a | 294 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 151.8 | 92.14 | 0.74067 | 151.0593 | 91.3993 | 3.7506 | 1.0659 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR89
EN661 | 20210206 | L1.2 | 13 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 150.1 | 89.22 | 0.74067 | 149.3593 | 88.4793 | 3.8403 | 0.83815 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR90
EN661 | 20210206 | L1.2 | 13 | 12 |  |  | a | 552 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 162.1 | 94.64 | 0.74067 | 161.3593 | 93.8993 | 2.2588 | 0.37669 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR93
EN661 | 20210206 | L1.2 | 13 | 12 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 142.4 | 85.08 | 0.74067 | 141.6593 | 84.3393 | 1.9262 | 0.44952 | 20220519 | 11-51 | 20220913 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR94
EN661 | 20210207 | MVCO | 14 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 139.2 | 85.45 | 0.70817 | 138.4918 | 84.7418 | 3.4256 | 1.1016 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210207 | MVCO | 14 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 229 | 139.5 | 0.70817 | 228.2918 | 138.7918 | 3.8027 | 1.1404 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210207 | MVCO | 14 | 4 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 217.3 | 135.7 | 0.70817 | 216.5918 | 134.9918 | 3.4671 | 1.3408 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (47%)
EN661 | 20210207 | MVCO | 14 | 4 |  |  | b | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 91.42 | 58.01 | 0.736 | 90.684 | 57.274 | 2.1293 | 0.93046 | 20220519 | 11-79 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (47%)
EN661 | 20210207 | MVCO | 14 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.5 | Cal_051922a | 1.8382 | 0.56187 | 206 | 125.1 | 0.70817 | 205.2918 | 124.3918 | 3.4373 | 0.99297 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210207 | MVCO | 14 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.5 | Cal_051922a | 1.8382 | 0.56187 | 208.7 | 127.7 | 0.70817 | 207.9918 | 126.9918 | 3.4416 | 1.0813 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN661 | 20210207 | MVCO | 14 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 177.8 | 113 | 0.81783 | 176.9822 | 112.1822 | 1.3766 | 0.62109 | 20220519 | 13-00 | 20230215 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN661 | 20210207 | MVCO | 14 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 135.8 | 87.99 | 0.70817 | 135.0918 | 87.2818 | 1.0157 | 0.53861 | 20220519 | 11-59 | 20221027 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | labNB has cruise as AR38C14N7 10b but EN661 analyzed same day and process of elimination shows this is cruise EN661 was was written incorrectly in labNB. Put as flag 2 because not great replication and it is an assumption it is this sample. <10s not as good a replication thought
EN661 | 20210207 | MVCO | 14 | 7 |  |  | a | 216 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 313.6 | 188 | 0.778 | 312.822 | 187.222 | 3.5824 | 0.89381 | 20220519 | 11-48 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR97; somewhat poor replication (22%) but this whole cruise is pretty variable so DF thinks within bio variation?
EN661 | 20210207 | MVCO | 14 | 7 |  |  | b | 230 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 281 | 174.1 | 0.778 | 280.222 | 173.322 | 2.8635 | 1.0282 | 20220519 | 11-48 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR98; reps off?; somewhat poor replication (22%) but this whole cruise is pretty variable so DF thinks within bio variation?
EN661 | 20210207 | MVCO | 14 | 7 |  |  | a | 552 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 162.5 | 96.54 | 0.778 | 161.722 | 95.762 | 2.2085 | 0.47919 | 20220519 | 11-48 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR99
EN661 | 20210207 | MVCO | 14 | 7 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 146.3 | 85.61 | 0.778 | 145.522 | 84.832 | 2.0395 | 0.35015 | 20220519 | 11-48 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR100
AR52A | 20210330 | L10 | 4 | 8 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 101.2 | 72.74 | 0.559 | 100.641 | 72.181 | 0.16588 | 0.18678 | 20220519 | 13-15 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/01; not counted as OOI sample but still had ID; OOI OffshoreSurf work | phae>chl; depth = 60m
AR52A | 20210330 | L10 | 4 | 9 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 178.5 | 112.8 | 0.559 | 177.941 | 112.241 | 2.2998 | 0.99361 | 20220519 | 13-15 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/02; not counted as OOI sample but still had ID; OOI OffshoreSurf work | sample was diluted; 2 acetone to 1 sample
AR52A | 20210330 | L10 | 4 | 9 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 212.7 | 132.6 | 0.57033 | 212.1297 | 132.0297 | 2.8039 | 1.0702 | 20220519 | 13-12 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/03; not counted as OOI sample but still had ID; OOI OffshoreSurf work | sample was diluted; 2 acetone to 1 sample
AR52A | 20210330 | L10 | 4 | 10 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 270.2 | 165.2 | 0.559 | 269.641 | 164.641 | 3.6755 | 1.1555 | 20220519 | 13-15 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/04 | sample was diluted; 2 acetone to 1 sample
AR52A | 20210330 | L10 | 4 | 11 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 280.5 | 171 | 0.559 | 279.941 | 170.441 | 3.8294 | 1.167 | 20220519 | 13-15 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/05 | sample was diluted; 2 acetone to 1 sample
AR52A | 20210330 | L10 | 4 | 12 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 221.6 | 135.4 | 0.559 | 221.041 | 134.841 | 3.0174 | 0.93915 | 20220519 | 13-15 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/06; not counted as OOI sample but still had ID; OOI OffshoreSurf work | sample was diluted; 2 acetone to 1 sample
AR52A | 20210330 |  | 5 | 4 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 178.7 | 116.9 | 0.60733 | 178.0927 | 116.2927 | 2.1613 | 1.2478 | 20220519 | 13-14 | 20230310 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/01; OOI Glider339 recovery | sample was diluted; 2 acetone to 1 sample
AR52A | 20210330 |  | 5 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 276.6 | 171.4 | 0.60733 | 275.9927 | 170.7927 | 3.6825 | 1.329 | 20220519 | 13-14 | 20230310 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/02; OOI Glider339 recovery | sample was diluted; 2 acetone to 1 sample
AR52A | 20210330 |  | 5 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 233.8 | 143.6 | 0.57033 | 233.2297 | 143.0297 | 3.1574 | 1.0394 | 20220519 | 13-12 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/03; OOI Glider339 recovery | sample was diluted; 2 acetone to 1 sample
AR52A | 20210331 | L8 | 6 | 4 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 219.4 | 156.5 | 0.60733 | 218.7927 | 155.8927 | 0.36662 | 0.39504 | 20220519 | 13-14 | 20230310 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/01; OOI Central Moor work | phae>chl; depth = 50m
AR52A | 20210331 | L8 | 6 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 192.5 | 119.7 | 0.559 | 191.941 | 119.141 | 2.5484 | 0.94754 | 20220519 | 13-15 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/02; OOI Central Moor work | sample was diluted; 2 acetone to 1 sample
AR52A | 20210331 | L8 | 6 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 245 | 150.2 | 0.559 | 244.441 | 149.641 | 3.3185 | 1.0724 | 20220519 | 13-15 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/03; OOI Central Moor work | sample was diluted; 2 acetone to 1 sample
AR52A | 20210331 | L8 | 6 | 9 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 278.9 | 169.5 | 0.57033 | 278.3297 | 168.9297 | 3.8295 | 1.1273 | 20220519 | 13-12 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/04; OOI Central Moor work | sample was diluted; 2 acetone to 1 sample
AR52A | 20210331 | L8 | 6 | 10 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 317.7 | 194.1 | 0.60733 | 317.0927 | 193.4927 | 4.3225 | 1.3497 | 20220519 | 13-14 | 20230310 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/05; not counted as OOI sample but still had ID; OOI Central Moor work | sample was diluted; 2 acetone to 1 sample
AR52A | 20210331 | L6 | 7 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 295.2 | 178.5 | 0.559 | 294.641 | 177.941 | 4.0851 | 1.1362 | 20220519 | 13-15 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/01; Inshore Moor work | sample was diluted; 2 acetone to 1 sample
AR52A | 20210331 | L6 | 7 | 9 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 292.6 | 178.4 | 0.559 | 292.041 | 177.841 | 3.9976 | 1.2207 | 20220519 | 13-15 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/02; Inshore Moor work | sample was diluted; 2 acetone to 1 sample
AR52A | 20210331 | L6 | 7 | 10 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 237.1 | 147.3 | 0.559 | 236.541 | 146.741 | 3.1434 | 1.1623 | 20220519 | 13-15 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/03; not counted as OOI sample but still had ID; OOI Inshore Moor work | sample was diluted; 2 acetone to 1 sample
AR52A | 20210402 |  | 17 | 3 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 275.2 | 183.5 | 0.559 | 274.641 | 182.941 | 1.069 | 0.71865 | 20220519 | 13-15 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 17/01; Glider339 deploy
AR52A | 20210402 |  | 17 | 4 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.25 | Cal_051922a | 1.8382 | 0.56187 | 265 | 162.3 | 0.559 | 264.441 | 161.741 | 4.7933 | 1.5345 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 17/02; Glider339 deploy | sample was diluted; 3 acetone to 1 sample
AR52A | 20210402 |  | 17 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.25 | Cal_051922a | 1.8382 | 0.56187 | 338.4 | 205.1 | 0.57033 | 337.8297 | 204.5297 | 6.2215 | 1.7804 | 20220519 | 13-12 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 17/03; Glider339 deploy | sample was diluted; 3 acetone to 1 sample
AR52A | 20210402 | L6 | 18 | 5 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 356.1 | 228.9 | 0.57033 | 355.5297 | 228.3297 | 1.4828 | 0.74834 | 20220519 | 13-13 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 18/01; Inshore Moor work
AR52A | 20210402 | L6 | 18 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 350.4 | 222.6 | 0.57033 | 349.8297 | 222.0297 | 1.4912 | 0.68043 | 20220519 | 13-13 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 18/02; Inshore Moor work
AR52A | 20210402 | L6 | 18 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 314.3 | 192.9 | 0.559 | 313.741 | 192.341 | 4.2496 | 1.3942 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 18/03; Inshore Moor work | sample was diluted; 2 acetone to 1 sample
AR52A | 20210402 | L6 | 18 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.25 | Cal_051922a | 1.8382 | 0.56187 | 258.6 | 157.6 | 0.559 | 258.041 | 157.041 | 4.714 | 1.43 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 18/04; Inshore Moor work | sample was diluted; 3 acetone to 1 sample
AR52A | 20210403 |  | 19 | 4 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 218.9 | 148.3 | 0.559 | 218.341 | 147.741 | 0.823 | 0.62066 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 19/01; deploy gliders 564/583
AR52A | 20210403 |  | 19 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 203.4 | 127.5 | 0.559 | 202.841 | 126.941 | 2.6569 | 1.0679 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 19/02; deploy gliders 564/583
AR52A | 20210403 |  | 19 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 277.6 | 173.7 | 0.57033 | 277.0297 | 173.1297 | 3.637 | 1.4431 | 20220519 | 13-13 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 19/03; deploy gliders 564/583 | sample was diluted; 2 acetone to 1 sample
AR52B | 20210409 | L10 | 1 | 7 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 246.3 | 150.9 | 0.57033 | 245.7297 | 150.3297 | 3.3363 | 1.0706 | 20220519 | 13-12 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 01/01; OffshoreProfiler | sample was diluted; 2 acetone to 1 sample
AR52B | 20210409 | L10 | 1 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 225 | 139.1 | 0.60733 | 224.3927 | 138.4927 | 3.0069 | 1.0568 | 20220519 | 13-14 | 20230310 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 01/02; OffshoreProfiler | sample was diluted; 2 acetone to 1 sample
AR52B | 20210410 | L12 | 3 | 7 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 193.3 | 119.8 | 0.60733 | 192.6927 | 119.1927 | 2.5704 | 0.92368 | 20220519 | 13-14 | 20230310 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 03/01; UpstreamOffshoreProfiler | sample was diluted; 2 acetone to 1 sample
AR52B | 20210410 | L12 | 3 | 8 |  |  | a | 1056 | 0 | 10 |  | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 354.1 | 301.3 | 0.559 | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 9 |  | OOI 03/02; UpstreamOffshoreProfiler; sample lost due to error during lab analysis | BAD replication is actually just false data. Rb of 354.1 is not Rb, it's the Ra reading from previous sample. Lost sample
AR52B | 20210411 | L9 | 4 | 5 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 101.3 | 80.05 | 0.60733 | 100.6927 | 79.4427 | 0.12386 | 0.26428 | 20220519 | 13-14 | 20230310 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/01; CentralOffshoreProfiler | phae>chl; depth = 53m
AR52B | 20210411 | L9 | 4 | 6 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 97.49 | 76.81 | 0.60733 | 96.8827 | 76.2027 | 0.12065 | 0.25201 | 20220519 | 13-14 | 20230310 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/02; CentralOffshoreProfiler | phae>chl; depth = 53m
AR52B | 20210411 | L9 | 4 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 160.9 | 101.7 | 0.57033 | 160.3297 | 101.1297 | 2.0723 | 0.89511 | 20220519 | 13-12 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/03; CentralOffshoreProfiler | sample was diluted; 2 acetone to 1 sample
AR52B | 20210411 | L9 | 4 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 171 | 108.9 | 0.559 | 170.441 | 108.341 | 2.1738 | 1.0052 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/04; CentralOffshoreProfiler | sample was diluted; 2 acetone to 1 sample
AR52B | 20210412 | L13 | 5 | 4 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 324.2 | 233.4 | 0.559 | 323.641 | 232.841 | 0.52924 | 0.60837 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/01; UpstreamInshoreProfiler | phae>chl; depth = 57m
AR52B | 20210412 | L13 | 5 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 364.6 | 241.1 | 0.559 | 364.041 | 240.541 | 1.441 | 0.91166 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/02; UpstreamInshoreProfiler
AR52B | 20210412 | L13 | 5 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 400.7 | 261.8 | 0.57033 | 400.1297 | 261.2297 | 1.6207 | 0.93432 | 20220519 | 13-12 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/03; UpstreamInshoreProfiler
AR52B | 20210412 | L13 | 5 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 177.8 | 116.4 | 0.559 | 177.241 | 115.841 | 1.0746 | 0.62488 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/04; UpstreamInshoreProfiler | sample was diluted; 2 acetone to 1 sample
AR52B | 20210412 | L13 | 5 | 8 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 187.6 | 122.5 | 0.559 | 187.041 | 121.941 | 1.1383 | 0.649 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/05; UpstreamInshoreProfiler | sample was diluted; 2 acetone to 1 sample
AR52B | 20210413 | L7 | 6 | 5 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 355.3 | 226 | 0.57033 | 354.7297 | 225.4297 | 1.5073 | 0.69552 | 20220519 | 13-12 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/01; CentralInshoreProfiler
AR52B | 20210413 | L7 | 6 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 345.8 | 221.1 | 0.559 | 345.241 | 220.541 | 1.455 | 0.70204 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/02; CentralInshoreProfiler
AR52B | 20210413 | L7 | 6 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 347.8 | 218.5 | 0.60733 | 347.1927 | 217.8927 | 1.5087 | 0.62246 | 20220519 | 13-14 | 20230310 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/03; CentralInshoreProfiler
AR52B | 20210413 | L7 | 6 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 350.2 | 216.5 | 0.60733 | 349.5927 | 215.8927 | 1.56 | 0.55156 | 20220519 | 13-14 | 20230310 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/04; CentralInshoreProfiler
AR52B | 20210413 |  | 7 | 6 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 186.2 | 115.6 | 0.57033 | 185.6297 | 115.0297 | 2.469 | 0.90306 | 20220519 | 13-12 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/01; AC-1 AUVdeploy | sample was diluted; 2 acetone to 1 sample
AR52B | 20210413 |  | 7 | 6 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 207.7 | 128.2 | 0.57033 | 207.1297 | 127.6297 | 2.7829 | 0.96209 | 20220519 | 13-12 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/02; AC-1 AUVdeploy | sample was diluted; 2 acetone to 1 sample
AR52B | 20210414 |  | 8 | 3 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 235.1 | 177.7 | 0.57033 | 234.5297 | 177.1297 | 0.33456 | 0.53086 | 20220519 | 13-12 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 08/01; AL-1 AUVdeploy | phae>chl; depth = 40m
AR52B | 20210414 |  | 8 | 4 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 334.5 | 217.8 | 0.57033 | 333.9297 | 217.2297 | 1.3617 | 0.763 | 20220519 | 13-12 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 08/02; AL-1 AUVdeploy
AR52B | 20210414 |  | 8 | 4 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 395.9 | 252.6 | 0.60733 | 395.2927 | 251.9927 | 1.6721 | 0.79263 | 20220519 | 13-14 | 20230310 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 08/03; AL-1 AUVdeploy
AR52B | 20210414 |  | 9 | 4 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 365.5 | 229.1 | 0.559 | 364.941 | 228.541 | 1.59 | 0.64316 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 09/01; AC-4 AUVrecovery
AR52B | 20210414 |  | 9 | 4 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 405.3 | 249.9 | 0.559 | 404.741 | 249.341 | 1.8133 | 0.62551 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 09/02; AC-4 AUVrecovery
AR52B | 20210414 |  | 9 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 325.4 | 200.2 | 0.60733 | 324.7927 | 199.5927 | 1.4609 | 0.49131 | 20220519 | 13-14 | 20230310 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 09/03; AC-4 AUVrecovery
AR52B | 20210415 |  | 10 | 3 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 390.4 | 267.7 | 0.559 | 389.841 | 267.141 | 0.71517 | 0.59002 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 10/01; AL-4 AUVrecovery
AR52B | 20210415 |  | 10 | 4 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 328 | 204.1 | 0.559 | 327.441 | 203.541 | 1.4457 | 0.5451 | 20220519 | 13-16 | 20230314 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 10/02; AL-4 AUVrecovery
AR52B | 20210415 |  | 10 | 4 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 329.4 | 203.8 | 0.57033 | 328.8297 | 203.2297 | 1.4655 | 0.52222 | 20220519 | 13-12 | 20230309 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 10/03; AL-4 AUVrecovery
EN668 | 20210716 | L1 | 1 | 1 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 236.4 | 149.8 | 0.7505 | 235.6495 | 149.0495 | 3.6795 | 1.629 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210716 | L1 | 1 | 1 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 235.2 | 147.9 | 0.736 | 234.464 | 147.164 | 3.7093 | 1.5321 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210716 | L1 | 1 | 4 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 235.2 | 147.1 | 0.7505 | 234.4495 | 146.3495 | 3.7433 | 1.4691 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210716 | L1 | 1 | 4 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 228.6 | 144.2 | 0.7505 | 227.8495 | 143.4495 | 3.586 | 1.523 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210716 | L1 | 1 | 4 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 345.8 | 217.7 | 0.62083 | 345.1792 | 217.0792 | 2.7214 | 1.1443 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210716 | L1 | 1 | 4 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 337.1 | 214.9 | 0.62083 | 336.4792 | 214.2792 | 2.5961 | 1.2198 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | 1 large black particle pn underside of filter, made it into exrtaction
EN668 | 20210716 | L1 | 1 | 4 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 238 | 140.9 | 0.83083 | 237.1692 | 140.0692 | 2.0417 | 0.42709 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR3; poor replication (31% range)
EN668 | 20210716 | L1 | 1 | 4 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 176.9 | 106.3 | 0.83083 | 176.0692 | 105.4692 | 1.4794 | 0.37318 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR4; poor replication (31% range)
EN668 | 20210716 | L1 | 1 | 4 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 86.62 | 51.38 | 0.76217 | 85.8578 | 50.6178 | 0.39474 | 0.080539 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR9
EN668 | 20210716 | L1 | 1 | 4 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 90.86 | 53.42 | 0.76217 | 90.0978 | 52.6578 | 0.41939 | 0.07505 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR10
EN668 | 20210716 | L1 | 1 | 11 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 165.9 | 105.7 | 0.7505 | 165.1495 | 104.9495 | 2.5578 | 1.18 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210716 | L1 | 1 | 11 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 167.4 | 108.1 | 0.7505 | 166.6495 | 107.3495 | 2.5196 | 1.3038 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210716 | L1 | 1 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 227 | 149.1 | 0.62083 | 226.3792 | 148.4792 | 1.6549 | 0.98916 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210716 | L1 | 1 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 233 | 155.2 | 0.62083 | 232.3792 | 154.5792 | 1.6528 | 1.0999 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210716 | L1 | 1 | 11 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 123.2 | 74.84 | 0.76217 | 122.4378 | 74.0778 | 1.0203 | 0.28979 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR5
EN668 | 20210716 | L1 | 1 | 11 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 140.3 | 85.99 | 0.76217 | 139.5378 | 85.2278 | 1.142 | 0.36022 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR6
EN668 | 20210716 | L1 | 1 | 11 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 75.71 | 45.29 | 0.83083 | 74.8792 | 44.4592 | 0.34013 | 0.076563 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR11
EN668 | 20210716 | L1 | 1 | 11 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 77.74 | 45.44 | 0.76217 | 76.9778 | 44.6778 | 0.36115 | 0.057592 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR12
EN668 | 20210716 | L1 | 1 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 125.8 | 77.53 | 0.736 | 125.064 | 76.794 | 1.0255 | 0.34207 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210716 | L1 | 1 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 122.7 | 77.58 | 0.7505 | 121.9495 | 76.8295 | 0.95854 | 0.40963 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210716 | L1 | 1 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 101.5 | 63.4 | 0.62083 | 100.8792 | 62.7792 | 0.80941 | 0.30855 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210716 | L1 | 1 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 105.4 | 65.44 | 0.62083 | 104.7792 | 64.8192 | 0.84892 | 0.30537 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210716 | L1 | 1 | 23 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.13 | 20.53 | 0.76217 | 33.3678 | 19.7678 | 0.28694 | 0.062668 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR1
EN668 | 20210716 | L1 | 1 | 23 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 32.12 | 19.76 | 0.83083 | 31.2892 | 18.9292 | 0.25901 | 0.073494 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR2
EN668 | 20210716 | L1 | 1 | 15 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.37 | 7.247 | 0.83083 | 10.5392 | 6.4162 | 0.046184 | 0.014061 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR7; poor replication (35% range) but low values and within bio variation
EN668 | 20210716 | L1 | 1 | 15 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.96 | 9.046 | 0.83083 | 14.1292 | 8.2152 | 0.066246 | 0.010891 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR8; poor replication (35% range) but low values and within bio variation
EN668 | 20210717 | L2 | 2 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 90.07 | 61.58 | 0.736 | 89.334 | 60.844 | 0.60525 | 0.47825 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN668 | 20210717 | L2 | 2 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 93.6 | 63.61 | 0.7505 | 92.8495 | 62.8595 | 0.63712 | 0.48228 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1
EN668 | 20210717 | L2 | 2 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 103.8 | 70.3 | 0.7505 | 103.0495 | 69.5495 | 0.71169 | 0.52684 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L2 | 2 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 106.2 | 69.98 | 0.736 | 105.464 | 69.244 | 0.76947 | 0.46362 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L2 | 2 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 177.1 | 115.4 | 0.736 | 176.364 | 114.664 | 1.3108 | 0.73115 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L2 | 2 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 185.9 | 120.1 | 0.7505 | 185.1495 | 119.3495 | 1.3979 | 0.72748 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L2 | 2 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 58.56 | 38.26 | 0.62083 | 57.9392 | 37.6392 | 0.43126 | 0.23901 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
EN668 | 20210717 | L2 | 2 | 7 |  |  | b | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN668 | 20210717 | L2 | 2 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 73.87 | 47.64 | 0.83083 | 73.0392 | 46.8092 | 0.55153 | 0.2735 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR17
EN668 | 20210717 | L2 | 2 | 7 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.91 | 44.75 | 0.76217 | 70.1478 | 43.9878 | 0.54819 | 0.22448 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR18
EN668 | 20210717 | L2 | 2 | 7 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 68.97 | 40.41 | 0.83083 | 68.1392 | 39.5792 | 0.31934 | 0.051622 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR23; poor replication (33% range) but DF thinks within bio variation
EN668 | 20210717 | L2 | 2 | 7 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 50.49 | 30.2 | 0.83083 | 49.6592 | 29.3692 | 0.22687 | 0.048397 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR24; poor replication (33% range) but DF thinks within bio variation
EN668 | 20210717 | L2 | 2 | 12 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 123.7 | 78.99 | 0.7505 | 122.9495 | 78.2395 | 1.8997 | 0.88689 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L2 | 2 | 12 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 120.4 | 79.06 | 0.7505 | 119.6495 | 78.3095 | 1.7565 | 1.0326 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L2 | 2 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 194.8 | 125.3 | 0.81783 | 193.9822 | 124.4822 | 1.4765 | 0.74028 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210717 | L2 | 2 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 177.4 | 115 | 0.62083 | 176.7792 | 114.3792 | 1.3256 | 0.7112 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210717 | L2 | 2 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 85.08 | 51.95 | 0.76217 | 84.3178 | 51.1878 | 0.69662 | 0.20559 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR15
EN668 | 20210717 | L2 | 2 | 12 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 100.3 | 62.5 | 0.83083 | 99.4692 | 61.6692 | 0.79211 | 0.29114 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR16
EN668 | 20210717 | L2 | 2 | 12 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.77 | 31.46 | 0.76217 | 52.0078 | 30.6978 | 0.23871 | 0.049535 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR21
EN668 | 20210717 | L2 | 2 | 12 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 47.29 | 28.27 | 0.83083 | 46.4592 | 27.4392 | 0.21305 | 0.044589 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR22
EN668 | 20210717 | L2 | 2 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.87 | 32.22 | 0.7505 | 52.1195 | 31.4695 | 0.4387 | 0.12171 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L2 | 2 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 54.93 | 33.46 | 0.7505 | 54.1795 | 32.7095 | 0.45612 | 0.12637 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L2 | 2 | 17 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 47.95 | 29.38 | 0.81783 | 47.1322 | 28.5622 | 0.39451 | 0.11412 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210717 | L2 | 2 | 17 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 42.78 | 25.77 | 0.62083 | 42.1592 | 25.1492 | 0.36137 | 0.086487 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210717 | L2 | 2 | 17 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.509 | 4.874 | 0.76217 | 6.7468 | 4.1118 | 0.055217 | 0.01701 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR13
EN668 | 20210717 | L2 | 2 | 17 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.843 | 5.876 | 0.83083 | 8.0122 | 5.0452 | 0.062174 | 0.026447 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR14
EN668 | 20210717 | L2 | 2 | 17 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.817 | 2.687 | 0.76217 | 3.0548 | 1.9248 | 0.012658 | 0.0054157 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR19
EN668 | 20210717 | L2 | 2 | 17 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.436 | 2.952 | 0.76217 | 3.6738 | 2.1898 | 0.016563 | 0.0039243 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR20
EN668 | 20210717 | L4 | 5 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 35.29 | 24.83 | 0.7505 | 34.5395 | 24.0795 | 0.22222 | 0.20659 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L4 | 5 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 32.74 | 22.93 | 0.736 | 32.004 | 22.194 | 0.20841 | 0.18682 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L4 | 5 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.47 | 54.07 | 0.7505 | 79.7195 | 53.3195 | 0.56085 | 0.38866 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L4 | 5 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 87.5 | 52.52 | 0.7505 | 86.7495 | 51.7695 | 0.74313 | 0.17878 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L4 | 5 | 4 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 42.69 | 27.78 | 0.83083 | 41.8592 | 26.9492 | 0.31351 | 0.16148 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR29
EN668 | 20210717 | L4 | 5 | 4 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 39.74 | 26.17 | 0.83083 | 38.9092 | 25.3392 | 0.28436 | 0.16073 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR30
EN668 | 20210717 | L4 | 5 | 4 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.73 | 14.04 | 0.76217 | 20.9678 | 13.2778 | 0.085984 | 0.038464 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR37
EN668 | 20210717 | L4 | 5 | 4 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 27.11 | 16.94 | 0.76217 | 26.3478 | 16.1778 | 0.11371 | 0.037915 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR38
EN668 | 20210717 | L4 | 5 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 196.1 | 125.9 | 0.7505 | 195.3495 | 125.1495 | 1.4914 | 0.73729 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L4 | 5 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 183 | 119.3 | 0.7505 | 182.2495 | 118.5495 | 1.3533 | 0.75785 | 20220519 | 11-75 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L4 | 5 | 9 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 209.7 | 133.8 | 0.7505 | 208.9495 | 133.0495 | 3.2249 | 1.5138 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L4 | 5 | 9 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 204.7 | 132.1 | 0.7505 | 203.9495 | 131.3495 | 3.0847 | 1.5934 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L4 | 5 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 267.9 | 170.3 | 0.81783 | 267.0822 | 169.4822 | 2.0734 | 0.94467 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210717 | L4 | 5 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 251.9 | 163.6 | 0.62083 | 251.2792 | 162.9792 | 1.8759 | 1.0264 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | C4N9<10b noticeably darker than 10a is nitex mesh not centered and sealed on oring?
EN668 | 20210717 | L4 | 5 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 154.2 | 98.25 | 0.83083 | 153.3692 | 97.4192 | 1.1765 | 0.54061 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR27
EN668 | 20210717 | L4 | 5 | 9 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 162.1 | 104.1 | 0.83083 | 161.2692 | 103.2692 | 1.2154 | 0.59858 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR28
EN668 | 20210717 | L4 | 5 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 178.6 | 108.6 | 0.76217 | 177.8378 | 107.8378 | 0.78411 | 0.22845 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR35
EN668 | 20210717 | L4 | 5 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 176.8 | 106.3 | 0.76217 | 176.0378 | 105.5378 | 0.78971 | 0.20125 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR36
EN668 | 20210717 | L4 | 5 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 60.51 | 36.79 | 0.7505 | 59.7595 | 36.0395 | 0.50392 | 0.13787 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L4 | 5 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 75.88 | 51.3 | 0.7505 | 75.1295 | 50.5495 | 0.52219 | 0.37799 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L4 | 5 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.06 | 23.86 | 0.62083 | 37.4392 | 23.2392 | 0.30167 | 0.11217 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L4 | 5 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.43 | 24.27 | 0.62083 | 37.8092 | 23.6492 | 0.30082 | 0.12032 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L4 | 5 | 14 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 24.68 | 14.63 | 0.76217 | 23.9178 | 13.8678 | 0.2106 | 0.032996 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR31
EN668 | 20210717 | L4 | 5 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23.09 | 13.71 | 0.76217 | 22.3278 | 12.9478 | 0.19656 | 0.030876 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR32
EN668 | 20210717 | L4 | 5 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 35.56 | 20.75 | 0.76217 | 34.7978 | 19.9878 | 0.1659 | 0.021783 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR39
EN668 | 20210717 | L4 | 5 | 14 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.4 | 16.81 | 0.83083 | 27.5692 | 15.9792 | 0.12936 | 0.020139 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR40
EN668 | 20210717 | L4 | 5 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 43.99 | 27.67 | 0.7505 | 43.2395 | 26.9195 | 0.34671 | 0.13267 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L4 | 5 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.53 | 28.06 | 0.7505 | 44.7795 | 27.3095 | 0.37114 | 0.11519 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L4 | 5 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.83 | 21.42 | 0.62083 | 34.2092 | 20.7992 | 0.28489 | 0.085502 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210717 | L4 | 5 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.23 | 23.4 | 0.81783 | 37.4122 | 22.5822 | 0.31505 | 0.087087 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Several more than 2 drops of acid used for Ra. Doesn't appear to have diluted sample though
EN668 | 20210717 | L4 | 5 | 18 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.813 | 4.033 | 0.83083 | 4.9822 | 3.2022 | 0.0373 | 0.018948 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR25
EN668 | 20210717 | L4 | 5 | 18 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.344 | 2.89 | 0.83083 | 3.5132 | 2.0592 | 0.030469 | 0.0057015 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR26
EN668 | 20210717 | L4 | 5 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.601 | 3.157 | 0.83083 | 3.7702 | 2.3262 | 0.016175 | 0.0056667 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR33
EN668 | 20210717 | L4 | 5 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.462 | 2.674 | 0.83083 | 3.6312 | 1.8432 | 0.019956 | -0.0027119 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR34
EN668 | 20210717 | L5 | 7 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 46.2 | 33.18 | 0.7505 | 45.4495 | 32.4295 | 0.2766 | 0.3009 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L5 | 7 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 39.45 | 28.52 | 0.7505 | 38.6995 | 27.7695 | 0.2322 | 0.26232 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L5 | 7 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.41 | 31.66 | 0.7505 | 44.6595 | 30.9095 | 0.29211 | 0.25832 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L5 | 7 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.94 | 32.44 | 0.7505 | 45.1895 | 31.6895 | 0.2868 | 0.27752 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L5 | 7 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 137.5 | 86.87 | 0.7505 | 136.7495 | 86.1195 | 1.0756 | 0.458 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L5 | 7 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 128.5 | 81.3 | 0.7505 | 127.7495 | 80.5495 | 1.0027 | 0.43168 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L5 | 7 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 334.2 | 206.2 | 0.7505 | 333.4495 | 205.4495 | 2.7193 | 0.93934 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L5 | 7 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 360.9 | 225.3 | 0.736 | 360.164 | 224.564 | 2.8807 | 1.1183 | 20220519 | 11-78 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L5 | 7 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 133.1 | 77.88 | 0.76217 | 132.3378 | 77.1178 | 1.1611 | 0.19814 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR41; poor replication (31% range) but DF thinks just within natural bio variation
EN668 | 20210717 | L5 | 7 | 9 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 176.4 | 100.5 | 0.76217 | 175.6378 | 99.7378 | 1.5905 | 0.16145 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR42; poor replication (31% range) but DF thinks just within natural bio variation
EN668 | 20210717 | L5 | 7 | 9 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 174.8 | 102.3 | 0.76217 | 174.0378 | 101.5378 | 0.81064 | 0.14103 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR47; somewhat poor replication (28% range) but within natural bio variation
EN668 | 20210717 | L5 | 7 | 9 |  |  | b | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 78.28 | 46.19 | 0.76217 | 77.5178 | 45.4278 | 1.0764 | 0.20091 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR48; somewhat poor replication (28% range) but within natural bio variation
EN668 | 20210717 | L5 | 7 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 47.57 | 30.04 | 0.7505 | 46.8195 | 29.2895 | 0.37241 | 0.14917 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L5 | 7 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.78 | 31.56 | 0.7505 | 49.0295 | 30.8095 | 0.38707 | 0.16158 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L5 | 7 | 13 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.37 | 7.365 | 0.76217 | 10.6078 | 6.6028 | 0.084212 | 0.032166 | 20220519 | 11-49 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR43
EN668 | 20210717 | L5 | 7 | 13 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.53 | 8.431 | 0.83083 | 12.6992 | 7.6002 | 0.10685 | 0.02665 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR44
EN668 | 20210717 | L5 | 7 | 13 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.994 | 5.865 | 0.76217 | 8.2318 | 5.1028 | 0.034986 | 0.012841 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR49
EN668 | 20210717 | L5 | 7 | 13 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 9.273 | 5.707 | 0.83083 | 8.4422 | 4.8762 | 0.039872 | 0.0058299 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR50
EN668 | 20210717 | L5 | 7 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.56 | 10.74 | 0.7505 | 15.8095 | 9.9895 | 0.12364 | 0.05425 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L5 | 7 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.33 | 10.75 | 0.7505 | 15.5795 | 9.9995 | 0.11854 | 0.059526 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210717 | L5 | 7 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.39 | 9.309 | 0.62083 | 13.7692 | 8.6882 | 0.10794 | 0.046775 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210717 | L5 | 7 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.02 | 10.42 | 0.81783 | 15.2022 | 9.6022 | 0.11897 | 0.052026 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210717 | L5 | 7 | 16 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.832 | 2.784 | 0.83083 | 3.0012 | 1.9532 | 0.021961 | 0.012347 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR45
EN668 | 20210717 | L5 | 7 | 16 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.879 | 2.955 | 0.83083 | 3.0482 | 2.1242 | 0.019363 | 0.01795 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR46
EN668 | 20210717 | L5 | 7 | 16 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.603 | 2.087 | 0.83083 | 1.7722 | 1.2562 | 0.00578 | 0.0060149 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR51; very poor replication (111 % range) but values are so low, basically 0 within range of flourometer
EN668 | 20210717 | L5 | 7 | 16 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.524 | 2.705 | 0.83083 | 3.6932 | 1.8742 | 0.020302 | -0.0027679 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR52; very poor replication (111 % range) but values are so low, basically 0 within range of flourometer
EN668 | 20210718 | L10 | 8 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.16 | 11.02 | 0.7505 | 14.4095 | 10.2695 | 0.087952 | 0.094926 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.82 | 10.78 | 0.7505 | 14.0695 | 10.0295 | 0.085827 | 0.092777 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.05 | 23.37 | 0.59517 | 33.4548 | 22.7748 | 0.22689 | 0.17868 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.31 | 22.2 | 0.7505 | 30.5595 | 21.4495 | 0.19354 | 0.18843 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 24.78 | 16.96 | 0.62083 | 24.1592 | 16.3392 | 0.16613 | 0.12484 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 26.56 | 17.73 | 0.62083 | 25.9392 | 17.1092 | 0.18759 | 0.11709 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.103 | 3.968 | 0.76217 | 4.3408 | 3.2058 | 0.023865 | 0.032639 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR57
EN668 | 20210718 | L10 | 8 | 8 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.739 | 3.689 | 0.76217 | 3.9768 | 2.9268 | 0.022003 | 0.029408 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR58
EN668 | 20210718 | L10 | 8 | 8 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.273 | 1.935 | 0.76217 | 1.5108 | 1.1728 | 0.0037793 | 0.0072132 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR63
EN668 | 20210718 | L10 | 8 | 8 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.171 | 1.849 | 0.76217 | 1.4088 | 1.0868 | 0.0036004 | 0.0065861 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR64
EN668 | 20210718 | L10 | 8 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 62.41 | 43.95 | 0.7505 | 61.6595 | 43.1995 | 0.39217 | 0.37712 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.33 | 43.47 | 0.7505 | 62.5795 | 42.7195 | 0.42191 | 0.33883 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 13 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 118.2 | 79.64 | 0.7505 | 117.4495 | 78.8895 | 1.6384 | 1.1713 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 13 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 116.7 | 79.91 | 0.7505 | 115.9495 | 79.1595 | 1.5632 | 1.2562 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 88.53 | 59.66 | 0.7505 | 87.7795 | 58.9095 | 0.61332 | 0.43573 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 79.8 | 54.59 | 0.7505 | 79.0495 | 53.8395 | 0.53557 | 0.4232 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | C8N15chlb dropped face down
EN668 | 20210718 | L10 | 8 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 53.61 | 37.08 | 0.62083 | 52.9892 | 36.4592 | 0.35117 | 0.29809 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | a few visible particles on filter, shouldn't be for <10… didn't pick off
EN668 | 20210718 | L10 | 8 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 50.89 | 34.77 | 0.62083 | 50.2692 | 34.1492 | 0.34246 | 0.26567 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 42.64 | 27.58 | 0.83083 | 41.8092 | 26.7492 | 0.31666 | 0.1548 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR55
EN668 | 20210718 | L10 | 8 | 15 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 43.74 | 27.59 | 0.76217 | 42.9778 | 26.8278 | 0.33843 | 0.13282 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR56
EN668 | 20210718 | L10 | 8 | 15 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 26.97 | 16.47 | 0.76217 | 26.2078 | 15.7078 | 0.11762 | 0.029874 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR61; poor replication (51%) but DF thinks within bio variation
EN668 | 20210718 | L10 | 8 | 15 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 43.15 | 25.43 | 0.76217 | 42.3878 | 24.6678 | 0.19849 | 0.03313 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR62; poor replication (51%) but DF thinks within bio variation
EN668 | 20210718 | L10 | 8 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 75.21 | 55.1 | 0.7505 | 74.4595 | 54.3495 | 0.42722 | 0.54063 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 72.36 | 52.46 | 0.7505 | 71.6095 | 51.7095 | 0.42276 | 0.49807 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 21 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.23 | 17.56 | 0.81783 | 27.4122 | 16.7422 | 0.22668 | 0.071465 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210718 | L10 | 8 | 21 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.39 | 18.23 | 0.7505 | 27.6395 | 17.4795 | 0.21584 | 0.09543 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210718 | L10 | 8 | 21 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.97 | 14.32 | 0.62083 | 21.3492 | 13.6992 | 0.16252 | 0.081434 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L10 | 8 | 21 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22.65 | 14.63 | 0.62083 | 22.0292 | 14.0092 | 0.17038 | 0.079094 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | 1 large particle visible -- picked off
EN668 | 20210718 | L10 | 8 | 21 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 9.56 | 6.471 | 0.83083 | 8.7292 | 5.6402 | 0.064731 | 0.034342 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR53
EN668 | 20210718 | L10 | 8 | 21 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.149 | 4.693 | 0.83083 | 6.3182 | 3.8622 | 0.051466 | 0.016375 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR54
EN668 | 20210718 | L10 | 8 | 21 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.32 | 2.369 | 0.76217 | 2.5578 | 1.6068 | 0.010614 | 0.0044188 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR59; poor replication (56% range) but values so low, bascially 0 and within noise
EN668 | 20210718 | L10 | 8 | 21 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.719 | 2.188 | 0.76217 | 1.9568 | 1.4258 | 0.0059265 | 0.007413 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR60; poor replication (56% range) but values so low, bascially 0 and within noise
EN668 | 20210718 | L11 | 11 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.44 | 13.86 | 0.7505 | 18.6895 | 13.1095 | 0.11854 | 0.11491 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L11 | 11 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.72 | 14.24 | 0.7505 | 18.9695 | 13.4895 | 0.11642 | 0.1238 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L11 | 11 | 3 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.43 | 4.429 | 0.76217 | 4.6678 | 3.6668 | 0.021048 | 0.043582 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR71; poor replication (57% range)
EN668 | 20210718 | L11 | 11 | 3 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.356 | 4.803 | 0.76217 | 4.5938 | 4.0408 | 0.011588 | 0.059391 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR72; poor replication (57% range)
EN668 | 20210718 | L11 | 11 | 3 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.538 | 2.648 | 0.83083 | 2.7072 | 1.8172 | 0.0099513 | 0.0070802 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR79
EN668 | 20210718 | L11 | 11 | 3 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.419 | 2.539 | 0.83083 | 2.5882 | 1.7082 | 0.0098395 | 0.0061704 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR80
EN668 | 20210718 | L11 | 11 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 153.8 | 97.56 | 0.7505 | 153.0495 | 96.8095 | 1.1948 | 0.52919 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L11 | 11 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 149.7 | 96.2 | 0.7505 | 148.9495 | 95.4495 | 1.1366 | 0.56318 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L11 | 11 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 98.86 | 64.56 | 0.62083 | 98.2392 | 63.9392 | 0.72868 | 0.40994 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L11 | 11 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 97.13 | 63.15 | 0.62083 | 96.5092 | 62.5292 | 0.72188 | 0.39163 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L11 | 11 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.77 | 40.08 | 0.76217 | 64.0078 | 39.3178 | 0.51915 | 0.17385 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR69
EN668 | 20210718 | L11 | 11 | 7 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 55.92 | 33.86 | 0.76217 | 55.1578 | 33.0978 | 0.46227 | 0.11911 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR70
EN668 | 20210718 | L11 | 11 | 7 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 62.98 | 36.04 | 0.76217 | 62.2178 | 35.2778 | 0.30177 | 0.029475 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR77
EN668 | 20210718 | L11 | 11 | 7 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.5 | 39.72 | 0.83083 | 65.6692 | 38.8892 | 0.29998 | 0.065176 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR78
EN668 | 20210718 | L11 | 11 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 212.7 | 156.5 | 0.59517 | 212.1048 | 155.9048 | 1.1939 | 1.5824 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L11 | 11 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 223.9 | 161.9 | 0.59517 | 223.3048 | 161.3048 | 1.3171 | 1.5553 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L11 | 11 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.19 | 26.67 | 0.59517 | 39.5948 | 26.0748 | 0.28722 | 0.17711 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L11 | 11 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 37.38 | 25.13 | 0.7505 | 36.6295 | 24.3795 | 0.26024 | 0.1739 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L11 | 11 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.1 | 22.28 | 0.62083 | 32.4792 | 21.6592 | 0.22986 | 0.15584 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210718 | L11 | 11 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.05 | 22.68 | 0.864 | 33.186 | 21.816 | 0.24155 | 0.14695 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210718 | L11 | 11 | 13 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.49 | 11.2 | 0.83083 | 14.6592 | 10.3692 | 0.089898 | 0.092242 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR67
EN668 | 20210718 | L11 | 11 | 13 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.47 | 9.967 | 0.76217 | 13.7078 | 9.2048 | 0.094362 | 0.067327 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR68
EN668 | 20210718 | L11 | 11 | 13 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.039 | 3.376 | 0.76217 | 3.2768 | 2.6138 | 0.0073997 | 0.017054 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR75
EN668 | 20210718 | L11 | 11 | 13 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.089 | 1.846 | 0.76217 | 1.3268 | 1.0838 | 0.002722 | 0.0074548 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR76
EN668 | 20210718 | L11 | 11 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.07 | 10.72 | 0.7505 | 16.3195 | 9.9695 | 0.1349 | 0.042634 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L11 | 11 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.34 | 11.08 | 0.7505 | 16.5895 | 10.3295 | 0.13299 | 0.050957 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210718 | L11 | 11 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.56 | 9.145 | 0.81783 | 13.7422 | 8.3272 | 0.11504 | 0.033251 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210718 | L11 | 11 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.15 | 9.126 | 0.81783 | 13.3322 | 8.3082 | 0.10673 | 0.041219 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210718 | L11 | 11 | 18 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.409 | 3.273 | 0.83083 | 3.5782 | 2.4422 | 0.023805 | 0.019093 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR65; lost a couple ml
EN668 | 20210718 | L11 | 11 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.533 | 3.027 | 0.83083 | 3.7022 | 2.1962 | 0.031666 | 0.0070422 | 20220519 | 11-54 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR66
EN668 | 20210718 | L11 | 11 | 18 |  |  | a | 500 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 1.964 | 1.415 | 0.76217 | 1.2018 | 0.65283 | 0.0067646 | -2.1789e-05 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR73; lost volume, re-measured with grad cylinder
EN668 | 20210718 | L11 | 11 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 1.712 | 1.149 | 0.76217 | 0.94983 | 0.38683 | 0.0063065 | -0.0026743 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR74
EN668 | 20210719 | L9 | 14 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.44 | 12.79 | 0.7505 | 15.6895 | 12.0395 | 0.077542 | 0.13686 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L9 | 14 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.66 | 11.1 | 0.59517 | 14.0648 | 10.5048 | 0.07563 | 0.11144 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L9 | 14 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 84.42 | 53.94 | 0.7505 | 83.6695 | 53.1895 | 0.64753 | 0.29966 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L9 | 14 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 89.54 | 58.25 | 0.59517 | 88.9448 | 57.6548 | 0.66474 | 0.36197 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L9 | 14 | 7 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 141.5 | 86.68 | 0.7505 | 140.7495 | 85.9295 | 2.3292 | 0.73121 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L9 | 14 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 277.3 | 165.3 | 0.5195 | 276.7805 | 164.7805 | 2.3794 | 0.55502 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L9 | 14 | 7 |  |  | a | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN668 | 20210719 | L9 | 14 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 112.3 | 72 | 0.62083 | 111.6792 | 71.3792 | 0.85615 | 0.41496 | 20220519 | 11-80 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | C# illegible on tiss cap-number not completely written. Process of elimination of what hasn't been run plus relative chl values assume it is this sample. Photo of tiss cap
EN668 | 20210719 | L9 | 14 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 211.8 | 127 | 0.83083 | 210.9692 | 126.1692 | 1.7831 | 0.44073 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR85
EN668 | 20210719 | L9 | 14 | 7 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 185.2 | 107.9 | 0.7595 | 184.4405 | 107.1405 | 1.6198 | 0.26214 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR86
EN668 | 20210719 | L9 | 14 | 7 |  |  | a | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 74.8 | 42.19 | 0.76217 | 74.0378 | 41.4278 | 1.0939 | 0.070995 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR91
EN668 | 20210719 | L9 | 14 | 7 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 281.8 | 160.3 | 0.76217 | 281.0378 | 159.5378 | 1.3585 | 0.13676 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR92
EN668 | 20210719 | L9 | 14 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 145.9 | 94.58 | 0.59517 | 145.3048 | 93.9848 | 1.0903 | 0.58341 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L9 | 14 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 119.2 | 74.71 | 0.7505 | 118.4495 | 73.9595 | 0.94516 | 0.3719 | 20220519 | 11-76 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L9 | 14 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 324.9 | 203.9 | 0.7505 | 324.1495 | 203.1495 | 2.5706 | 1.0471 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L9 | 14 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 335.3 | 210 | 0.59517 | 334.7048 | 209.4048 | 2.6619 | 1.0671 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L9 | 14 | 14 |  |  | a | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN668 | 20210719 | L9 | 14 | 14 |  |  | b | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN668 | 20210719 | L9 | 14 | 14 |  |  | a | 291 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 287.3 | 175.8 | 0.83083 | 286.4692 | 174.9692 | 2.3606 | 0.74452 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR83
EN668 | 20210719 | L9 | 14 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 304.5 | 191.9 | 0.83083 | 303.6692 | 191.0692 | 2.3596 | 0.99668 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR84
EN668 | 20210719 | L9 | 14 | 14 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 304.1 | 181.3 | 0.7595 | 303.3405 | 180.5405 | 1.3731 | 0.31907 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR89; poor replication (35% range) but DF thinks within bio variation
EN668 | 20210719 | L9 | 14 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 144.7 | 86.69 | 0.76217 | 143.9378 | 85.9278 | 1.9494 | 0.47108 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR90; poor replication (35% range) but DF thinks within bio variation
EN668 | 20210719 | L9 | 14 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 57.83 | 37.5 | 0.864 | 56.966 | 36.636 | 0.4319 | 0.22051 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210719 | L9 | 14 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 60.02 | 38.35 | 0.59517 | 59.4248 | 37.7548 | 0.46036 | 0.21197 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210719 | L9 | 14 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.02 | 10.68 | 0.59517 | 15.4248 | 10.0848 | 0.11344 | 0.066145 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L9 | 14 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.38 | 10.19 | 0.7505 | 14.6295 | 9.4395 | 0.11026 | 0.057839 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | N20chla WSW leaked, assume that means funnel leaked, sample definitely compromised
EN668 | 20210719 | L9 | 14 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.27 | 7.849 | 0.81783 | 11.4522 | 7.0312 | 0.093921 | 0.031289 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210719 | L9 | 14 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.38 | 6.939 | 0.62083 | 10.7592 | 6.3182 | 0.094346 | 0.018167 | 20220519 | 11-81 | 20221201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210719 | L9 | 14 | 20 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.394 | 3.415 | 0.76217 | 3.6318 | 2.6528 | 0.020515 | 0.026083 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR81
EN668 | 20210719 | L9 | 14 | 20 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.574 | 3.201 | 0.76217 | 3.8118 | 2.4388 | 0.02887 | 0.014116 | 20220519 | 11-50 | 20220819 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR82
EN668 | 20210719 | L9 | 14 | 20 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.319 | 1.477 | 0.7595 | 1.5595 | 0.7175 | 0.0094317 | -0.0026947 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR87; neg. phaeo, below detection?; poor replication (47% range) but values are so low
EN668 | 20210719 | L9 | 14 | 20 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.626 | 2.102 | 0.7595 | 1.8665 | 1.3425 | 0.0058483 | 0.0067115 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR88; poor replication (47% range) but values are so low
EN668 | 20210719 | L8 | 15 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.63 | 27.21 | 0.59517 | 36.0348 | 26.6148 | 0.20012 | 0.27383 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | C15N4 chls accidentally filtered 6hr late
EN668 | 20210719 | L8 | 15 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.87 | 25.54 | 0.59517 | 34.2748 | 24.9448 | 0.19821 | 0.246 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | C15N4 chls accidentally filtered 6hr late
EN668 | 20210719 | L8 | 15 | 4 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.94 | 8.487 | 0.7595 | 11.1805 | 7.7275 | 0.072359 | 0.063379 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR97
EN668 | 20210719 | L8 | 15 | 4 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.17 | 13.22 | 0.7595 | 16.4105 | 12.4605 | 0.083056 | 0.13657 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR98
EN668 | 20210719 | L8 | 15 | 4 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.979 | 4.807 | 0.7595 | 5.2195 | 4.0475 | 0.013104 | 0.024831 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR105; poor replication (52% range) but values are so low and likely in noise of flourometer
EN668 | 20210719 | L8 | 15 | 4 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.138 | 6.132 | 0.7595 | 7.3785 | 5.3725 | 0.02243 | 0.027925 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR106; poor replication (52% range) but values are so low and likely in noise of flourometer
EN668 | 20210719 | L8 | 15 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.04 | 22.36 | 0.5195 | 35.5205 | 21.8405 | 0.29062 | 0.09831 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC | there were two sets of capsules labeled "EN668 C15 N8 chla/chlb" --2 sets ofN8WSW and N8<10s look at CTD fluor profile to decide which were corect and which were N7WSW and N14<10s. These values fit the 55m approx fluor
EN668 | 20210719 | L8 | 15 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.21 | 20.98 | 0.5195 | 33.6905 | 20.4605 | 0.28106 | 0.083295 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC | there were two sets of capsules labeled "EN668 C15 N8 chla/chlb"--2 sets ofN8WSW and N8<10s look at CTD fluor profile to decide which were corect and which were N7WSW and N14<10s. These values fit the 55m approx fluor
EN668 | 20210719 | L8 | 15 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 171.9 | 113.5 | 0.59517 | 171.3048 | 112.9048 | 1.2407 | 0.76993 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC | there were two sets of capsules labeled "EN668 C15 N8 chla/chlb" -- I have to assume that one set should be <10s so due to value differences on page 63 of lab notebook and page 73 of labnotebook, I made the values on pg 73 the WSW and 63 the 10s
EN668 | 20210719 | L8 | 15 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 171 | 112.7 | 0.59517 | 170.4048 | 112.1048 | 1.2385 | 0.7578 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC | there were two sets of capsules labeled "EN668 C15 N8 chla/chlb" -- I have to assume that one set should be <10s so due to value differences on page 63 of lab notebook and page 73 of labnotebook, I made the values on pg 73 the WSW and 63 the 10s
EN668 | 20210719 | L8 | 15 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 92.87 | 61.66 | 0.864 | 92.006 | 60.796 | 0.66304 | 0.41961 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC | deduced these tiss caps were actuallly correctly labled
EN668 | 20210719 | L8 | 15 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 97.91 | 64.35 | 0.864 | 97.046 | 63.486 | 0.71296 | 0.41759 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC | deduced these tiss caps were actuallly correctly labled
EN668 | 20210719 | L8 | 15 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 92.09 | 56.13 | 0.7595 | 91.3305 | 55.3705 | 0.75612 | 0.21981 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR95
EN668 | 20210719 | L8 | 15 | 8 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 83.68 | 52.48 | 0.7595 | 82.9205 | 51.7205 | 0.65381 | 0.2547 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR96
EN668 | 20210719 | L8 | 15 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 75.96 | 46.83 | 0.7595 | 75.2005 | 46.0705 | 0.3263 | 0.10628 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR103; poor replication (58% range) within bio variation? Or maybe something was written down incorrecly from lab notebook?
EN668 | 20210719 | L8 | 15 | 8 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 43.75 | 27.78 | 0.7595 | 42.9905 | 27.0205 | 0.17889 | 0.074823 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR104; poor replication (58% range) within bio variation? Or maybe something was written down incorrecly from lab notebook?
EN668 | 20210719 | L8 | 15 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 208.8 | 149.4 | 0.5195 | 208.2805 | 148.8805 | 1.2619 | 1.3893 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | C15N12 chls accidentally filtered 6hr late
EN668 | 20210719 | L8 | 15 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 199.2 | 144.3 | 0.5195 | 198.6805 | 143.7805 | 1.1663 | 1.3941 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | C15N12 chls accidentally filtered 6hr late
EN668 | 20210719 | L8 | 15 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.02 | 34.87 | 0.59517 | 51.4248 | 34.2748 | 0.36434 | 0.24602 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L8 | 15 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 53.46 | 35.75 | 0.5195 | 52.9405 | 35.2305 | 0.37624 | 0.25114 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L8 | 15 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.25 | 17.35 | 0.81783 | 27.4322 | 16.5322 | 0.23156 | 0.062839 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC | tiss cap says C15N8<10a but those already entered. This value could align with C15N14<10a?
EN668 | 20210719 | L8 | 15 | 14 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 30.25 | 19.15 | 0.864 | 29.386 | 18.286 | 0.23995 | 0.091398 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC | C15N14<10b spilled approx 5ml; tiss cap says C15N8<10b but that already exists. This value falls in line with CTD fluor profile and matches other value also estimated to be this depth
EN668 | 20210719 | L8 | 15 | 14 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.367 | 5.004 | 0.7595 | 6.6075 | 4.2445 | 0.049517 | 0.02504 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR99
EN668 | 20210719 | L8 | 15 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.333 | 4.892 | 0.7595 | 6.5735 | 4.1325 | 0.051152 | 0.021438 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR100
EN668 | 20210719 | L8 | 15 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.322 | 4.356 | 0.7595 | 5.5625 | 3.5965 | 0.022022 | 0.011747 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DFthenTC | TAR107; poor replication (94% range) but values so low within noise
EN668 | 20210719 | L8 | 15 | 14 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.846 | 2.139 | 0.7595 | 2.0865 | 1.3795 | 0.0078908 | 0.0050152 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DFthenTC | TAR108; poor replication (94% range) but values so low
EN668 | 20210719 | L8 | 15 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.02 | 9.915 | 0.5195 | 14.5005 | 9.3955 | 0.10845 | 0.058862 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L8 | 15 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.42 | 9.882 | 0.7505 | 14.6695 | 9.1315 | 0.11765 | 0.044961 | 20220519 | 11-77 | 20221107 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210719 | L8 | 15 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.64 | 8.112 | 0.864 | 11.776 | 7.248 | 0.096194 | 0.032877 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210719 | L8 | 15 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.91 | 8.222 | 0.864 | 12.046 | 7.358 | 0.099593 | 0.031437 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210719 | L8 | 15 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.332 | 1.563 | 0.7595 | 1.5725 | 0.8035 | 0.01617 | -0.0020075 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR93; neg phaeo
EN668 | 20210719 | L8 | 15 | 18 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.503 | 2.502 | 0.7595 | 2.7435 | 1.7425 | 0.020976 | 0.0096318 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR94
EN668 | 20210719 | L8 | 15 | 18 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.19 | 1.583 | 0.7595 | 1.4305 | 0.8235 | 0.0067747 | 0.00092961 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR101; ; high z-score but values so low that within noise
EN668 | 20210719 | L8 | 15 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 1.571 | 1.409 | 0.7595 | 0.8115 | 0.6495 | 0.0018147 | 0.0042839 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR102; ; high z-score but values so low that within noise, though could this Rb be an accidental re-read of previous Ra?
EN668 | 20210720 | L7 | 17 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.76 | 13.58 | 0.5195 | 16.2405 | 13.0605 | 0.067557 | 0.16502 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L7 | 17 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.21 | 14.78 | 0.5195 | 16.6905 | 14.2605 | 0.051624 | 0.20233 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L7 | 17 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.92 | 9.052 | 0.59517 | 11.3248 | 8.4568 | 0.060929 | 0.089669 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L7 | 17 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 9.928 | 7.254 | 0.59517 | 9.3328 | 6.6588 | 0.056807 | 0.061772 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L7 | 17 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 60.35 | 39.84 | 0.5195 | 59.8305 | 39.3205 | 0.43572 | 0.26449 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L7 | 17 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 59.93 | 40.03 | 0.5195 | 59.4105 | 39.5105 | 0.42276 | 0.28084 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L7 | 17 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.84 | 27.21 | 0.81783 | 40.0222 | 26.3922 | 0.28956 | 0.18043 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210720 | L7 | 17 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.22 | 27.29 | 0.864 | 39.356 | 26.426 | 0.27469 | 0.1959 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210720 | L7 | 17 | 6 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.62 | 13.4 | 0.7595 | 20.8605 | 12.6405 | 0.17225 | 0.049785 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR113
EN668 | 20210720 | L7 | 17 | 6 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 18.47 | 11.79 | 0.7595 | 17.7105 | 11.0305 | 0.14046 | 0.053959 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR114
EN668 | 20210720 | L7 | 17 | 6 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.96 | 18.26 | 0.7595 | 28.2005 | 17.5005 | 0.11964 | 0.044385 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR119
EN668 | 20210720 | L7 | 17 | 6 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22.44 | 14.15 | 0.7595 | 21.6805 | 13.3905 | 0.092692 | 0.032811 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR120
EN668 | 20210720 | L7 | 17 | 12 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 130.4 | 81.02 | 0.59517 | 129.8048 | 80.4248 | 2.0981 | 0.7663 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | tube broke and parafilm/acetone rubbed off label. DF knows the cruise is correct and thinks its cast17 -- might be another niskin but chl value lines up with replicate
EN668 | 20210720 | L7 | 17 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 254.5 | 156.8 | 0.5195 | 253.9805 | 156.2805 | 2.0756 | 0.70745 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L7 | 17 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 107.2 | 70.11 | 0.864 | 106.336 | 69.246 | 0.78795 | 0.44517 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210720 | L7 | 17 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 105.1 | 68.23 | 0.81783 | 104.2822 | 67.4122 | 0.78328 | 0.41719 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210720 | L7 | 17 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 115.3 | 64.9 | 0.7595 | 114.5405 | 64.1405 | 1.0598 | 0.070759 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR111
EN668 | 20210720 | L7 | 17 | 12 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 145.6 | 84.01 | 0.7595 | 144.8405 | 83.2505 | 1.2906 | 0.17171 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR112
EN668 | 20210720 | L7 | 17 | 12 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 271 | 153.7 | 0.7595 | 270.2405 | 152.9405 | 1.3139 | 0.12211 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR117; should this be moved to a QF of 2? The value is larger than either 5 rep
EN668 | 20210720 | L7 | 17 | 12 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 135.2 | 77.76 | 0.7595 | 134.4405 | 77.0005 | 0.64342 | 0.079587 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | 1 | TAR118; reps mismatchf; in LabNotebook, rep 118 was run second time after vortexing more, value recorded is second running. First number is really low compared to 117 rep. Still poor matchup this sample seems suspcious so toss
EN668 | 20210720 | L7 | 17 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 56.43 | 36.54 | 0.5195 | 55.9105 | 36.0205 | 0.42255 | 0.2189 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L7 | 17 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 46.74 | 31.14 | 0.59517 | 46.1448 | 30.5448 | 0.33141 | 0.21253 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L7 | 17 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.77 | 24.34 | 0.5195 | 38.2505 | 23.8205 | 0.30656 | 0.11764 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L7 | 17 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.99 | 25.21 | 0.5195 | 38.4705 | 24.6905 | 0.29275 | 0.14694 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L7 | 17 | 22 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.16 | 12.45 | 0.59517 | 18.5648 | 11.8548 | 0.14255 | 0.06856 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L7 | 17 | 22 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 18.57 | 11.88 | 0.5195 | 18.0505 | 11.3605 | 0.14212 | 0.060182 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L7 | 17 | 22 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.14 | 9.23 | 0.81783 | 13.3222 | 8.4122 | 0.10431 | 0.045493 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210720 | L7 | 17 | 22 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.45 | 9.285 | 0.81783 | 13.6322 | 8.4672 | 0.10973 | 0.041055 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210720 | L7 | 17 | 22 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.066 | 2.604 | 0.7595 | 3.3065 | 1.8445 | 0.030741 | 0.0017691 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR109
EN668 | 20210720 | L7 | 17 | 22 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.338 | 3.447 | 0.7595 | 4.5785 | 2.6875 | 0.039626 | 0.007581 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR110
EN668 | 20210720 | L7 | 17 | 22 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.546 | 1.842 | 0.7595 | 1.7865 | 1.0825 | 0.0078859 | 0.0022783 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR115; high z-score but values so low that within noise
EN668 | 20210720 | L7 | 17 | 22 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 1.598 | 1.443 | 0.7595 | 0.8385 | 0.6835 | 0.0017362 | 0.0046816 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR116; high z-score but values so low that within noise
EN668 | 20210720 | L6 | 18 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.96 | 18.35 | 0.4865 | 25.4735 | 17.8635 | 0.16167 | 0.15644 | 20220519 | 11-65 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L6 | 18 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.25 | 22.05 | 0.4865 | 30.7635 | 21.5635 | 0.19545 | 0.18855 | 20220519 | 11-65 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L6 | 18 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.48 | 33.34 | 0.4865 | 48.9935 | 32.8535 | 0.34288 | 0.24217 | 20220519 | 11-65 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L6 | 18 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.42 | 42.9 | 0.4865 | 63.9335 | 42.4135 | 0.45718 | 0.29812 | 20220519 | 11-65 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L6 | 18 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 27.36 | 18.84 | 0.4865 | 26.8735 | 18.3535 | 0.181 | 0.14583 | 20220519 | 11-65 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L6 | 18 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.73 | 25.21 | 0.4865 | 34.2435 | 24.7235 | 0.20225 | 0.23803 | 20220519 | 11-65 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L6 | 18 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.99 | 24.73 | 0.4865 | 36.5035 | 24.2435 | 0.26046 | 0.17127 | 20220519 | 11-65 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (57% range); volume filtered was checked..was something incorrectly recorded from lab NB?
EN668 | 20210720 | L6 | 18 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 62.84 | 40.68 | 0.4865 | 62.3535 | 40.1935 | 0.47077 | 0.24499 | 20220519 | 11-65 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (57% range); volume filtered was checked..was something incorrectly recorded from lab NB?
EN668 | 20210720 | L6 | 18 | 5 |  |  | a | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN668 | 20210720 | L6 | 18 | 5 |  |  | b | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN668 | 20210720 | L6 | 18 | 5 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.94 | 11.68 | 0.7595 | 17.1805 | 10.9205 | 0.13118 | 0.060645 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR127
EN668 | 20210720 | L6 | 18 | 5 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.28 | 13.58 | 0.7595 | 19.5205 | 12.8205 | 0.1404 | 0.084799 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR128
EN668 | 20210720 | L6 | 18 | 5 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 9.695 | 5.882 | 0.7595 | 8.9355 | 5.1225 | 0.042712 | 0.0053867 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR135; poor replication (43% range) but values are low and maybe just within noise of flourometer?
EN668 | 20210720 | L6 | 18 | 5 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.35 | 8.402 | 0.7595 | 13.5905 | 7.6425 | 0.066385 | 0.0051146 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR136; oor replication (43% range) but values are low and maybe just within noise of flourometer?
EN668 | 20210720 | L6 | 18 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 159.2 | 99.69 | 0.4865 | 158.7135 | 99.2035 | 1.2643 | 0.50235 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L6 | 18 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 161.2 | 103.5 | 0.4865 | 160.7135 | 103.0135 | 1.2258 | 0.60865 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L6 | 18 | 10 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 218.8 | 131.9 | 0.4865 | 218.3135 | 131.4135 | 3.6923 | 0.98812 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L6 | 18 | 10 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 212.1 | 132.3 | 0.4865 | 211.6135 | 131.8135 | 3.3906 | 1.304 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L6 | 18 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 206.9 | 131 | 0.81783 | 206.0822 | 130.1822 | 1.6124 | 0.70582 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | not great replication but phaeo good match and values less than WSW which is pretty hgih
EN668 | 20210720 | L6 | 18 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 200.6 | 127.3 | 0.81783 | 199.7822 | 126.4822 | 1.5572 | 0.69517 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | not great replication but phaeo good match and values less than WSW which is pretty hgih
EN668 | 20210720 | L6 | 18 | 10 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 182.4 | 106.2 | 0.7595 | 181.6405 | 105.4405 | 1.5968 | 0.25533 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR125
EN668 | 20210720 | L6 | 18 | 10 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 197.2 | 113.8 | 0.7595 | 196.4405 | 113.0405 | 1.7536 | 0.23876 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR126
EN668 | 20210720 | L6 | 18 | 10 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 192.1 | 109.6 | 0.7595 | 191.3405 | 108.8405 | 0.92245 | 0.097663 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR133; poor replication (33% range) but DF thinks that this is just within bio variation
EN668 | 20210720 | L6 | 18 | 10 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 267.5 | 153.7 | 0.7595 | 266.7405 | 152.9405 | 1.2724 | 0.16102 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR134; poor replication (33% range) but DF thinks that this is just within bio variation
EN668 | 20210720 | L6 | 18 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 78.02 | 50.74 | 0.4865 | 77.5335 | 50.2535 | 0.57955 | 0.31536 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | logsheet note says chl bottle vol not written but all sosik chl bott should be 290
EN668 | 20210720 | L6 | 18 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 81.15 | 50.59 | 0.4865 | 80.6635 | 50.1035 | 0.64923 | 0.24301 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L6 | 18 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 54.06 | 34.81 | 0.81783 | 53.2422 | 33.9922 | 0.40895 | 0.19637 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210720 | L6 | 18 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 59.04 | 37.22 | 0.81783 | 58.2222 | 36.4022 | 0.46355 | 0.18469 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210720 | L6 | 18 | 14 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.76 | 18.35 | 0.7595 | 28.0005 | 17.5905 | 0.21814 | 0.090843 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR123
EN668 | 20210720 | L6 | 18 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.97 | 18.39 | 0.7595 | 28.2105 | 17.6305 | 0.22246 | 0.088283 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR124
EN668 | 20210720 | L6 | 18 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.15 | 21.17 | 0.7595 | 33.3905 | 20.4105 | 0.1454 | 0.04625 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR131
EN668 | 20210720 | L6 | 18 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 27.68 | 16.35 | 0.7595 | 26.9205 | 15.5905 | 0.12691 | 0.019475 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR132
EN668 | 20210720 | L6 | 18 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22.15 | 13.89 | 0.4865 | 21.6635 | 13.4035 | 0.17548 | 0.06321 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L6 | 18 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22.53 | 13.98 | 0.4865 | 22.0435 | 13.4935 | 0.18164 | 0.058651 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210720 | L6 | 18 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.52 | 12.22 | 0.81783 | 18.7022 | 11.4022 | 0.15508 | 0.047965 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | not great replication but lower than WSW and low values
EN668 | 20210720 | L6 | 18 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.79 | 11.94 | 0.81783 | 18.9722 | 11.1222 | 0.16677 | 0.031294 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | not great replication but lower than WSW and low values
EN668 | 20210720 | L6 | 18 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.561 | 2.576 | 0.7595 | 2.8015 | 1.8165 | 0.020711 | 0.011305 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR121
EN668 | 20210720 | L6 | 18 | 18 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.793 | 2.704 | 0.7595 | 3.0335 | 1.9445 | 0.02282 | 0.011336 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR122
EN668 | 20210720 | L6 | 18 | 18 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 1.343 | 1.28 | 0.7595 | 0.5835 | 0.5205 | 0.00070314 | 0.0041664 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR129
EN668 | 20210720 | L6 | 18 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.198 | 1.421 | 0.7595 | 1.4385 | 0.6615 | 0.0087036 | -0.0024924 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR130
EN668 | 20210720 | d6a | 20 | 1 |  |  | a | 290 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 236.8 | 137.7 | 0.4865 | 236.3135 | 137.2135 | 12.6319 | 2.029 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | reps a bit off but within reason for such high values
EN668 | 20210720 | d6a | 20 | 1 |  |  | b | 290 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 268.5 | 159.3 | 0.4865 | 268.0135 | 158.8135 | 13.9193 | 3.0495 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | reps a bit off but within reason for such high values
EN668 | 20210720 | d6a | 20 | 1 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 132.6 | 85.02 | 0.81783 | 131.7822 | 84.2022 | 1.0108 | 0.48866 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | pour replication but such high concentrated WSW that these seem reasonable
EN668 | 20210720 | d6a | 20 | 1 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 116.8 | 76.83 | 0.81783 | 115.9822 | 76.0122 | 0.84914 | 0.50448 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | pour replication but such high concentrated WSW that these seem reasonable
EN668 | 20210720 | d6a | 20 | 1 |  |  | a | 293 | 5 | 5 | 1 | 0 | 0.25 | Cal_051922a | 1.8382 | 0.56187 | 376.7 | 227.1 | 0.7595 | 375.9405 | 226.3405 | 12.5824 | 3.375 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR137
EN668 | 20210720 | d6a | 20 | 1 |  |  | b | 294 | 5 | 5 | 1 | 0 | 0.25 | Cal_051922a | 1.8382 | 0.56187 | 368.1 | 211.7 | 0.7595 | 367.3405 | 210.9405 | 13.1096 | 1.7115 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR138
EN668 | 20210720 | d6a | 20 | 1 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.1 | Cal_051922a | 1.8382 | 0.56187 | 253.1 | 137 | 0.7595 | 252.3405 | 136.2405 | 13.005 | -0.21256 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR139; negative phaeo
EN668 | 20210720 | d6a | 20 | 1 |  |  | b | 552 | 20 | 5 | 1 | 0 | 0.1 | Cal_051922a | 1.8382 | 0.56187 | 270.8 | 150.2 | 0.7595 | 270.0405 | 149.4405 | 13.4601 | 0.5209 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR140
EN668 | 20210720 | d6a | 20 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.88 | 14.05 | 0.4865 | 21.3935 | 13.5635 | 0.16634 | 0.075194 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | such low values samples are probably fine but could chla and <10b been swapped?
EN668 | 20210720 | d6a | 20 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23 | 14.65 | 0.4865 | 22.5135 | 14.1635 | 0.17739 | 0.074832 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | such low values samples are probably fine but could chla and <10b been swapped?
EN668 | 20210720 | d6a | 20 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22 | 14.18 | 0.81783 | 21.1822 | 13.3622 | 0.16613 | 0.071821 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | such low values samples are probably fine but could chla and <10b been swapped?
EN668 | 20210720 | d6a | 20 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.83 | 13.25 | 0.81783 | 21.0122 | 12.4322 | 0.18228 | 0.039114 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | such low values samples are probably fine but could chla and <10b been swapped?
EN668 | 20210720 | d6a | 20 | 16 |  |  | a | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.625 | 4.084 | 0.7595 | 5.8655 | 3.3245 | 0.053612 | 0.0051847 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR141
EN668 | 20210720 | d6a | 20 | 16 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.8 | 3.806 | 0.7595 | 5.0405 | 3.0465 | 0.041927 | 0.011769 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR142
EN668 | 20210720 | d6a | 20 | 16 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.96 | 2.118 | 0.7595 | 2.2005 | 1.3585 | 0.0094317 | 0.0033241 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR143
EN668 | 20210720 | d6a | 20 | 16 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.564 | 2.347 | 0.7595 | 1.8045 | 1.5875 | 0.0024307 | 0.012475 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 1 | TAR144
EN668 | 20210721 | L3 | 21 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.92 | 44.43 | 0.4865 | 66.4335 | 43.9435 | 0.47778 | 0.30476 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | L3 | 21 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.63 | 42.95 | 0.4865 | 63.1435 | 42.4635 | 0.43933 | 0.31685 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | L3 | 21 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 68.38 | 45.5 | 0.4865 | 67.8935 | 45.0135 | 0.48607 | 0.31552 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | L3 | 21 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.61 | 43.5 | 0.4865 | 66.1235 | 43.0135 | 0.49096 | 0.27502 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | L3 | 21 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 120.9 | 77.59 | 0.4865 | 120.4135 | 77.1035 | 0.92009 | 0.45296 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | L3 | 21 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 121.8 | 78.67 | 0.4865 | 121.3135 | 78.1835 | 0.91627 | 0.47601 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | L3 | 21 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 198.8 | 132.7 | 0.4865 | 198.3135 | 132.2135 | 1.4043 | 0.95019 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | logsheet note-tissue cap labeled incorrectly-says C19N11 but should be C21N11
EN668 | 20210721 | L3 | 21 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 193 | 131.3 | 0.4865 | 192.5135 | 130.8135 | 1.3108 | 1.0187 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | logsheet note-tissue cap labeled incorrectly-says C19N11 but should be C21N11
EN668 | 20210721 | L3 | 21 | 11 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 59.46 | 38.2 | 0.7595 | 58.7005 | 37.4405 | 0.44703 | 0.21288 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR149
EN668 | 20210721 | L3 | 21 | 11 |  |  | b | 292 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.09 | 40.22 | 0.7595 | 64.3305 | 39.4605 | 0.52473 | 0.17317 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR150
EN668 | 20210721 | L3 | 21 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 50.05 | 29.31 | 0.7595 | 49.2905 | 28.5505 | 0.23232 | 0.035758 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR155; this rep is half the other rep? 79% range
EN668 | 20210721 | L3 | 21 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 111.2 | 63.18 | 0.7595 | 110.4405 | 62.4205 | 0.5379 | 0.048205 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR156; this rep is double the other rep? 79 % range. Based on that fact that this is larger than either 5 rep, I think this one shold be tossed and assume other rep is fine
EN668 | 20210721 | L3 | 21 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 97.74 | 60.29 | 0.4865 | 97.2535 | 59.8035 | 0.7956 | 0.26937 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | L3 | 21 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 94.57 | 59.88 | 0.4865 | 94.0835 | 59.3935 | 0.73697 | 0.32071 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | L3 | 21 | 16 |  |  | a | 295 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.95 | 17.14 | 0.7595 | 28.1905 | 16.3805 | 0.24664 | 0.040115 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR147
EN668 | 20210721 | L3 | 21 | 16 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 24.65 | 15.4 | 0.7595 | 23.8905 | 14.6405 | 0.19384 | 0.063332 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR148
EN668 | 20210721 | L3 | 21 | 16 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.85 | 13.02 | 0.7595 | 21.0905 | 12.2605 | 0.09873 | 0.016182 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR153
EN668 | 20210721 | L3 | 21 | 16 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.12 | 11.05 | 0.7595 | 18.3605 | 10.2905 | 0.090232 | 0.0062158 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR154
EN668 | 20210721 | L3 | 21 | 21 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 26.6 | 16.97 | 0.4865 | 26.1135 | 16.4835 | 0.20458 | 0.088953 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | L3 | 21 | 21 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 27.74 | 17.85 | 0.4865 | 27.2535 | 17.3635 | 0.21011 | 0.0991 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | L3 | 21 | 21 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.13 | 13.41 | 0.81783 | 20.3122 | 12.5922 | 0.16401 | 0.060233 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210721 | L3 | 21 | 21 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.96 | 13.42 | 0.81783 | 20.1422 | 12.6022 | 0.16018 | 0.064235 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210721 | L3 | 21 | 21 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.24 | 5.544 | 0.7595 | 7.4805 | 4.7845 | 0.056688 | 0.027641 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR145
EN668 | 20210721 | L3 | 21 | 21 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.034 | 4.039 | 0.7595 | 5.2745 | 3.2795 | 0.041806 | 0.015801 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR146
EN668 | 20210721 | L3 | 21 | 21 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.72 | 2.613 | 0.7595 | 2.9605 | 1.8535 | 0.0124 | 0.0050035 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR151; poor replication (41%) but values extremely low and within nosie of flourometer, basically 0
EN668 | 20210721 | L3 | 21 | 21 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.377 | 2.644 | 0.7595 | 2.6175 | 1.8845 | 0.008181 | 0.0094496 | 20220519 | 11-42 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR152; poor replication (41%) but values extremely low and within nosie of flourometer, basically 0
EN668 | 20210721 | MVCO | 25 | 2 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 222.4 | 137.4 | 0.4865 | 221.9135 | 136.9135 | 3.6115 | 1.2647 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | MVCO | 25 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 224 | 140.3 | 0.4865 | 223.5135 | 139.8135 | 3.5563 | 1.4233 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | MVCO | 25 | 5 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 228.3 | 144.9 | 0.4865 | 227.8135 | 144.4135 | 3.5436 | 1.5998 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | MVCO | 25 | 5 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 230.8 | 145.6 | 0.4865 | 230.3135 | 145.1135 | 3.62 | 1.5483 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | MVCO | 25 | 14 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 224 | 137.9 | 0.4865 | 223.5135 | 137.4135 | 3.6583 | 1.2358 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | MVCO | 25 | 14 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 221.7 | 138.9 | 0.4865 | 221.2135 | 138.4135 | 3.5181 | 1.4116 | 20220519 | 11-66 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN668 | 20210721 | MVCO | 25 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 299.7 | 190.2 | 0.81783 | 298.8822 | 189.3822 | 2.3263 | 1.0462 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN668 | 20210721 | MVCO | 25 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 275.4 | 173.9 | 0.81783 | 274.5822 | 173.0822 | 2.1563 | 0.92592 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61A | 20211030 | L10 | 5 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 72.54 | 62.07 | 0.6765 | 71.8635 | 61.3935 | 0.061026 | 0.23893 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 05/01; not counted as OOI sample but still have ID; OffshoreSurfaceMoor | phae>chl; depth 61m
AR61A | 20211030 | L10 | 5 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 180.7 | 118.7 | 0.6765 | 180.0235 | 118.0235 | 0.72343 | 0.43093 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/02; not counted as OOI sample but still have ID; OffshoreSurfaceMoor
AR61A | 20211030 | L10 | 5 | 9 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 185.4 | 121.1 | 0.6765 | 184.7235 | 120.4235 | 0.75027 | 0.42757 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/03; OffshoreSurfMoor work
AR61A | 20211030 | L10 | 5 | 10 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 191 | 123.1 | 0.6765 | 190.3235 | 122.4235 | 0.79228 | 0.40513 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/04; OffshoreSurfMoor work
AR61A | 20211101 | L6 | 13 | 3 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 168.6 | 108.2 | 0.6765 | 167.9235 | 107.5235 | 0.7041 | 0.34658 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 13/01; InshoreSurfMoor work
AR61A | 20211101 | L6 | 13 | 4 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 315 | 207.4 | 0.62 | 314.38 | 206.78 | 0.62775 | 0.38349 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 13/02; InshoreSurfMoor work
AR61A | 20211101 | L6 | 13 | 5 |  |  | a | 1056 | 0 | 15 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 150 | 90.46 | 0.6765 | 149.3235 | 89.7835 | 3.1263 | 0.82543 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 13/03; InshoreSurfMoor work
AR61A | 20211101 | L6 | 13 | 6 |  |  | a | 1056 | 0 | 15 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 146.5 | 86.81 | 0.6765 | 145.8235 | 86.1335 | 3.1342 | 0.6569 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 13/04; InshoreSurfMoor work
AR61A | 20211101 | L6 | 13 | 7 |  |  | a | 1057 | 0 | 15 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 144.2 | 87.73 | 0.6765 | 143.5235 | 87.0535 | 2.9623 | 0.86565 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 13/05; InshoreSurfMoor work; poor replication
AR61A | 20211101 | L6 | 13 | 8 |  |  | a | 1056 | 0 | 15 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 167.1 | 99.58 | 0.6765 | 166.4235 | 98.9035 | 3.5453 | 0.80782 | 20220519 | 13-23 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 13/06; InshoreSurfMoor work; poor replication
AR61A | 20211103 | L8 | 14 | 6 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 189 | 126.7 | 0.6765 | 188.3235 | 126.0235 | 0.72625 | 0.5052 | 20220519 | 13-24 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 14/01; CentralSurfMoor work
AR61A | 20211103 | L8 | 14 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 189.8 | 125.3 | 0.6765 | 189.1235 | 124.6235 | 0.7526 | 0.46632 | 20220519 | 13-24 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 14/02; CentralSurfMoor work
AR61A | 20211103 | L8 | 14 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 157.6 | 104 | 0.6765 | 156.9235 | 103.3235 | 0.62542 | 0.38517 | 20220519 | 13-24 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 14/03; CentralSurfMoor work
AR61A | 20211103 | L8 | 14 | 9 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 155.4 | 103.4 | 0.6765 | 154.7235 | 102.7235 | 0.60675 | 0.39797 | 20220519 | 13-24 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 14/04; CentralSurfMoor work
AR61A | 20211104 |  | 15 | 5 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 247.8 | 164.6 | 0.6765 | 247.1235 | 163.9235 | 0.96988 | 0.63191 | 20220519 | 13-24 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 15/01; AUV deploy AC1
AR61A | 20211104 |  | 15 | 5 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 255.6 | 170.6 | 0.6765 | 254.9235 | 169.9235 | 0.9918 | 0.67019 | 20220519 | 13-24 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 15/02; AUV deploy AC1
AR61A | 20211104 |  | 16 | 3 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 195.6 | 128.5 | 0.6765 | 194.9235 | 127.8235 | 0.7822 | 0.46684 | 20220519 | 13-24 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 16/01; AUV deploy AL-1
AR61A | 20211104 |  | 16 | 4 |  |  | a | 1056 | 0 | 15 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 149.2 | 94.24 | 0.6765 | 148.5235 | 93.5635 | 2.8858 | 1.2323 | 20220519 | 13-24 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 16/02; AUV deploy AL-1; poor replication
AR61A | 20211105 |  | 16 | 4 |  |  | b | 1056 | 0 | 15 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 170.7 | 105.9 | 0.6765 | 170.0235 | 105.2235 | 3.4025 | 1.2288 | 20220519 | 13-24 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 16/03; AUV deploy AL-1; poor replication
AR61A | 20211105 |  | 17 | 4 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 176.6 | 119.2 | 0.6765 | 175.9235 | 118.5235 | 0.66913 | 0.48904 | 20220519 | 13-24 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 17/01; AUV recover AC-4
AR61A | 20211105 |  | 17 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 132.8 | 89.67 | 0.6765 | 132.1235 | 88.9935 | 0.50325 | 0.36718 | 20220519 | 13-24 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 17/02; AUV recover AC-4
AR61A | 20211106 |  | 18 | 3 |  |  | a | 1057 | 0 | 15 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 147.7 | 88.06 | 0.6765 | 147.0235 | 87.3835 | 3.1286 | 0.71387 | 20220519 | 13-24 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 18/01; AUV recover AL-4
AR61A | 20211106 |  | 18 | 4 |  |  | a | 1056 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 374.2 | 224.4 | 0.6765 | 373.5235 | 223.7235 | 2.6219 | 0.66044 | 20220519 | 13-24 | 20230328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 18/02; AUV recover AL-4
AR61B | 20211108 | L1 | 1 | 1 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 185.4 | 121.7 | 0.632 | 184.768 | 121.068 | 2.7065 | 1.6054 | 20220519 | 11-84 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L1 | 1 | 1 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 193.1 | 123.3 | 0.719 | 192.381 | 122.581 | 2.9657 | 1.4001 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L1 | 1 | 3 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 194.9 | 126.8 | 0.632 | 194.268 | 126.168 | 2.8935 | 1.6001 | 20220519 | 11-84 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L1 | 1 | 3 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 194 | 124.7 | 0.632 | 193.368 | 124.068 | 2.9445 | 1.4743 | 20220519 | 11-84 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L1 | 1 | 6 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 187.8 | 120.2 | 0.632 | 187.168 | 119.568 | 2.8722 | 1.3863 | 20220519 | 11-84 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L1 | 1 | 6 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 192.4 | 120.9 | 0.632 | 191.768 | 120.268 | 3.0379 | 1.2455 | 20220519 | 11-84 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L1 | 1 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 241.9 | 153.6 | 0.81783 | 241.0822 | 152.7822 | 1.8759 | 0.84485 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211108 | L1 | 1 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 249.7 | 159.8 | 0.81783 | 248.8822 | 158.9822 | 1.9099 | 0.92127 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211108 | L2 | 2 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 112.2 | 89.67 | 0.632 | 111.568 | 89.038 | 0.47863 | 1.1069 | 20220519 | 11-84 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L2 | 2 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 121.1 | 91.05 | 0.632 | 120.468 | 90.418 | 0.63839 | 0.97176 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L2 | 2 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 261 | 166.6 | 0.632 | 260.368 | 165.968 | 2.0055 | 0.95007 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L2 | 2 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 267.4 | 170.9 | 0.632 | 266.768 | 170.268 | 2.0501 | 0.98203 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L2 | 2 | 4 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 146.1 | 89.12 | 0.802 | 145.298 | 88.318 | 1.1981 | 0.35854 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR5
AR61B | 20211108 | L2 | 2 | 4 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 137.7 | 84.64 | 0.802 | 136.898 | 83.838 | 1.1157 | 0.36201 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR6
AR61B | 20211108 | L2 | 2 | 4 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.41 | 48.95 | 0.802 | 79.608 | 48.148 | 0.3524 | 0.09969 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR11; poor replication (45 % range) but within the bio variation of size fractions
AR61B | 20211108 | L2 | 2 | 4 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 123.3 | 73.4 | 0.802 | 122.498 | 72.598 | 0.55896 | 0.12271 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR1; poor replication (45 % range) but within the bio variation of size fractions
AR61B | 20211108 | L2 | 2 | 8 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 159.1 | 100.9 | 0.632 | 158.468 | 100.268 | 2.4728 | 1.0983 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L2 | 2 | 8 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 159.3 | 99.12 | 0.632 | 158.668 | 98.488 | 2.557 | 0.95075 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L2 | 2 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 162 | 98.28 | 0.802 | 161.198 | 97.478 | 1.3398 | 0.37827 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR3
AR61B | 20211108 | L2 | 2 | 8 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 174.9 | 106.4 | 0.802 | 174.098 | 105.598 | 1.4354 | 0.41945 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR4
AR61B | 20211108 | L2 | 2 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 172.2 | 102.6 | 0.802 | 171.398 | 101.798 | 0.77963 | 0.17621 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR9
AR61B | 20211108 | L2 | 2 | 8 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 190.4 | 112.8 | 0.802 | 189.598 | 111.998 | 0.86766 | 0.18204 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR10
AR61B | 20211108 | L2 | 2 | 12 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 159.3 | 94.42 | 0.719 | 158.581 | 93.701 | 2.7567 | 0.58057 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC | poor replication
AR61B | 20211108 | L2 | 2 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 294.9 | 184.3 | 0.81783 | 294.0822 | 183.4822 | 2.3496 | 0.9178 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC | poor replication
AR61B | 20211108 | L2 | 2 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 188.6 | 120.7 | 0.632 | 187.968 | 120.068 | 1.4425 | 0.69566 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211108 | L2 | 2 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 183.1 | 116.3 | 0.81783 | 182.2822 | 115.4822 | 1.4191 | 0.63737 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211108 | L2 | 2 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 180.4 | 108.2 | 0.802 | 179.598 | 107.398 | 1.5181 | 0.37481 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR1
AR61B | 20211108 | L2 | 2 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 170.8 | 101.3 | 0.802 | 169.998 | 100.498 | 1.4614 | 0.30997 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR2
AR61B | 20211108 | L2 | 2 | 12 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 193.5 | 116.3 | 0.802 | 192.698 | 115.498 | 0.86476 | 0.21972 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR7
AR61B | 20211108 | L2 | 2 | 12 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 168.1 | 97.55 | 0.802 | 167.298 | 96.748 | 0.78741 | 0.11773 | 20220519 | 11-55 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR8
AR61B | 20211108 | L3 | 3 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 79.74 | 58.1 | 0.632 | 79.108 | 57.468 | 0.45973 | 0.56366 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L3 | 3 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 94.72 | 66 | 0.632 | 94.088 | 65.368 | 0.61014 | 0.55393 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L3 | 3 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 197.9 | 124.2 | 0.632 | 197.268 | 123.568 | 1.5657 | 0.63478 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L3 | 3 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 199.9 | 126.9 | 0.632 | 199.268 | 126.268 | 1.5508 | 0.69773 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L3 | 3 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 206.8 | 128 | 0.632 | 206.168 | 127.368 | 1.6741 | 0.5941 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L3 | 3 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 187.3 | 117.8 | 0.632 | 186.668 | 117.168 | 1.4765 | 0.61003 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L3 | 3 | 9 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 192.6 | 119.8 | 0.719 | 191.881 | 119.081 | 3.0932 | 1.148 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L3 | 3 | 9 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 211.9 | 131.8 | 0.719 | 211.181 | 131.081 | 3.4033 | 1.2652 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L3 | 3 | 11 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 198 | 123.7 | 0.719 | 197.281 | 122.981 | 3.1569 | 1.2232 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L3 | 3 | 11 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 208.1 | 128.1 | 0.632 | 207.468 | 127.468 | 3.3991 | 1.1408 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211108 | L3 | 3 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 206.4 | 130.5 | 0.81783 | 205.5822 | 129.6822 | 1.6124 | 0.69692 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211108 | L3 | 3 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 206.1 | 130.9 | 0.81783 | 205.2822 | 130.0822 | 1.5976 | 0.71891 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211109 | L4 | 4 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 150 | 100.2 | 0.719 | 149.281 | 99.481 | 1.058 | 0.71358 | 20220519 | 11-87 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L4 | 4 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 151.1 | 103.7 | 0.632 | 150.468 | 103.068 | 1.007 | 0.82844 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L4 | 4 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 78.87 | 58.17 | 0.632 | 78.238 | 57.538 | 0.43976 | 0.58487 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L4 | 4 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 91.86 | 67.76 | 0.632 | 91.228 | 67.128 | 0.51199 | 0.68342 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L4 | 4 | 8 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 182.3 | 110.6 | 0.632 | 181.668 | 109.968 | 3.0464 | 0.87015 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L4 | 4 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 360.6 | 215.9 | 0.719 | 359.881 | 215.181 | 3.0741 | 0.75786 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L4 | 4 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 256.5 | 152.9 | 0.802 | 255.698 | 152.098 | 2.1784 | 0.50243 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR17
AR61B | 20211109 | L4 | 4 | 8 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 293 | 172.1 | 0.802 | 292.198 | 171.298 | 2.5421 | 0.47708 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR18
AR61B | 20211109 | L4 | 4 | 8 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 251.9 | 144.7 | 0.802 | 251.098 | 143.898 | 1.1986 | 0.15007 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR23; poor replication (38% range) but within bio variation?
AR61B | 20211109 | L4 | 4 | 8 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 381.7 | 224.2 | 0.802 | 380.898 | 223.398 | 1.761 | 0.33277 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR24; poor replication (38% range) but within bio variation?
AR61B | 20211109 | L4 | 4 | 12 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 196.9 | 120.3 | 0.632 | 196.268 | 119.668 | 3.2546 | 1.0074 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | somewhat poor replication (15% range)
AR61B | 20211109 | L4 | 4 | 12 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 226.7 | 137.7 | 0.719 | 225.981 | 136.981 | 3.7815 | 1.0972 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | somewhat poor replication (15% range)
AR61B | 20211109 | L4 | 4 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 351.4 | 205 | 0.802 | 350.598 | 204.198 | 3.0783 | 0.52077 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR15
AR61B | 20211109 | L4 | 4 | 12 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 321 | 192.3 | 0.802 | 320.198 | 191.498 | 2.6969 | 0.66683 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR16
AR61B | 20211109 | L4 | 4 | 12 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 161 | 93.14 | 0.802 | 160.198 | 92.338 | 2.2804 | 0.32064 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR21
AR61B | 20211109 | L4 | 4 | 12 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 138 | 81.68 | 0.802 | 137.198 | 80.878 | 1.8926 | 0.38562 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR22
AR61B | 20211109 | L4 | 4 | 16 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR61B | 20211109 | L4 | 4 | 16 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 196.1 | 120.9 | 0.719 | 195.381 | 120.181 | 3.1951 | 1.0852 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211109 | L4 | 4 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 155.2 | 98.17 | 0.81783 | 154.3822 | 97.3522 | 1.2116 | 0.52207 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211109 | L4 | 4 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 151.3 | 94.65 | 0.81783 | 150.4822 | 93.8322 | 1.2035 | 0.46746 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211109 | L4 | 4 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 359.1 | 213.9 | 0.802 | 358.298 | 213.098 | 3.0531 | 0.70287 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR13
AR61B | 20211109 | L4 | 4 | 16 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 329.2 | 195.1 | 0.802 | 328.398 | 194.298 | 2.8197 | 0.60491 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR14
AR61B | 20211109 | L4 | 4 | 16 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 168.9 | 98.29 | 0.802 | 168.098 | 97.488 | 2.3728 | 0.3733 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR19
AR61B | 20211109 | L4 | 4 | 16 |  |  | b | 552 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 144.1 | 85.74 | 0.802 | 143.298 | 84.938 | 1.9541 | 0.42987 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR20
AR61B | 20211109 | L5 | 6 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 204 | 129.1 | 0.632 | 203.368 | 128.468 | 1.5912 | 0.69654 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L5 | 6 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 206.2 | 136.9 | 0.632 | 205.568 | 136.268 | 1.4722 | 0.95441 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L5 | 6 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 322.1 | 198.2 | 0.632 | 321.468 | 197.568 | 2.6322 | 0.88609 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L5 | 6 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 332.1 | 203.5 | 0.719 | 331.381 | 202.781 | 2.732 | 0.87908 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L5 | 6 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 360.6 | 224.5 | 0.632 | 359.968 | 223.868 | 2.8914 | 1.0953 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L5 | 6 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 365.4 | 223.2 | 0.632 | 364.768 | 222.568 | 3.0209 | 0.94252 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L5 | 6 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 353.7 | 218.1 | 0.719 | 352.981 | 217.381 | 2.8807 | 0.99036 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication; possible mixup N10 and N12 but will never know so flag as questionable
AR61B | 20211109 | L5 | 6 | 10 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 200.3 | 119.1 | 0.719 | 199.581 | 118.381 | 3.4501 | 0.76615 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication; possible mixup N10 and N12 but will never know so flag as questionable
AR61B | 20211109 | L5 | 6 | 12 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 195 | 119.3 | 0.719 | 194.281 | 118.581 | 3.2164 | 1.007 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication; possible mixup N10 and N12 but will never know so flag as questionable
AR61B | 20211109 | L5 | 6 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 327.5 | 204.2 | 0.632 | 326.868 | 203.568 | 2.6194 | 1.0057 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication; possible mixup N10 and N12 but will never know so flag as questionable
AR61B | 20211109 | L5 | 6 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 135.8 | 87.65 | 0.81783 | 134.9822 | 86.8322 | 1.0229 | 0.52338 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211109 | L5 | 6 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 141.1 | 91.51 | 0.81783 | 140.2822 | 90.6922 | 1.0535 | 0.56153 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211109 | L10 | 7 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.48 | 12.47 | 0.632 | 16.848 | 11.838 | 0.10643 | 0.10438 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L10 | 7 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.53 | 12.61 | 0.632 | 15.898 | 11.978 | 0.083278 | 0.13002 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L10 | 7 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.75 | 14.31 | 0.632 | 19.118 | 13.678 | 0.11557 | 0.12801 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L10 | 7 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.82 | 14.55 | 0.632 | 19.188 | 13.918 | 0.11196 | 0.13589 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L10 | 7 | 14 |  |  | a | 194 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 91.92 | 58.43 | 0.719 | 91.201 | 57.711 | 1.0635 | 0.47273 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | filled late didn't filter 106ml
AR61B | 20211109 | L10 | 7 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 130.7 | 83.7 | 0.632 | 130.068 | 83.068 | 0.99848 | 0.48078 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L10 | 7 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 161.8 | 103.1 | 0.632 | 161.168 | 102.468 | 1.247 | 0.57769 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L10 | 7 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 158 | 98.68 | 0.719 | 157.281 | 97.961 | 1.2602 | 0.48426 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211109 | L10 | 7 | 22 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 126.2 | 77.12 | 0.632 | 125.568 | 76.488 | 1.0427 | 0.31942 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 07/01; OffshoreSurfMoor
AR61B | 20211109 | L10 | 7 | 22 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 110.3 | 69.84 | 0.719 | 109.581 | 69.121 | 0.85955 | 0.37135 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 07/02; OffshoreSurfMoor
AR61B | 20211109 | L10 | 7 | 22 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 51.52 | 32.44 | 0.81783 | 50.7022 | 31.6222 | 0.40534 | 0.15778 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC | poor replication
AR61B | 20211109 | L10 | 7 | 22 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 58.48 | 36.21 | 0.81783 | 57.6622 | 35.3922 | 0.47311 | 0.15715 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC | poor replication
AR61B | 20211109 | L12 | 8 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 149.9 | 91.73 | 0.632 | 149.268 | 91.098 | 1.2358 | 0.38648 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 08/01; UpstreamOffshoreProfiler; reps off
AR61B | 20211109 | L12 | 8 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 108.6 | 93.84 | 0.632 | 107.968 | 93.208 | 0.31357 | 1.3463 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | OOI 08/02; UpstreamOffshoreProfiler; reps off…likely error recording Rb value from fluorometer. Likely Rb was 148.6 which would make values fit. Flag as 3 and throw out though can't assume
AR61B | 20211110 | L11 | 9 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.29 | 11.95 | 0.632 | 13.658 | 11.318 | 0.049712 | 0.15184 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L11 | 9 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.51 | 13.44 | 0.632 | 15.878 | 12.808 | 0.06522 | 0.16286 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L11 | 9 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 59.86 | 40.31 | 0.719 | 59.141 | 39.591 | 0.41533 | 0.2897 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L11 | 9 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 59.25 | 41.22 | 0.719 | 58.531 | 40.501 | 0.38304 | 0.3382 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L11 | 9 | 9 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 46.58 | 30.67 | 0.802 | 45.778 | 29.868 | 0.3334 | 0.19125 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR31
AR61B | 20211110 | L11 | 9 | 9 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 47.84 | 32.59 | 0.802 | 47.038 | 31.788 | 0.31957 | 0.23881 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR32
AR61B | 20211110 | L11 | 9 | 9 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.76 | 35.44 | 0.802 | 51.958 | 34.638 | 0.19331 | 0.13075 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR39
AR61B | 20211110 | L11 | 9 | 9 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 44.68 | 29.2 | 0.802 | 43.878 | 28.398 | 0.17309 | 0.093077 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR40
AR61B | 20211110 | L11 | 9 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.08 | 43.37 | 0.719 | 63.361 | 42.651 | 0.43997 | 0.31955 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (51% range)
AR61B | 20211110 | L11 | 9 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 104.2 | 69.05 | 0.719 | 103.481 | 68.331 | 0.74674 | 0.47009 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (51% range)
AR61B | 20211110 | L11 | 9 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 50.65 | 33.84 | 0.802 | 49.848 | 33.038 | 0.35346 | 0.22885 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR29
AR61B | 20211110 | L11 | 9 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 54.42 | 36.5 | 0.802 | 53.618 | 35.698 | 0.3768 | 0.2524 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR30
AR61B | 20211110 | L11 | 9 | 12 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 50.47 | 31.21 | 0.802 | 49.668 | 30.408 | 0.21535 | 0.06965 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR37
AR61B | 20211110 | L11 | 9 | 12 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 62.7 | 39.69 | 0.802 | 61.898 | 38.888 | 0.25728 | 0.1072 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR38
AR61B | 20211110 | L11 | 9 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 179.6 | 117.1 | 0.632 | 178.968 | 116.468 | 1.3278 | 0.74628 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L11 | 9 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 146.5 | 94.42 | 0.632 | 145.868 | 93.788 | 1.1064 | 0.56376 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L11 | 9 | 16 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 130.1 | 81.08 | 0.802 | 129.298 | 80.278 | 1.0272 | 0.3829 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR27
AR61B | 20211110 | L11 | 9 | 16 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 132.2 | 81.62 | 0.802 | 131.398 | 80.818 | 1.0635 | 0.36092 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR28
AR61B | 20211110 | L11 | 9 | 16 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 188.1 | 114.1 | 0.802 | 187.298 | 113.298 | 0.82892 | 0.23491 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR35; poor replication (31% range) but within bio variation
AR61B | 20211110 | L11 | 9 | 16 |  |  | b | 548 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 135.6 | 82.07 | 0.802 | 134.798 | 81.268 | 0.60181 | 0.16405 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR36; poor replication (31% range) but within bio variation
AR61B | 20211110 | L11 | 9 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 178.5 | 113.2 | 0.632 | 177.868 | 112.568 | 1.3873 | 0.61734 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L11 | 9 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 156.6 | 98.82 | 0.632 | 155.968 | 98.188 | 1.2275 | 0.52102 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L11 | 9 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 43.86 | 29.15 | 0.81783 | 43.0422 | 28.3322 | 0.3125 | 0.19203 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211110 | L11 | 9 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 42.5 | 28.16 | 0.81783 | 41.6822 | 27.3422 | 0.30464 | 0.18226 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211110 | L11 | 9 | 20 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 119.7 | 75.53 | 0.802 | 118.898 | 74.728 | 0.92875 | 0.38837 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR25
AR61B | 20211110 | L11 | 9 | 20 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 128.3 | 78.86 | 0.802 | 127.498 | 78.058 | 1.0396 | 0.33625 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR26
AR61B | 20211110 | L11 | 9 | 20 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 181.6 | 113.9 | 0.802 | 180.798 | 113.098 | 0.7556 | 0.3025 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR33
AR61B | 20211110 | L11 | 9 | 20 |  |  | b | 548 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 159.9 | 99.14 | 0.802 | 159.098 | 98.338 | 0.68309 | 0.24363 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR34
AR61B | 20211110 | L9 | 11 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.24 | 5.899 | 0.719 | 6.521 | 5.18 | 0.028489 | 0.063756 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L9 | 11 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.949 | 4.978 | 0.632 | 5.317 | 4.346 | 0.020628 | 0.056765 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L9 | 11 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.45 | 11.81 | 0.719 | 13.731 | 11.091 | 0.056085 | 0.14142 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 11/01
AR61B | 20211110 | L9 | 11 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.53 | 11.57 | 0.719 | 13.811 | 10.851 | 0.062883 | 0.13035 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 11/02
AR61B | 20211110 | L9 | 11 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.31 | 12.27 | 0.632 | 15.678 | 11.638 | 0.085827 | 0.12142 | 20220519 | 11-85 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L9 | 11 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16 | 12.53 | 0.719 | 15.281 | 11.811 | 0.073718 | 0.13661 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L9 | 11 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.04 | 28.83 | 0.632 | 39.408 | 28.198 | 0.23815 | 0.264 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L9 | 11 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.76 | 29.1 | 0.719 | 40.041 | 28.381 | 0.24771 | 0.2577 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L9 | 11 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 95.97 | 60.66 | 0.632 | 95.338 | 60.028 | 0.75014 | 0.31883 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L9 | 11 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 95.25 | 61.66 | 0.632 | 94.618 | 61.028 | 0.7136 | 0.37318 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L9 | 11 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 108.1 | 68.94 | 0.632 | 107.468 | 68.308 | 0.83193 | 0.38449 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 11/03
AR61B | 20211110 | L9 | 11 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 107 | 69.79 | 0.632 | 106.368 | 69.158 | 0.7905 | 0.44105 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 11/04
AR61B | 20211110 | L9 | 11 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 116.5 | 74.73 | 0.719 | 115.781 | 74.011 | 0.88738 | 0.4306 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L9 | 11 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 119.6 | 76.35 | 0.632 | 118.968 | 75.718 | 0.91882 | 0.42956 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L9 | 11 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 124.9 | 79.93 | 0.632 | 124.268 | 79.298 | 0.95536 | 0.45677 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 11/05
AR61B | 20211110 | L9 | 11 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 122.7 | 77.8 | 0.632 | 122.068 | 77.168 | 0.95387 | 0.42033 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 11/06
AR61B | 20211110 | L9 | 11 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 79.29 | 51.25 | 0.632 | 78.658 | 50.618 | 0.59569 | 0.30571 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | 0
AR61B | 20211110 | L9 | 11 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 79.88 | 51.43 | 0.81783 | 79.0622 | 50.6122 | 0.6044 | 0.29689 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211110 | L8 | 12 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 114.3 | 71.96 | 0.632 | 113.668 | 71.328 | 0.89948 | 0.37071 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 12/01
AR61B | 20211110 | L8 | 12 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 102 | 66.35 | 0.632 | 101.368 | 65.718 | 0.75736 | 0.41294 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 12/02
AR61B | 20211110 | L8 | 12 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 114.6 | 70.54 | 0.719 | 113.881 | 69.821 | 0.93603 | 0.30734 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L8 | 12 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 106 | 66.69 | 0.719 | 105.281 | 65.971 | 0.83511 | 0.33969 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L8 | 12 | 6 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 93.99 | 57 | 0.802 | 93.188 | 56.198 | 0.77514 | 0.21201 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR47
AR61B | 20211110 | L8 | 12 | 6 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 86.63 | 51.88 | 0.802 | 85.828 | 51.078 | 0.7282 | 0.16902 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR48
AR61B | 20211110 | L8 | 12 | 6 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 104.9 | 63.53 | 0.802 | 104.098 | 62.728 | 0.46173 | 0.12513 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR55
AR61B | 20211110 | L8 | 12 | 6 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 130.8 | 78.6 | 0.802 | 129.998 | 77.798 | 0.58366 | 0.14551 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR56
AR61B | 20211110 | L8 | 12 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 114.6 | 72.45 | 0.632 | 113.968 | 71.818 | 0.89545 | 0.38348 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L8 | 12 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 105.1 | 66.31 | 0.632 | 104.468 | 65.678 | 0.82407 | 0.34552 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211110 | L8 | 12 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 124.3 | 75.39 | 0.802 | 123.498 | 74.588 | 1.0284 | 0.28623 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR45
AR61B | 20211110 | L8 | 12 | 10 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 110.8 | 67.73 | 0.802 | 109.998 | 66.928 | 0.90562 | 0.27402 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR46
AR61B | 20211110 | L8 | 12 | 10 |  |  | a | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 38.02 | 22.78 | 0.802 | 37.218 | 21.978 | 0.51121 | 0.10676 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR53
AR61B | 20211110 | L8 | 12 | 10 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 151.3 | 89.43 | 0.802 | 150.498 | 88.628 | 0.69178 | 0.13889 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR54
AR61B | 20211110 | L8 | 12 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 184.4 | 112.3 | 0.632 | 183.768 | 111.668 | 1.5317 | 0.45685 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | TC after DF2 | OOI 12/03; poor replication (26% range)
AR61B | 20211110 | L8 | 12 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 148.8 | 93.48 | 0.632 | 148.168 | 92.848 | 1.1752 | 0.47819 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | TC after DF2 | OOI 12/04; poor replication (26% range)
AR61B | 20211110 | L8 | 12 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 144.4 | 87.52 | 0.802 | 143.598 | 86.718 | 1.196 | 0.33245 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR43
AR61B | 20211110 | L8 | 12 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 150.5 | 90.16 | 0.802 | 149.698 | 89.358 | 1.2644 | 0.30518 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR44
AR61B | 20211110 | L8 | 12 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 232.7 | 138.7 | 0.802 | 231.898 | 137.898 | 1.0529 | 0.24186 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR51
AR61B | 20211110 | L8 | 12 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 198.1 | 116.9 | 0.802 | 197.298 | 116.098 | 0.90957 | 0.18055 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR52
AR61B | 20211110 | L8 | 12 | 17 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 132.2 | 81.02 | 0.632 | 131.568 | 80.388 | 2.1746 | 0.68851 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (27% range)
AR61B | 20211110 | L8 | 12 | 17 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 102.1 | 63.32 | 0.632 | 101.468 | 62.688 | 1.6477 | 0.58497 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (27% range)
AR61B | 20211110 | L8 | 12 | 21 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 193.9 | 115 | 0.632 | 193.268 | 114.368 | 3.3524 | 0.72095 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 12/05; reps off
AR61B | 20211110 | L8 | 12 | 21 |  |  | b | 285 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 183.6 | 110.6 | 0.632 | 182.968 | 109.968 | 3.1561 | 0.82922 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 12/06; spilled approx 5ml C12N21chlb; reps off
AR61B | 20211110 | L8 | 12 | 21 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 110 | 68.9 | 0.81783 | 109.1822 | 68.0822 | 0.87314 | 0.33926 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | TC | somewhat poor replication but high WSW value maybe makes higher variation feasible
AR61B | 20211110 | L8 | 12 | 21 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 95.07 | 60.99 | 0.81783 | 94.2522 | 60.1722 | 0.72401 | 0.34753 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | TC | somewhat poor replication but high WSW value maybe makes higher variation feasible
AR61B | 20211110 | L8 | 12 | 21 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 303.4 | 178 | 0.802 | 302.598 | 177.198 | 2.6368 | 0.48645 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR41
AR61B | 20211110 | L8 | 12 | 21 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 288.7 | 166.8 | 0.802 | 287.898 | 165.998 | 2.5632 | 0.36264 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR42
AR61B | 20211110 | L8 | 12 | 21 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 166.4 | 99.9 | 0.802 | 165.598 | 99.098 | 2.2347 | 0.55676 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR49
AR61B | 20211110 | L8 | 12 | 21 |  |  | b | 552 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 148.4 | 88.47 | 0.802 | 147.598 | 87.668 | 2.0066 | 0.45392 | 20220519 | 11-56 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR50
AR61B | 20211111 | L6 | 18 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.49 | 48.03 | 0.719 | 64.771 | 47.311 | 0.37093 | 0.47158 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 18/01; InshoreProfilerMoor; the reps for this whole cast seem off but nothing seems askew on logsheets. Maybe just bio variation at this station?
AR61B | 20211111 | L6 | 18 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76.3 | 54.16 | 0.632 | 75.668 | 53.528 | 0.47035 | 0.48287 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 18/02; InshoreProfilerMoor; the reps for this whole cast seem off but nothing seems askew on logsheets. Maybe just bio variation at this station?
AR61B | 20211111 | L6 | 18 | 5 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 291.8 | 175.8 | 0.719 | 291.081 | 175.081 | 4.9287 | 1.307 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | the reps for this whole cast seem off but nothing seems askew on logsheets. Maybe just bio variation at this station?
AR61B | 20211111 | L6 | 18 | 5 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 238.5 | 143.3 | 0.719 | 237.781 | 142.581 | 4.0449 | 1.0332 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | the reps for this whole cast seem off but nothing seems askew on logsheets. Maybe just bio variation at this station?
AR61B | 20211111 | L6 | 18 | 8 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 258.3 | 153.6 | 0.719 | 257.581 | 152.881 | 4.4486 | 0.99641 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | OOI 18/03; InshoreProfilerMoor; the reps for this whole cast seem off but nothing seems askew on logsheets. Maybe just bio variation at this station?
AR61B | 20211111 | L6 | 18 | 8 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 307 | 185.5 | 0.632 | 306.368 | 184.868 | 5.1624 | 1.4218 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | OOI 18/04; InshoreProfilerMoor; the reps for this whole cast seem off but nothing seems askew on logsheets. Maybe just bio variation at this station?
AR61B | 20211111 | L6 | 18 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 190.2 | 109 | 0.802 | 189.398 | 108.198 | 5.1221 | 0.59901 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR61; poor replication (41% range); based on sample logs, clear that 61 is good, so keep this value seems like there was confusion with 62
AR61B | 20211111 | L6 | 18 | 8 |  |  | b | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 129.6 | 76 | 0.802 | 128.798 | 75.198 | 3.3811 | 0.5951 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR62; poor replication (41%) There is a note in logsheet about sample mixup -- DF think this should be switched with TAR 66. No DF doesn't think this is the case because of the volume filtered makese the reps worse if they're switched. Is there a lab NB error? DF thinks this is the wrong sample and should be tossed based on confusion with samples 62-65
AR61B | 20211111 | L6 | 18 | 8 |  |  | a | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 332.5 | 192.7 | 0.802 | 331.698 | 191.898 | 4.6894 | 0.70632 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR67
AR61B | 20211111 | L6 | 18 | 8 |  |  | b | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 252.1 | 144.3 | 0.802 | 251.298 | 143.498 | 3.616 | 0.41882 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR68
AR61B | 20211111 | L6 | 18 | 10 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 324 | 194 | 0.719 | 323.281 | 193.281 | 5.5235 | 1.3603 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | the reps for this whole cast seem off but nothing seems askew on logsheets. Maybe just bio variation at this station? | very high negative phaeo toss. Assume other rep is fine
AR61B | 20211111 | L6 | 18 | 10 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 355.3 | 210.7 | 0.632 | 354.668 | 210.068 | 6.1439 | 1.3379 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | the reps for this whole cast seem off but nothing seems askew on logsheets. Maybe just bio variation at this station?
AR61B | 20211111 | L6 | 18 | 14 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 350.4 | 209.4 | 0.632 | 349.768 | 208.768 | 5.9909 | 1.4445 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | OOI 18/05; InshoreProfilerMoor; the reps for this whole cast seem off but nothing seems askew on logsheets. Maybe just bio variation at this station?
AR61B | 20211111 | L6 | 18 | 14 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 289.6 | 171.5 | 0.632 | 288.968 | 170.868 | 5.0179 | 1.0677 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | OOI 18/06; InshoreProfilerMoor; the reps for this whole cast seem off but nothing seems askew on logsheets. Maybe just bio variation at this station?
AR61B | 20211111 | L6 | 18 | 14 |  |  | a | 294 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 202.5 | 118.2 | 0.802 | 201.698 | 117.398 | 5.2996 | 0.8869 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR59
AR61B | 20211111 | L6 | 18 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 213.2 | 124.3 | 0.802 | 212.398 | 123.498 | 5.6079 | 0.9223 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR60
AR61B | 20211111 | L6 | 18 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 224.9 | 126.8 | 0.802 | 224.098 | 125.998 | 3.2966 | 0.2526 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR64; poor replication (42% range); DF thinks ths is the correct value for this sample type based on logsheets
AR61B | 20211111 | L6 | 18 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 359.3 | 208.7 | 0.802 | 358.498 | 207.898 | 5.0609 | 0.79538 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR66; Poor replication (42 % range) There is a note in logsheet about sample mixup -- DF thinks this should be switched with TAR 62. No DF doesn't think this is the case because of the volume filtered makese the reps worse if they're switched. Is there a lab NB error? Or something got switched with this sample? unsure which one this could be. toss
AR61B | 20211111 | L6 | 18 | 19 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 268.7 | 159.7 | 0.632 | 268.068 | 159.068 | 4.6313 | 1.0341 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | s
AR61B | 20211111 | L6 | 18 | 19 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 269.9 | 158.4 | 0.719 | 269.181 | 157.681 | 4.7375 | 0.87844 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 18/08; InshoreProfilerMoor
AR61B | 20211111 | L6 | 18 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 131.6 | 82.32 | 0.81783 | 130.7822 | 81.5022 | 1.0469 | 0.40446 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211111 | L6 | 18 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 149.5 | 95.54 | 0.81783 | 148.6822 | 94.7222 | 1.1463 | 0.54046 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211111 | L6 | 18 | 19 |  |  | a | 293 | 5 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | 9 |  | TAR57; this sample was thrown out (it was left out overnight and not extracted)
AR61B | 20211111 | L6 | 18 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 196 | 113.3 | 0.802 | 195.198 | 112.498 | 5.2168 | 0.73176 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR58
AR61B | 20211111 | L6 | 18 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 310.1 | 178.8 | 0.802 | 309.298 | 177.998 | 4.4123 | 0.6017 | [6 empty] | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR63
AR61B | 20211111 | L6 | 18 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 311.4 | 180.1 | 0.802 | 310.598 | 179.298 | 4.3963 | 0.63601 | 20220519 | 11-57 | 20220914 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR65
AR61B | 20211112 | L13 | 29 | 6 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 357.9 | 211.8 | 0.81783 | 357.0822 | 210.9822 | 6.2076 | 1.3067 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 29/01; UpstreamInshoreProfiler
AR61B | 20211112 | L13 | 29 | 6 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 379.3 | 234.5 | 0.81783 | 378.4822 | 233.6822 | 6.1524 | 2.1704 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 29/02; UpstreamInshoreProfiler
AR61B | 20211112 | L13 | 29 | 8 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 363.9 | 224.8 | 0.81783 | 363.0822 | 223.9822 | 5.9102 | 2.0671 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 29/03; UpstreamInshoreProfiler
AR61B | 20211112 | L13 | 29 | 8 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 345.6 | 204.6 | 0.81783 | 344.7822 | 203.7822 | 5.9909 | 1.267 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 29/04; UpstreamInshoreProfiler
AR61B | 20211114 | L7 | 32 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.84 | 44.18 | 0.632 | 63.208 | 43.548 | 0.41766 | 0.35783 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (35% range)
AR61B | 20211114 | L7 | 32 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 47.73 | 33.9 | 0.719 | 47.011 | 33.181 | 0.29381 | 0.29707 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (35% range)
AR61B | 20211114 | L7 | 32 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 106.1 | 69.58 | 0.632 | 105.468 | 68.948 | 0.77584 | 0.45197 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 32/01; CentralInshoreProfiler
AR61B | 20211114 | L7 | 32 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 100.3 | 66.59 | 0.632 | 99.668 | 65.958 | 0.71615 | 0.45842 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 32/02; CentralInshoreProfiler
AR61B | 20211114 | L7 | 32 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 164.5 | 105.2 | 0.719 | 163.781 | 104.481 | 1.2598 | 0.60079 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211114 | L7 | 32 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 159.2 | 103.8 | 0.632 | 158.568 | 103.168 | 1.1769 | 0.66027 | 20220519 | 11-86 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211114 | L7 | 32 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 336.7 | 201.4 | 0.719 | 335.981 | 200.681 | 2.8744 | 0.69934 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211114 | L7 | 32 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 326.2 | 197.9 | 0.719 | 325.481 | 197.181 | 2.7256 | 0.78573 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AR61B | 20211114 | L7 | 32 | 15 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 198.7 | 122.3 | 0.719 | 197.981 | 121.581 | 3.2461 | 1.0841 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 32/03;volume entry missing on logsheet but confident sampled and vol is correct
AR61B | 20211114 | L7 | 32 | 15 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 200.7 | 121.5 | 0.719 | 199.981 | 120.781 | 3.3651 | 0.9366 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 32/04; volume entry missing on logsheet but confident sampled and vol is correct
AR61B | 20211114 | L7 | 32 | 18 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 186.5 | 114.2 | 0.719 | 185.781 | 113.481 | 3.0719 | 0.96978 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 32/05; CentralInshoreProfiler
AR61B | 20211114 | L7 | 32 | 18 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 177.7 | 109.4 | 0.719 | 176.981 | 108.681 | 2.902 | 0.96878 | 20220519 | 11-88 | 20221204 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | OOI 32/06; CentralInshoreProfiler
AR61B | 20211114 | L7 | 32 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 125.9 | 79.56 | 0.81783 | 125.0822 | 78.7422 | 0.98446 | 0.41777 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR61B | 20211114 | L7 | 32 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 124.9 | 79.81 | 0.81783 | 124.0822 | 78.9922 | 0.95791 | 0.44878 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220216 | L1 | 1 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 414.9 | 244.3 | 0.5395 | 414.3605 | 243.7605 | 3.6243 | 0.71657 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220216 | L1 | 1 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 209.9 | 124 | 0.667 | 209.233 | 123.333 | 3.6498 | 0.74282 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220216 | L1 | 1 | 2 |  |  | a | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [9 empty] | sample exists on cruise logsheet but might not have been sampled. They might have meant to have entered on surf niskin only. Time will tell
AT46 | 20220216 | L1 | 1 | 2 |  |  | b | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [9 empty] | sample exists on cruise logsheet but might not have been sampled. They might have meant to have entered on surf niskin only. Time will tell
AT46 | 20220216 | L1 | 1 | 2 |  |  | a | 290 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 181.4 | 106.5 | 0.93767 | 180.4623 | 105.5623 | 1.5912 | 0.28864 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR5; very slight leak >1ml; this rep is smaller than the other and had a leak -- adjust volume? Small leak though | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220216 | L1 | 1 | 2 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 235 | 133.9 | 0.93767 | 234.0623 | 132.9623 | 2.1258 | 0.21772 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR6 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220216 | L1 | 1 | 2 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 430.6 | 238.8 | 0.73067 | 429.8693 | 238.0693 | 2.1446 | 0.08676 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR11; these values are larger than the 5s could they be switched?
AT46 | 20220216 | L1 | 1 | 2 |  |  | b | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 172.8 | 98.73 | 0.73067 | 172.0693 | 97.9993 | 2.4846 | 0.27093 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR12; these values are larger than the 5s, could they be swtiched?
AT46 | 20220216 | L1 | 1 | 7 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 217.2 | 127 | 0.667 | 216.533 | 126.333 | 3.8325 | 0.66697 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220216 | L1 | 1 | 7 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 225.6 | 131.5 | 0.7225 | 224.8775 | 130.7775 | 3.9982 | 0.65955 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220216 | L1 | 1 | 7 |  |  | a | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [9 empty] | sample not on logsheet but possibly sampled because high priority station. confident on volume
AT46 | 20220216 | L1 | 1 | 7 |  |  | b | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [9 empty] | sample not on logsheet but possibly sampled because high priority station. confident on volume
AT46 | 20220216 | L1 | 1 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 248.4 | 138.9 | 0.93767 | 247.4623 | 137.9623 | 2.3024 | 0.12922 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR3 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220216 | L1 | 1 | 7 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 306.1 | 171 | 0.93767 | 305.1623 | 170.0623 | 2.8407 | 0.15672 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR4 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220216 | L1 | 1 | 7 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 147.6 | 81.87 | 0.73067 | 146.8693 | 81.1393 | 2.2088 | 0.076764 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR9
AT46 | 20220216 | L1 | 1 | 7 |  |  | b | 548 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 389.6 | 214.5 | 0.73067 | 388.8693 | 213.7693 | 1.9686 | 0.045984 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR10; leak 2 ml
AT46 | 20220216 | L1 | 1 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 387.8 | 230 | 0.667 | 387.133 | 229.333 | 3.3524 | 0.73158 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220216 | L1 | 1 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 390 | 231 | 0.667 | 389.333 | 230.333 | 3.3778 | 0.72389 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220216 | L1 | 1 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 108.6 | 68.13 | 0.864 | 107.736 | 67.266 | 0.85976 | 0.33811 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | sample not on logsheet but N2<10s logged. Possibly mixup on logsheet and they meant to log N13. confident on volume
AT46 | 20220216 | L1 | 1 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 98.52 | 62.18 | 0.864 | 97.656 | 61.316 | 0.77202 | 0.31989 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | sample not on logsheet but N2<10s logged. Possibly mixup on logsheet and they meant to log N13. confident on volume
AT46 | 20220216 | L1 | 1 | 13 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 277.8 | 157.6 | 0.93767 | 276.8623 | 156.6623 | 2.5274 | 0.23383 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR1 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220216 | L1 | 1 | 13 |  |  | b | 283 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 262.9 | 149.1 | 0.93767 | 261.9623 | 148.1623 | 2.4774 | 0.22631 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR2; spilled about 10ml | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220216 | L1 | 1 | 13 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 350.8 | 191.3 | 0.73067 | 350.0693 | 190.5693 | 1.7866 | 0.0027222 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR7;
AT46 | 20220216 | L1 | 1 | 13 |  |  | b | 552 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 152.6 | 86.86 | 0.73067 | 151.8693 | 86.1293 | 2.2012 | 0.2162 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR8;
AT46 | 20220217 | L4 | 2 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 207.5 | 124.1 | 0.7225 | 206.7775 | 123.3775 | 1.7718 | 0.42531 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L4 | 2 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 199.1 | 121.7 | 0.7225 | 198.3775 | 120.9775 | 1.6443 | 0.51004 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L4 | 2 | 2 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.16 | 26.81 | 0.5395 | 39.6205 | 26.2705 | 0.28361 | 0.18421 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L4 | 2 | 2 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.96 | 23.91 | 0.864 | 34.096 | 23.046 | 0.23475 | 0.17565 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L4 | 2 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 171.2 | 102 | 0.7225 | 170.4775 | 101.2775 | 1.4701 | 0.33343 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L4 | 2 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 179.4 | 107.8 | 0.7225 | 178.6775 | 107.0775 | 1.5211 | 0.38573 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L4 | 2 | 8 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 87.57 | 52.96 | 0.7225 | 86.8475 | 52.2375 | 1.4705 | 0.38994 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L4 | 2 | 8 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 110.4 | 66.34 | 0.7225 | 109.6775 | 65.6175 | 1.8721 | 0.46497 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L4 | 2 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 35.61 | 24.15 | 0.5395 | 35.0705 | 23.6105 | 0.24346 | 0.17699 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L4 | 2 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 37.56 | 25.21 | 0.864 | 36.696 | 24.346 | 0.26237 | 0.17118 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L4 | 2 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 117.7 | 65.67 | 0.93767 | 116.7623 | 64.7323 | 1.094 | 0.046917 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR17; reps a bit off but not bad? | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L4 | 2 | 8 |  |  | b | 290 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 165 | 96.61 | 0.93767 | 164.0623 | 95.6723 | 1.4529 | 0.25082 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR18; >3ml leak | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L4 | 2 | 8 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 227.4 | 123.7 | 0.93767 | 226.4623 | 122.7623 | 1.1595 | -0.0088956 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR23; somewhat negative phaeo | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L4 | 2 | 8 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 257.7 | 137.7 | 0.93767 | 256.7623 | 136.7623 | 1.3417 | -0.059934 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR24; somewhat negative phaeo | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L4 | 2 | 14 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 97.61 | 58.18 | 0.7225 | 96.8875 | 57.4575 | 1.6753 | 0.37106 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L4 | 2 | 14 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 100.4 | 59.2 | 0.7225 | 99.6775 | 58.4775 | 1.7505 | 0.33219 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L4 | 2 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 188.6 | 103.4 | 0.93767 | 187.6623 | 102.4623 | 1.7915 | 0.014469 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR15; reps a bit off but not bad; this rep is higher than either of the two WSW samples | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L4 | 2 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 146.4 | 82.52 | 0.93767 | 145.4623 | 81.5823 | 1.3432 | 0.094739 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR16; reps a bit off but not bad | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L4 | 2 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 193.4 | 109.8 | 0.73067 | 192.6693 | 109.0693 | 0.93645 | 0.087668 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | TAR21; poor replication (96% range)
AT46 | 20220217 | L4 | 2 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.29 | 35.26 | 0.93767 | 63.3523 | 34.3223 | 0.32518 | -0.0029081 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR22; neg phaeo, poor replication (96% range)…this value seems low compared to size fractionation ratios at other depths. DF thinks toss and assume other rep is fine | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L4 | 2 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 215.3 | 126.3 | 0.667 | 214.633 | 125.633 | 1.8907 | 0.34651 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L4 | 2 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 181.4 | 108.8 | 0.7225 | 180.6775 | 108.0775 | 1.5423 | 0.38229 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L4 | 2 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.83 | 25.04 | 0.5395 | 36.2905 | 24.5005 | 0.25047 | 0.18583 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L4 | 2 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.8 | 24.68 | 0.5395 | 36.2605 | 24.1405 | 0.25748 | 0.17241 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L4 | 2 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 148.3 | 87.28 | 0.93767 | 147.3623 | 86.3423 | 1.2831 | 0.23877 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR13 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L4 | 2 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 165.4 | 95.75 | 0.93767 | 164.4623 | 94.8123 | 1.4645 | 0.2066 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR14 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L4 | 2 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 227.7 | 129 | 0.73067 | 226.9693 | 128.2693 | 1.1056 | 0.098805 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR19
AT46 | 20220217 | L4 | 2 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 270.8 | 150.5 | 0.73067 | 270.0693 | 149.7693 | 1.3427 | 0.058515 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | ; higher than one of the 5 reps but bio variation and if 20s are averaged, the value is lower than the averaged 5 value
AT46 | 20220217 | L5 | 3 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76.56 | 48.79 | 0.7225 | 75.8375 | 48.0675 | 0.58996 | 0.26602 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L5 | 3 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 77.42 | 49.74 | 0.7225 | 76.6975 | 49.0175 | 0.58804 | 0.28485 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L5 | 3 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 68.7 | 44.77 | 0.7225 | 67.9775 | 44.0475 | 0.50838 | 0.27602 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L5 | 3 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 82.94 | 51.59 | 0.7225 | 82.2175 | 50.8675 | 0.66601 | 0.23983 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L5 | 3 | 4 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23.58 | 14.31 | 0.93767 | 22.6423 | 13.3723 | 0.19426 | 0.040637 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR27; poor replication (89% range) DF thinks something is up with this sample given 20 values | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L5 | 3 | 4 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 59.11 | 34.9 | 0.73067 | 58.3793 | 34.1693 | 0.50906 | 0.093194 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR28; poor replication (89% range)
AT46 | 20220217 | L5 | 3 | 4 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.49 | 22.58 | 0.73067 | 37.7593 | 21.8493 | 0.17822 | 0.02694 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR33
AT46 | 20220217 | L5 | 3 | 4 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.12 | 27.33 | 0.93767 | 48.1823 | 26.3923 | 0.24408 | 0.0037316 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR34 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L5 | 3 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 136.3 | 83.37 | 0.667 | 135.633 | 82.703 | 1.1245 | 0.3483 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L5 | 3 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 123.9 | 76.69 | 0.7225 | 123.1775 | 75.9675 | 1.0029 | 0.34987 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L5 | 3 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 109 | 67.28 | 0.7225 | 108.2775 | 66.5575 | 0.88631 | 0.29893 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L5 | 3 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 111.4 | 69.17 | 0.7225 | 110.6775 | 68.4475 | 0.89715 | 0.32176 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L5 | 3 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 50.6 | 31.32 | 0.73067 | 49.8693 | 30.5893 | 0.4054 | 0.13376 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR29; poor replication (47% range) but within biological variation of size fractions so okay
AT46 | 20220217 | L5 | 3 | 10 |  |  | b | 290 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.38 | 19.61 | 0.73067 | 30.6493 | 18.8793 | 0.25005 | 0.086155 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR30; poor replication (47% range) but within biological variation of size fractions so okay
AT46 | 20220217 | L5 | 3 | 10 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.91 | 19.75 | 0.93767 | 33.9723 | 18.8123 | 0.16951 | 0.0068124 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR35; poor replication (40% range) but within bio variation of size fractions | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L5 | 3 | 10 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.63 | 29.41 | 0.93767 | 51.6923 | 28.4723 | 0.25963 | 0.0072306 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR36;poor replication (40% range) and higher than one of the 5 reps, but averaged 20 values are less than averaged 5 values so okay and within bio variation | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L5 | 3 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 108.9 | 66.95 | 0.7225 | 108.1775 | 66.2275 | 0.8912 | 0.28817 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L5 | 3 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 133.1 | 82.57 | 0.667 | 132.433 | 81.903 | 1.0735 | 0.38504 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L5 | 3 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 51.75 | 33.51 | 0.5395 | 51.2105 | 32.9705 | 0.3875 | 0.19964 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L5 | 3 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 50.71 | 33.13 | 0.5395 | 50.1705 | 32.5905 | 0.37348 | 0.20689 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L5 | 3 | 13 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.8 | 26.98 | 0.93767 | 44.8623 | 26.0423 | 0.39572 | 0.063285 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR25 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L5 | 3 | 13 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.36 | 22.66 | 0.93767 | 37.4223 | 21.7223 | 0.33012 | 0.052747 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR26 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L5 | 3 | 13 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 79.61 | 47.09 | 0.73067 | 78.8793 | 46.3593 | 0.36295 | 0.070764 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR31
AT46 | 20220217 | L5 | 3 | 13 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 60.06 | 35.35 | 0.73067 | 59.3293 | 34.6193 | 0.27679 | 0.048271 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR32
AT46 | 20220217 | L6 | 6 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 187.6 | 115.2 | 0.7225 | 186.8775 | 114.4775 | 1.5381 | 0.50051 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L6 | 6 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 204.4 | 124.1 | 0.7225 | 203.6775 | 123.3775 | 1.7059 | 0.49117 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L6 | 6 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 144.2 | 90.72 | 0.7225 | 143.4775 | 89.9975 | 1.1361 | 0.46652 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L6 | 6 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 156.6 | 96.95 | 0.7225 | 155.8775 | 96.2275 | 1.2672 | 0.44638 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L6 | 6 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 157.4 | 97.73 | 0.7225 | 156.6775 | 97.0075 | 1.2676 | 0.45985 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L6 | 6 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 159.5 | 97.94 | 0.7225 | 158.7775 | 97.2175 | 1.3078 | 0.42344 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L6 | 6 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 184.1 | 113.9 | 0.7225 | 183.3775 | 113.1775 | 1.4914 | 0.5241 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L6 | 6 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 183.1 | 109.6 | 0.7225 | 182.3775 | 108.8775 | 1.5615 | 0.37742 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L6 | 6 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.16 | 42.11 | 0.5395 | 62.6205 | 41.5705 | 0.44719 | 0.29309 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L6 | 6 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 68.98 | 45.24 | 0.5395 | 68.4405 | 44.7005 | 0.50434 | 0.29168 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L6 | 6 | 7 |  |  | a | 295 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76.29 | 42.8 | 0.93767 | 75.3523 | 41.8623 | 0.69941 | 0.03343 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR41; lab notebook says this sample was diluted but was it actually 45 that was diluted? DF thinks this is the case yes and updated dilution factor to reflect this | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L6 | 6 | 7 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76.29 | 43.48 | 0.93767 | 75.3523 | 42.5423 | 0.68989 | 0.059942 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR42 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L6 | 6 | 7 |  |  | a | 495 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 181.6 | 98.59 | 0.93767 | 180.6623 | 97.6523 | 1.0332 | -0.014361 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR47 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L6 | 6 | 7 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 190 | 102.4 | 0.93767 | 189.0623 | 101.4623 | 0.97948 | -0.028513 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR48 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L6 | 6 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 184.4 | 111 | 0.7225 | 183.6775 | 110.2775 | 1.5593 | 0.40447 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L6 | 6 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 185.8 | 114 | 0.667 | 185.133 | 113.333 | 1.5253 | 0.49288 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L6 | 6 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 179.2 | 110 | 0.7225 | 178.4775 | 109.2775 | 1.4701 | 0.47589 | 20220519 | 11-94 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L6 | 6 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 176.2 | 108.6 | 0.7225 | 175.4775 | 107.8775 | 1.4361 | 0.48495 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L6 | 6 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.61 | 42.88 | 0.864 | 64.746 | 42.016 | 0.48288 | 0.26533 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L6 | 6 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.27 | 41.92 | 0.5395 | 62.7305 | 41.3805 | 0.45357 | 0.28333 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L6 | 6 | 13 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 77.26 | 43.44 | 0.7307 | 76.5293 | 42.7093 | 0.71113 | 0.041647 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR39; clearly something up with this sample given it is loewr than the 20s.
AT46 | 20220217 | L6 | 6 | 13 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 174.8 | 96.51 | 0.93767 | 173.8623 | 95.5723 | 1.6406 | 0.038194 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR40; higher than both WSW -- just bio variation and that most of the chl signal is in the >5? | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L6 | 6 | 13 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 95.39 | 52.21 | 0.93767 | 94.4523 | 51.2723 | 1.451 | -0.006769 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR45; was this sample actually diluted? Yes DF thinks so; neg phaeo | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L6 | 6 | 13 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 264.6 | 146.1 | 0.93767 | 263.6623 | 145.1623 | 1.3274 | 0.035632 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR46 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L6 | 6 | 18 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 86.94 | 53.03 | 0.7225 | 86.2175 | 52.3075 | 1.4408 | 0.42218 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L6 | 6 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 165.1 | 100.6 | 0.7225 | 164.3775 | 99.8775 | 1.3703 | 0.40835 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L6 | 6 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 68.96 | 45.05 | 0.5395 | 68.4205 | 44.5105 | 0.50795 | 0.28469 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L6 | 6 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.54 | 43.81 | 0.5395 | 65.0005 | 43.2705 | 0.46164 | 0.30892 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L6 | 6 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 163.3 | 93.97 | 0.73067 | 162.5693 | 93.2393 | 1.4578 | 0.1856 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR37; higher than both WSW replicates but averaged 5 values are lower than averaged WSW values
AT46 | 20220217 | L6 | 6 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 122.5 | 70.14 | 0.73067 | 121.7693 | 69.4093 | 1.101 | 0.12241 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR38
AT46 | 20220217 | L6 | 6 | 18 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 164.5 | 90.27 | 0.93767 | 163.5623 | 89.3323 | 0.82848 | 0.0072772 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR43 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L6 | 6 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 207.9 | 119.6 | 0.93767 | 206.9623 | 118.6623 | 0.9891 | 0.12509 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR44 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L11 | 7 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.65 | 15.88 | 0.667 | 20.983 | 15.213 | 0.12258 | 0.14833 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L11 | 7 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.31 | 13.96 | 0.7225 | 18.5875 | 13.2375 | 0.11366 | 0.12207 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L11 | 7 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 94.21 | 62.59 | 0.7225 | 93.4875 | 61.8675 | 0.67175 | 0.42998 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L11 | 7 | 4 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.78 | 32.96 | 0.7225 | 49.0575 | 32.2375 | 0.71466 | 0.4335 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220217 | L11 | 7 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 118.7 | 76.73 | 0.7225 | 117.9775 | 76.0075 | 0.89162 | 0.46191 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | *THIS DEPTH DOESN'T HAVE <10 RECORDED
AT46 | 20220217 | L11 | 7 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 125.1 | 79.97 | 0.7225 | 124.3775 | 79.2475 | 0.95876 | 0.45247 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | *THIS DEPTH DOESN'T HAVE <10 RECORDED
AT46 | 20220217 | L11 | 7 | 5 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 79.03 | 48.55 | 0.7595 | 78.2705 | 47.7905 | 0.6409 | 0.20144 | 20220519 | 11-40 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR55
AT46 | 20220217 | L11 | 7 | 5 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 85.84 | 52.69 | 0.7595 | 85.0805 | 51.9305 | 0.69704 | 0.21826 | 20220519 | 11-40 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR56
AT46 | 20220217 | L11 | 7 | 5 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 128.5 | 77.94 | 0.7595 | 127.7405 | 77.1805 | 0.5643 | 0.15777 | 20220519 | 11-40 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR63
AT46 | 20220217 | L11 | 7 | 5 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 131.1 | 77.42 | 0.7595 | 130.3405 | 76.6605 | 0.6013 | 0.11851 | 20220519 | 11-40 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR64
AT46 | 20220217 | L11 | 7 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 130.7 | 83.81 | 0.667 | 130.033 | 83.143 | 0.99615 | 0.48445 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | *THIS DEPTH DOESN'T HAVE <10 RECORDED
AT46 | 20220217 | L11 | 7 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 136.3 | 80.85 | 0.7225 | 135.5775 | 80.1275 | 1.178 | 0.2489 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | *THIS DEPTH DOESN'T HAVE <10 RECORDED
AT46 | 20220217 | L11 | 7 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 83.18 | 49.67 | 0.7595 | 82.4205 | 48.9105 | 0.70461 | 0.15747 | 20220519 | 11-40 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR53
AT46 | 20220217 | L11 | 7 | 9 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 84.21 | 50.79 | 0.7595 | 83.4505 | 50.0305 | 0.70272 | 0.1791 | 20220519 | 11-40 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR54
AT46 | 20220217 | L11 | 7 | 9 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 85.49 | 50.72 | 0.7595 | 84.7305 | 49.9605 | 0.38807 | 0.079343 | 20220519 | 11-40 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR61; poor replication (42% range) but within the bio variation of size fractions
AT46 | 20220217 | L11 | 7 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 130.5 | 77.36 | 0.7595 | 129.7405 | 76.6005 | 0.59525 | 0.124 | 20220519 | 11-40 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR62; poor replication (42% range) but within the bio variation of size fractions
AT46 | 20220217 | L11 | 7 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 124 | 77.28 | 0.7225 | 123.2775 | 76.5575 | 0.99254 | 0.37079 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | *THIS DEPTH DOESN'T HAVE <10 RECORDED
AT46 | 20220217 | L11 | 7 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 116.7 | 72.02 | 0.7225 | 115.9775 | 71.2975 | 0.9492 | 0.32046 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | *THIS DEPTH DOESN'T HAVE <10 RECORDED
AT46 | 20220217 | L11 | 7 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 61.35 | 37.24 | 0.73067 | 60.6193 | 36.5093 | 0.50696 | 0.13654 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR51
AT46 | 20220217 | L11 | 7 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 67.23 | 41.37 | 0.7595 | 66.4705 | 40.6105 | 0.54375 | 0.17203 | 20220519 | 11-40 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR52
AT46 | 20220217 | L11 | 7 | 14 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 110.2 | 62.66 | 0.93767 | 109.2623 | 61.7223 | 0.53059 | 0.046856 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR59 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L11 | 7 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 127.8 | 72.33 | 0.93767 | 126.8623 | 71.3923 | 0.62135 | 0.048995 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR60; higher than both 5 replicates | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220217 | L11 | 7 | 19 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AT46 | 20220217 | L11 | 7 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 105.9 | 68.25 | 0.7225 | 105.1775 | 67.5275 | 0.79985 | 0.40267 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L11 | 7 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 57.6 | 37.4 | 0.5395 | 57.0605 | 36.8605 | 0.42914 | 0.22727 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | filter wetter than normal going into extraction
AT46 | 20220217 | L11 | 7 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 59.3 | 38.35 | 0.5395 | 58.7605 | 37.8105 | 0.44507 | 0.22826 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220217 | L11 | 7 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.97 | 50.6 | 0.73067 | 80.2393 | 49.8693 | 0.63858 | 0.24039 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR49
AT46 | 20220217 | L11 | 7 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.46 | 41.98 | 0.73067 | 69.7293 | 41.2493 | 0.59884 | 0.1282 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR50
AT46 | 20220217 | L11 | 7 | 19 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 121.2 | 71.83 | 0.7595 | 120.4405 | 71.0705 | 0.55102 | 0.11389 | 20220519 | 11-40 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR57
AT46 | 20220217 | L11 | 7 | 19 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 109.9 | 63.67 | 0.93767 | 108.9623 | 62.7323 | 0.51785 | 0.071184 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR58 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220218 | L9 | 8 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.07 | 22.15 | 0.7225 | 30.3475 | 21.4275 | 0.1895 | 0.19208 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L9 | 8 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 27.75 | 19.25 | 0.7225 | 27.0275 | 18.5275 | 0.18058 | 0.14936 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L9 | 8 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 39.34 | 28.42 | 0.667 | 38.673 | 27.753 | 0.23199 | 0.26223 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L9 | 8 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.98 | 25.72 | 0.7225 | 34.2575 | 24.9975 | 0.19672 | 0.24843 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L9 | 8 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 41.45 | 28.82 | 0.667 | 40.783 | 28.153 | 0.26832 | 0.23303 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L9 | 8 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 41.14 | 29.2 | 0.667 | 40.473 | 28.533 | 0.25366 | 0.25445 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L9 | 8 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 105 | 68.26 | 0.667 | 104.333 | 67.593 | 0.78052 | 0.42317 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L9 | 8 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 115.8 | 72.51 | 0.7225 | 115.0775 | 71.7875 | 0.91967 | 0.35872 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L9 | 8 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 149.7 | 94.46 | 0.667 | 149.033 | 93.793 | 1.1735 | 0.49672 | 20220519 | 11-90 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L9 | 8 | 12 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76.27 | 49.78 | 0.7225 | 75.5475 | 49.0575 | 1.1255 | 0.62169 | 20220519 | 11-93 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L9 | 8 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 248.8 | 149.3 | 0.667 | 248.133 | 148.633 | 2.1138 | 0.53303 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L9 | 8 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 244 | 150.7 | 0.667 | 243.333 | 150.033 | 1.9821 | 0.68967 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L9 | 8 | 13 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 179.3 | 103.7 | 0.7595 | 178.5405 | 102.9405 | 1.5842 | 0.22399 | 20220519 | 11-40 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR67
AT46 | 20220218 | L9 | 8 | 13 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 199.3 | 115.8 | 0.7595 | 198.5405 | 115.0405 | 1.7557 | 0.27191 | 20220519 | 11-40 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR68
AT46 | 20220218 | L9 | 8 | 13 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 207.1 | 112.1 | 0.93767 | 206.1623 | 111.1623 | 1.0641 | -0.020378 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR73; neg phaeo | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220218 | L9 | 8 | 13 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 273.6 | 155.4 | 0.73067 | 272.8693 | 154.6693 | 1.324 | 0.12826 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR74
AT46 | 20220218 | L9 | 8 | 18 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 162.8 | 95.16 | 0.7225 | 162.0775 | 94.4375 | 2.8739 | 0.48953 | 20220519 | 11-94 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | Sosik log says N18 TAR log says N19 were they actually different N or does someone have typo?
AT46 | 20220218 | L9 | 8 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 290.4 | 172.8 | 0.667 | 289.733 | 172.133 | 2.4983 | 0.56699 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | Sosik log says N18 TAR log says N19 were they actually different N or does someone have typo?
AT46 | 20220218 | L9 | 8 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 244.4 | 133 | 0.7595 | 243.6405 | 132.2405 | 2.3424 | -0.011577 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR69;negative phaeo
AT46 | 20220218 | L9 | 8 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 213.3 | 120.7 | 0.7595 | 212.5405 | 119.9405 | 1.9471 | 0.16693 | 20220519 | 11-41 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR70
AT46 | 20220218 | L9 | 8 | 19 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 396.2 | 218.3 | 0.73067 | 395.4693 | 217.5693 | 1.9891 | 0.050042 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR75
AT46 | 20220218 | L9 | 8 | 19 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 351.7 | 201.1 | 0.73067 | 350.9693 | 200.3693 | 1.6839 | 0.19408 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR76
AT46 | 20220218 | L9 | 8 | 22 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 282.4 | 170.9 | 0.667 | 281.733 | 170.233 | 2.3687 | 0.66274 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L9 | 8 | 22 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 255.1 | 156.9 | 0.667 | 254.433 | 156.233 | 2.0862 | 0.69598 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L9 | 8 | 22 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 82.42 | 55.21 | 0.5395 | 81.8805 | 54.6705 | 0.57806 | 0.39551 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220218 | L9 | 8 | 22 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76 | 50.66 | 0.864 | 75.136 | 49.796 | 0.53833 | 0.34843 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220218 | L9 | 8 | 22 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 227.4 | 127.8 | 0.7595 | 226.6405 | 127.0405 | 2.0943 | 0.14489 | 20220519 | 11-40 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR65
AT46 | 20220218 | L9 | 8 | 22 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 246.7 | 139.4 | 0.7595 | 245.9405 | 138.6405 | 2.2562 | 0.18744 | 20220519 | 11-40 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR66
AT46 | 20220218 | L9 | 8 | 22 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 230.3 | 124.3 | 0.93767 | 229.3623 | 123.3623 | 1.1874 | -0.029042 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR71; neg phaeo | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220218 | L9 | 8 | 22 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 273.1 | 151.5 | 0.93767 | 272.1623 | 150.5623 | 1.3572 | 0.051424 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR72 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220218 | L10 | 9 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 102.4 | 65.52 | 0.7225 | 101.6775 | 64.7975 | 0.78349 | 0.37041 | 20220519 | 11-94 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L10 | 9 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 96.66 | 62.63 | 0.667 | 95.993 | 61.963 | 0.72294 | 0.38048 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L10 | 9 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 133.9 | 82.78 | 0.7225 | 133.1775 | 82.0575 | 1.086 | 0.37526 | 20220519 | 11-94 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L10 | 9 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 137.3 | 84.04 | 0.667 | 136.633 | 83.373 | 1.1315 | 0.35322 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L10 | 9 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 170.7 | 104.8 | 0.7225 | 169.9775 | 104.0775 | 1.4 | 0.4534 | 20220519 | 11-94 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L10 | 9 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 181.5 | 109.7 | 0.7225 | 180.7775 | 108.9775 | 1.5253 | 0.41531 | 20220519 | 11-94 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L10 | 9 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 54.04 | 37.26 | 0.5395 | 53.5005 | 36.7205 | 0.35648 | 0.29743 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220218 | L10 | 9 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 55.69 | 38.16 | 0.5395 | 55.1505 | 37.6205 | 0.37241 | 0.29753 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220218 | L10 | 9 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 121.7 | 68.61 | 0.93767 | 120.7623 | 67.6723 | 1.1163 | 0.076448 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR81; poor replication (76% range)-- were tubes switched during sampling? | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220218 | L10 | 9 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 53.74 | 30.02 | 0.93767 | 52.8023 | 29.0823 | 0.49876 | 0.013836 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR82; poor replication (76% range) were tubes switched during sampling? Or did something happen to this one? DF thinks this one is compromised given the values of the 20s. Assume other replicate is fine | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220218 | L10 | 9 | 12 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 199.6 | 112.9 | 0.93767 | 198.6623 | 111.9623 | 0.96941 | 0.079962 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR87 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220218 | L10 | 9 | 12 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 220.3 | 121.6 | 0.93767 | 219.3623 | 120.6623 | 1.1036 | 0.027328 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR88 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220218 | L10 | 9 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 235.1 | 145.8 | 0.667 | 234.433 | 145.133 | 1.8971 | 0.68739 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L10 | 9 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 229 | 139.9 | 0.667 | 228.333 | 139.233 | 1.8929 | 0.58657 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L10 | 9 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 83.76 | 56.75 | 0.864 | 82.896 | 55.886 | 0.57381 | 0.4214 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220218 | L10 | 9 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 74.45 | 50.31 | 0.5395 | 73.9105 | 49.7705 | 0.51284 | 0.37347 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220218 | L10 | 9 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 98.65 | 56.28 | 0.93767 | 97.7123 | 55.3423 | 0.89091 | 0.084532 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR79; poor replication (69%),should this be 85 value? No DF doesn't think so. Based on 20 values, toss this sample and assume other one is fine | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220218 | L10 | 9 | 16 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 197.1 | 109.8 | 0.93767 | 196.1623 | 108.8623 | 1.8356 | 0.083116 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR80; poor replication (69%) | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220218 | L10 | 9 | 16 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 328.5 | 187.6 | 0.63617 | 327.8638 | 186.9638 | 1.5783 | 0.17722 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR85; poor replication (33% range), is this really sample tube #79 value? No DF doesn't think so, thinks it's just within bio variation
AT46 | 20220218 | L10 | 9 | 16 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 219.5 | 119 | 0.93767 | 218.5623 | 118.0623 | 1.1258 | -0.017198 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR86; poor replication (33% range), DF thinks within bio variation but neg phaeo | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220218 | L10 | 9 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 272.9 | 166.5 | 0.667 | 272.233 | 165.833 | 2.2604 | 0.69274 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L10 | 9 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 248.9 | 152.3 | 0.667 | 248.233 | 151.633 | 2.0522 | 0.64806 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220218 | L10 | 9 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 107.4 | 69.02 | 0.864 | 106.536 | 68.156 | 0.81536 | 0.39836 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220218 | L10 | 9 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 96.04 | 61.39 | 0.864 | 95.176 | 60.526 | 0.73612 | 0.34172 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220218 | L10 | 9 | 20 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 168.1 | 95.06 | 0.93767 | 167.1623 | 94.1223 | 1.5358 | 0.12316 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR76-ID# SHOULD BE TAR77, MIGHT BE 2 TUBES LABELED TAR76 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220218 | L10 | 9 | 20 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 218.5 | 121.3 | 0.93767 | 217.5623 | 120.3623 | 2.0438 | 0.077644 | 20220519 | 11-43 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR78 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220218 | L10 | 9 | 20 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 302.9 | 173.5 | 0.63617 | 302.2638 | 172.8638 | 1.4495 | 0.17364 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR83; poor replication (40% range) but within bio variation
AT46 | 20220218 | L10 | 9 | 20 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 200.6 | 114.2 | 0.63617 | 199.9638 | 113.5638 | 0.96431 | 0.098148 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR84; poor replication (40% range) but within bio variation
AT46 | 20220219 | L8 | 13 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.45 | 22.97 | 0.667 | 30.783 | 22.303 | 0.18015 | 0.21702 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L8 | 13 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.79 | 22.5 | 0.7225 | 31.0675 | 21.7775 | 0.19736 | 0.19045 | 20220519 | 11-94 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L8 | 13 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.55 | 25.21 | 0.667 | 35.883 | 24.543 | 0.24091 | 0.19615 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L8 | 13 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.23 | 22.38 | 0.667 | 30.563 | 21.713 | 0.18801 | 0.19865 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L8 | 13 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.13 | 13.72 | 0.864 | 18.266 | 12.856 | 0.11493 | 0.11401 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220219 | L8 | 13 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23.34 | 16.78 | 0.864 | 22.476 | 15.916 | 0.13936 | 0.14407 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220219 | L8 | 13 | 6 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22.14 | 14.61 | 0.73067 | 21.4093 | 13.8793 | 0.15779 | 0.086005 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR95
AT46 | 20220219 | L8 | 13 | 6 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.55 | 14.91 | 0.73067 | 19.8193 | 14.1793 | 0.11819 | 0.13088 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR96
AT46 | 20220219 | L8 | 13 | 6 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 32.1 | 19.1 | 0.93767 | 31.1623 | 18.1623 | 0.14536 | 0.024872 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR103 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220219 | L8 | 13 | 6 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 29.12 | 18.57 | 0.73067 | 28.3893 | 17.8393 | 0.11775 | 0.049149 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR104
AT46 | 20220219 | L8 | 13 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 69.6 | 45.1 | 0.667 | 68.933 | 44.433 | 0.52049 | 0.27077 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L8 | 13 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.77 | 42.31 | 0.667 | 63.103 | 41.643 | 0.4559 | 0.28567 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L8 | 13 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 29.24 | 20.7 | 0.5395 | 28.7005 | 20.1605 | 0.18143 | 0.17759 | 20220519 | 13-05 | 20230220 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220219 | L8 | 13 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.13 | 21.65 | 0.864 | 30.266 | 20.786 | 0.2014 | 0.16876 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220219 | L8 | 13 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.28 | 21.43 | 0.93767 | 35.3423 | 20.4923 | 0.31225 | 0.04894 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR93 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220219 | L8 | 13 | 10 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 32.33 | 19.44 | 0.93767 | 31.3923 | 18.5023 | 0.27104 | 0.055078 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR94 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220219 | L8 | 13 | 10 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 37.03 | 20.91 | 0.93767 | 36.0923 | 19.9723 | 0.18024 | 0.0069506 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR101; poor replication (57% range) but within bio variation | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220219 | L8 | 13 | 10 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.75 | 11.81 | 0.93767 | 19.8123 | 10.8723 | 0.09996 | 0.0019415 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR102; poor replication (57% range) but within bio variation | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220219 | L8 | 13 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 148.8 | 90.75 | 0.667 | 148.133 | 90.083 | 1.2332 | 0.37095 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L8 | 13 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 157.2 | 97.81 | 0.667 | 156.533 | 97.143 | 1.2617 | 0.46821 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L8 | 13 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 102.4 | 61.39 | 0.73067 | 101.6693 | 60.6593 | 0.86231 | 0.20684 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR91
AT46 | 20220219 | L8 | 13 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 97.16 | 55.58 | 0.93767 | 96.2223 | 54.6423 | 0.87132 | 0.088503 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR92 | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220219 | L8 | 13 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 150.6 | 86.82 | 0.73067 | 149.8693 | 86.0893 | 0.71444 | 0.093909 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR99
AT46 | 20220219 | L8 | 13 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 153 | 88.55 | 0.73067 | 152.2693 | 87.8193 | 0.72194 | 0.10265 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR100
AT46 | 20220219 | L8 | 13 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 246.4 | 146.3 | 0.667 | 245.733 | 145.633 | 2.1266 | 0.46686 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L8 | 13 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 206.4 | 127.4 | 0.667 | 205.733 | 126.733 | 1.6783 | 0.57854 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L8 | 13 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.32 | 46.79 | 0.864 | 69.456 | 45.926 | 0.49988 | 0.31797 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220219 | L8 | 13 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 67.01 | 45.14 | 0.864 | 66.146 | 44.276 | 0.46461 | 0.32385 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220219 | L8 | 13 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 173 | 101.5 | 0.73067 | 172.2693 | 100.7693 | 1.5034 | 0.2727 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR89
AT46 | 20220219 | L8 | 13 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 131.3 | 77.53 | 0.73067 | 130.5693 | 76.7993 | 1.1306 | 0.22302 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR90
AT46 | 20220219 | L8 | 13 | 19 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 214.7 | 107 | 0.73067 | 213.9693 | 106.2693 | 1.202 | -0.20783 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR97; neg phaeo
AT46 | 20220219 | L8 | 13 | 19 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 241.1 | 140.4 | 0.73067 | 240.3693 | 139.6693 | 1.128 | 0.18344 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR98
AT46 | 20220219 | L7 | 14 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 68.39 | 43.68 | 0.667 | 67.723 | 43.013 | 0.52495 | 0.24102 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | poor replication (43% range)
AT46 | 20220219 | L7 | 14 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.65 | 33.73 | 0.667 | 48.983 | 33.063 | 0.33821 | 0.25057 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | poor replication (43% range)
AT46 | 20220219 | L7 | 14 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 145.1 | 88.41 | 0.667 | 144.433 | 87.743 | 1.2043 | 0.35817 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L7 | 14 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 133.3 | 83.59 | 0.667 | 132.633 | 82.923 | 1.0561 | 0.42063 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L7 | 14 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 170.4 | 104.3 | 0.667 | 169.733 | 103.633 | 1.4043 | 0.44123 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L7 | 14 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 171.4 | 106.6 | 0.667 | 170.733 | 105.933 | 1.3766 | 0.50981 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L7 | 14 | 11 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 143.6 | 84.05 | 0.93767 | 142.6623 | 83.1123 | 1.2521 | 0.21275 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR109; poor replication (87% range) | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220219 | L7 | 14 | 11 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 55.72 | 32.29 | 0.93767 | 54.7823 | 31.3523 | 0.49098 | 0.05974 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR110; poor replication (87% range), based on 20 values, something is clearly up with this sample. Toss and assume other rep is fine | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220219 | L7 | 14 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 224.5 | 131.8 | 0.73067 | 223.7693 | 131.0693 | 1.0384 | 0.1923 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR115
AT46 | 20220219 | L7 | 14 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 168.6 | 96.4 | 0.73067 | 167.8693 | 95.6693 | 0.80875 | 0.089545 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR116
AT46 | 20220219 | L7 | 14 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 181.1 | 112 | 0.667 | 180.433 | 111.333 | 1.468 | 0.51462 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L7 | 14 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 176.5 | 109.6 | 0.667 | 175.833 | 108.933 | 1.4212 | 0.51862 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L7 | 14 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 216 | 131.9 | 0.667 | 215.333 | 131.233 | 1.7866 | 0.55033 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L7 | 14 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 218.4 | 136.8 | 0.667 | 217.733 | 136.133 | 1.7335 | 0.6907 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L7 | 14 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 46.62 | 26.55 | 0.93767 | 45.6823 | 25.6123 | 0.42201 | 0.029423 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR107; poor replication (72% range). Something is up with samples from N15 and N21 from this cast | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220219 | L7 | 14 | 15 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 94.19 | 50.97 | 0.93767 | 93.2523 | 50.0323 | 0.90569 | -0.026841 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR108; neg phaeo, poor replication (72% range). Something is up with samples from N15 and N21 from this cast | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220219 | L7 | 14 | 15 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 111.2 | 61.88 | 0.73067 | 110.4693 | 61.1493 | 0.55246 | 0.021707 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR113; poor replication (75% range). Something is up with samples from N15 and N21 from this cast
AT46 | 20220219 | L7 | 14 | 15 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 245.7 | 137.6 | 0.73067 | 244.9693 | 136.8693 | 1.2109 | 0.07426 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR114; poor replication (75% range). Something is up with samples from N15 and N21 from this cast
AT46 | 20220219 | L7 | 14 | 21 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 193.2 | 117.7 | 0.667 | 192.533 | 117.033 | 1.6039 | 0.48016 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L7 | 14 | 21 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 211.1 | 131.6 | 0.667 | 210.433 | 130.933 | 1.6889 | 0.64271 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L7 | 14 | 21 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 69.58 | 45.82 | 0.864 | 68.716 | 44.956 | 0.50477 | 0.29581 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | not the best replication
AT46 | 20220219 | L7 | 14 | 21 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 61.37 | 42.34 | 0.864 | 60.506 | 41.476 | 0.40428 | 0.33432 | 20220519 | 13-03 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | not the best replication
AT46 | 20220219 | L7 | 14 | 21 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 159.5 | 88.58 | 0.93767 | 158.5623 | 87.6423 | 1.4912 | 0.053521 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC after DF2 | TAR105; poor replication (77% range). Something is up with samples from N15 and N21 from this cast | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220219 | L7 | 14 | 21 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 74.7 | 43.13 | 0.93767 | 73.7623 | 42.1923 | 0.66382 | 0.079846 | 20220519 | 11-44 | 20220817 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | TC after DF2 | TAR106; poor replication (77% range). Something is up with samples from N15 and N21 from this cast | DF messed up and didn't run the blanks after sample running so only 3 values for the acetone blank
AT46 | 20220219 | L7 | 14 | 21 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 257.3 | 146.5 | 0.73067 | 256.5693 | 145.7693 | 1.2366 | 0.12712 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | TC after DF2 | TAR111; poor replication (62% range). Something is up with samples from N15 and N21 from this cast
AT46 | 20220219 | L7 | 14 | 21 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 134.8 | 76.9 | 0.73067 | 134.0693 | 76.1693 | 0.64857 | 0.06663 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC after DF2 | TAR112;poor replication (62% range). Something is up with samples from N15 and N21 from this cast
AT46 | 20220219 | L4 | 15 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 177.8 | 108.5 | 0.667 | 177.133 | 107.833 | 1.4722 | 0.44804 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L4 | 15 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 178 | 110.7 | 0.667 | 177.333 | 110.033 | 1.4297 | 0.52971 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L4 | 15 | 1 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 101.7 | 60.68 | 0.73067 | 100.9693 | 59.9493 | 0.86252 | 0.19412 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR119; 5 values are smaller than the 20 values
AT46 | 20220219 | L4 | 15 | 1 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 124.2 | 72.03 | 0.63617 | 123.5638 | 71.3938 | 1.097 | 0.16139 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR120; 5 values are smaller than the 20 valules
AT46 | 20220219 | L4 | 15 | 1 |  |  | a | 449 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 189.5 | 109.5 | 0.63617 | 188.8638 | 108.8638 | 1.0977 | 0.15442 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR123; 20 values are larger than the 5 values
AT46 | 20220219 | L4 | 15 | 1 |  |  | b | 449 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 188.4 | 104.5 | 0.63617 | 187.7638 | 103.8638 | 1.1512 | 0.043399 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR124; 20 values are larger than the 5 values
AT46 | 20220219 | L4 | 15 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 195.4 | 117.6 | 0.667 | 194.733 | 116.933 | 1.6528 | 0.42952 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L4 | 15 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 176.6 | 109.7 | 0.667 | 175.933 | 109.033 | 1.4212 | 0.5204 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L4 | 15 | 3 |  |  | a | 290 | 5 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20220519 | [5 empty] | Turner Designs Aquafluor Handheld Serial Number 800446 | 9 | DF2 | TAR117; might be missing?
AT46 | 20220219 | L4 | 15 | 3 |  |  | b | 290 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 118.9 | 68.5 | 0.63617 | 118.2638 | 67.8638 | 1.0707 | 0.1378 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR118
AT46 | 20220219 | L4 | 15 | 3 |  |  | a | 548 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 191.2 | 110 | 0.63617 | 190.5638 | 109.3638 | 0.91289 | 0.11774 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR121
AT46 | 20220219 | L4 | 15 | 3 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 181.9 | 104.2 | 0.63617 | 181.2638 | 103.5638 | 0.87036 | 0.10206 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR122
AT46 | 20220219 | L3 | 16 | 1 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 211.6 | 126.4 | 0.667 | 210.933 | 125.733 | 3.62 | 0.85804 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L3 | 16 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 145.1 | 87.49 | 0.667 | 144.433 | 86.823 | 3.6717 | 0.96674 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L3 | 16 | 4 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 194.8 | 117.6 | 0.667 | 194.133 | 116.933 | 3.2801 | 0.88453 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DFthenTC | somewhat poor replication (14% range) but also large chl values, could just be natural variation?
AT46 | 20220219 | L3 | 16 | 4 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 222.2 | 132.7 | 0.667 | 221.533 | 132.033 | 3.8027 | 0.89972 | 20220519 | 11-91 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DFthenTC | higher than WSW; somewhat poor replication (14% range) but also large chl values, could just be natural variation?
AT46 | 20220219 | L3 | 16 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 376.5 | 225.7 | 0.667 | 375.833 | 225.033 | 3.2036 | 0.80371 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L3 | 16 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 402.8 | 247.4 | 0.667 | 402.133 | 246.733 | 3.3014 | 1.0924 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L3 | 16 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.9 | 50.61 | 0.864 | 80.036 | 49.746 | 0.64349 | 0.24238 | 20220519 | 13-04 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | sample never recorded on cruis logsheet but samples exist. Assume vol filt 290
AT46 | 20220219 | L3 | 16 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 77.79 | 50.11 | 0.864 | 76.926 | 49.246 | 0.58804 | 0.28892 | 20220519 | 13-04 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | sample never recorded on cruis logsheet but samples exist. Assume vol filt 290
AT46 | 20220219 | L3 | 16 | 7 |  |  | a | 290 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 353.2 | 207.1 | 0.73067 | 352.4693 | 206.3693 | 3.1038 | 0.5712 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR129
AT46 | 20220219 | L3 | 16 | 7 |  |  | b | 290 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 326.7 | 189.2 | 0.73067 | 325.9693 | 188.4693 | 2.9211 | 0.43514 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR130
AT46 | 20220219 | L3 | 16 | 7 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 187.1 | 106.1 | 0.63617 | 186.4638 | 105.4638 | 2.722 | 0.24881 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR135
AT46 | 20220219 | L3 | 16 | 7 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 183.8 | 103.1 | 0.63617 | 183.1638 | 102.4638 | 2.7119 | 0.17439 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR136
AT46 | 20220219 | L3 | 16 | 10 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 207.3 | 126.4 | 0.667 | 206.633 | 125.733 | 3.4373 | 1.0407 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L3 | 16 | 10 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 216.7 | 129.8 | 0.667 | 216.033 | 129.133 | 3.6923 | 0.9069 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220219 | L3 | 16 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 389.2 | 229 | 0.63617 | 388.5638 | 228.3638 | 3.3685 | 0.65654 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR127
AT46 | 20220219 | L3 | 16 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 350.5 | 203.5 | 0.6362 | 349.8638 | 202.8638 | 3.0804 | 0.48299 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR128
AT46 | 20220219 | L3 | 16 | 10 |  |  | a | 448 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 225.5 | 125.6 | 0.6362 | 224.8638 | 124.9638 | 4.1214 | 0.20008 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR133; poor replication (36% range); this sample is higher than both WSW and 5 values so something is up, was the dilution factor different? DF thinks it should be tossed
AT46 | 20220219 | L3 | 16 | 10 |  |  | b | 449 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 162.7 | 93.03 | 0.6362 | 162.0638 | 92.3938 | 2.8679 | 0.32018 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR134;poor replication (36% range)
AT46 | 20220219 | L3 | 16 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 391 | 235.2 | 0.667 | 390.333 | 234.533 | 3.3099 | 0.86667 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | N12 ran out of water so WSW from N12 but other sfrom N13
AT46 | 20220219 | L3 | 16 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 142.6 | 83.72 | 0.667 | 141.933 | 83.053 | 3.7526 | 0.68439 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | N12 ran out of water so WSW from N12 but other sfrom N13
AT46 | 20220219 | L3 | 16 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.26 | 52.39 | 0.864 | 79.396 | 51.526 | 0.59208 | 0.32549 | 20220519 | 13-04 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220219 | L3 | 16 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 67.85 | 43.14 | 0.864 | 66.986 | 42.276 | 0.52495 | 0.2279 | 20220519 | 13-04 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AT46 | 20220219 | L3 | 16 | 13 |  |  | a | 290 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 334.8 | 189 | 0.63617 | 334.1638 | 188.3638 | 3.0974 | 0.25693 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR125; for this set, the 20s are higher than the 5s -- should they be switched? DF thinks yes they should be switched
AT46 | 20220219 | L3 | 16 | 13 |  |  | b | 290 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 372.1 | 214.2 | 0.73067 | 371.3693 | 213.4693 | 3.3545 | 0.44695 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR126
AT46 | 20220219 | L3 | 16 | 13 |  |  | a | 448 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 563.5 | 323.9 | 0.63617 | 562.8638 | 323.2638 | 3.295 | 0.43144 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR131; larger than either 5 rep; dilution and not dil RbRa available. Non dil makes more sense so use that one.
AT46 | 20220219 | L3 | 16 | 13 |  |  | b | 448 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 207.9 | 121.6 | 0.63617 | 207.2638 | 120.9638 | 3.5604 | 0.62283 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR132; larger than either 5 rep
AT46 | 20220220 | L2 | 17 | 2 |  |  | a | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 291.4 | 178.1 | 0.667 | 290.733 | 177.433 | 7.2209 | 2.2582 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | seems like this whole cast is pretty variable
AT46 | 20220220 | L2 | 17 | 2 |  |  | b | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 283.1 | 174.3 | 0.667 | 282.433 | 173.633 | 6.9341 | 2.342 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220220 | L2 | 17 | 4 |  |  | a | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 196.6 | 121.9 | 0.667 | 195.933 | 121.233 | 4.7609 | 1.7158 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (33% range), seems like maybe just a variable cast?
AT46 | 20220220 | L2 | 17 | 4 |  |  | b | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 272.4 | 167.3 | 0.667 | 271.733 | 166.633 | 6.6983 | 2.2038 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | poor replication (33% range), seems like maybe just a variable cast?
AT46 | 20220220 | L2 | 17 | 4 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 101.3 | 68.85 | 0.864 | 100.436 | 67.986 | 0.68938 | 0.52131 | 20220519 | 13-04 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC | when taking readings might have been mix up with test tubes. Not 100% cofident these are both correct but values somewhat reasonable for <10s sample never recorded on cruis logsheet but samples exist. Assume vol filt 290
AT46 | 20220220 | L2 | 17 | 4 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 68.82 | 45.06 | 0.864 | 67.956 | 44.196 | 0.50477 | 0.28227 | 20220519 | 13-04 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | TC | when taking readings might have been mix up with test tubes. Not 100% cofident these are both correct but values somewhat reasonable for <10s sample never recorded on cruis logsheet but samples exist. Assume vol filt 290
AT46 | 20220220 | L2 | 17 | 4 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.82 | 10.97 | 0.73067 | 16.0893 | 10.2393 | 0.12301 | 0.057467 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR141; throw out, clearly something wrong with sample; reps are way off -- is this the right sample?
AT46 | 20220220 | L2 | 17 | 4 |  |  | b | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 270.1 | 157.9 | 0.63617 | 269.4638 | 157.2638 | 7.0776 | 1.238 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR142; reps are way off, is this the right sample?
AT46 | 20220220 | L2 | 17 | 4 |  |  | a | 371 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 190.1 | 109.5 | 0.73067 | 189.3693 | 108.7693 | 4.0154 | 0.52682 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR147; filtered long time no progress
AT46 | 20220220 | L2 | 17 | 4 |  |  | b | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 325.1 | 182.7 | 0.73067 | 324.3693 | 181.9693 | 4.7766 | 0.33994 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR148
AT46 | 20220220 | L2 | 17 | 9 |  |  | a | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 283.4 | 175.1 | 0.667 | 282.733 | 174.433 | 6.9023 | 2.4166 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220220 | L2 | 17 | 9 |  |  | b | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 285.4 | 173.8 | 0.667 | 284.733 | 173.133 | 7.1126 | 2.1368 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220220 | L2 | 17 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 108.1 | 74.59 | 0.864 | 107.236 | 73.726 | 0.7119 | 0.60101 | 20220519 | 13-04 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | very poor replication
AT46 | 20220220 | L2 | 17 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.6 | 32.34 | 0.864 | 48.736 | 31.476 | 0.36668 | 0.19384 | 20220519 | 13-04 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | very poor replication. Also this value very low relative to high WSW
AT46 | 20220220 | L2 | 17 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 278.7 | 165.8 | 0.63617 | 278.0638 | 165.1638 | 7.1218 | 1.6115 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR139;arger than either WSW rep
AT46 | 20220220 | L2 | 17 | 9 |  |  | b | 294 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 353.7 | 207.9 | 0.63617 | 353.0638 | 207.2638 | 9.1658 | 1.7563 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR140;larger than either WSW rep
AT46 | 20220220 | L2 | 17 | 9 |  |  | a | 500 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 90.46 | 52.97 | 0.63617 | 89.8238 | 52.3338 | 1.3858 | 0.23578 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR145; much lower than 20s at other depths of this cast
AT46 | 20220220 | L2 | 17 | 9 |  |  | b | 460 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 340.9 | 207.3 | 0.73067 | 340.1693 | 206.5693 | 1.7893 | 0.52976 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR146; filtered long time no progress; much lower than 20s at other depths of this cast
AT46 | 20220220 | L2 | 17 | 14 |  |  | a | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 248.5 | 152.1 | 0.667 | 247.833 | 151.433 | 6.1439 | 1.9462 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220220 | L2 | 17 | 14 |  |  | b | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 269.1 | 164.8 | 0.667 | 268.433 | 164.133 | 6.6474 | 2.1212 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220220 | L2 | 17 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 54.49 | 36.99 | 0.5395 | 53.9505 | 36.4505 | 0.37178 | 0.27733 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replication
AT46 | 20220220 | L2 | 17 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.02 | 52.94 | 0.864 | 79.156 | 52.076 | 0.5753 | 0.35207 | 20220519 | 13-04 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replication
AT46 | 20220220 | L2 | 17 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 194.5 | 111.4 | 0.63617 | 193.8638 | 110.7638 | 5.242 | 0.61483 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR137
AT46 | 20220220 | L2 | 17 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 174.2 | 108.3 | 0.63617 | 173.5638 | 107.6638 | 4.157 | 1.5359 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR138
AT46 | 20220220 | L2 | 17 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 398.6 | 240.5 | 0.73067 | 397.8693 | 239.7693 | 5.3129 | 1.4411 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR143; poor replication (54% range)
AT46 | 20220220 | L2 | 17 | 14 |  |  | b | 524 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 183.5 | 97.09 | 0.73067 | 182.7693 | 96.3593 | 3.0479 | -0.19885 | 20220519 | 11-46 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR144; filtered long time no progress; poor replication (54% range); neg phaeo and much lower concen
AT46 | 20220221 | MVCO | 23 | 2 |  |  | a | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 229.4 | 153.4 | 0.667 | 228.733 | 152.733 | 4.8437 | 3.3158 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220221 | MVCO | 23 | 2 |  |  | b | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 214.4 | 145.5 | 0.667 | 213.733 | 144.833 | 4.3912 | 3.3463 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220221 | MVCO | 23 | 3 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 238.7 | 151.9 | 0.667 | 238.033 | 151.233 | 3.688 | 1.6983 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220221 | MVCO | 23 | 3 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 235.1 | 152.3 | 0.667 | 234.433 | 151.633 | 3.5181 | 1.8825 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220221 | MVCO | 23 | 7 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 234.7 | 153.5 | 0.667 | 234.033 | 152.833 | 3.4501 | 1.9932 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220221 | MVCO | 23 | 7 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 236.4 | 156 | 0.667 | 235.733 | 155.333 | 3.4161 | 2.1162 | 20220519 | 11-92 | 20221205 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
AT46 | 20220221 | MVCO | 23 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 59.44 | 42.61 | 0.864 | 58.576 | 41.746 | 0.35754 | 0.38587 | 20220519 | 13-04 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | very poor replication; <10a half of <10b could vol extracted or dilution be wrong?
AT46 | 20220221 | MVCO | 23 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 96.89 | 66.87 | 0.864 | 96.026 | 66.006 | 0.63775 | 0.53767 | 20220519 | 13-04 | 20230217 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | very poor replication; <10a half of <10b could vol extracted or dilution be wrong?
AR66A | 20220409 |  | 1 | 9 |  |  | a | 1057 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 340.6 | 217.4 | 0.62 | 339.98 | 216.78 | 2.1543 | 1.0232 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 01/01; Gliders 525 and 538
AR66A | 20220409 |  | 1 | 10 |  |  | a | 1056 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 372.3 | 236.7 | 0.62 | 371.68 | 236.08 | 2.3733 | 1.0903 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 01/02; Gliders 525 and 538
AR66A | 20220409 | L10 | 2 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 115.6 | 73.23 | 0.62 | 114.98 | 72.61 | 1.4832 | 0.6474 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 02/01 | poor replication between replicates/rep depths but maybe samples were accidentally switche. N7a & N9 pair and N7b & N8 pair
AR66A | 20220409 | L10 | 2 | 7 |  |  | b | 1056 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 326.4 | 206.5 | 0.62 | 325.78 | 205.88 | 2.0985 | 0.92198 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 02/02 | poor replication between replicates/rep depths but maybe samples were accidentally switche. N7a & N9 pair and N7b & N8 pair
AR66A | 20220409 | L10 | 2 | 8 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 149.8 | 94.71 | 0.62 | 149.18 | 94.09 | 1.9266 | 0.83163 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 02/03 | poor replication between replicates/rep depths but maybe samples were accidentally switche. N7a & N9 pair and N7b & N8 pair
AR66A | 20220409 | L10 | 2 | 9 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 389.4 | 248.1 | 0.62 | 388.78 | 247.48 | 1.6487 | 0.77183 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 02/04 | poor replication between replicates/rep depths but maybe samples were accidentally switche. N7a & N9 pair and N7b & N8 pair
AR66A | 20220409 |  | 3 | 2 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 183.3 | 133.9 | 0.62 | 182.68 | 133.28 | 0.28793 | 0.36324 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 03/01; Glider 339 | phae>chl; depth 84m
AR66A | 20220409 |  | 3 | 3 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 307.9 | 211.8 | 0.62 | 307.28 | 211.18 | 0.56066 | 0.4721 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 03/02; Glider 339
AR66A | 20220409 |  | 3 | 4 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 182.6 | 121.8 | 0.62 | 181.98 | 121.18 | 0.70943 | 0.47581 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 03/02; Glider 339
AR66A | 20220410 | L8 | 4 | 5 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 362.6 | 232.5 | 0.62 | 361.98 | 231.88 | 1.5166 | 0.74923 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/01
AR66A | 20220410 | L8 | 4 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 352.8 | 225.8 | 0.62 | 352.18 | 225.18 | 1.4819 | 0.72058 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/02
AR66A | 20220410 | L8 | 4 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 344 | 216.1 | 0.62 | 343.38 | 215.48 | 1.4924 | 0.6152 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/03
AR66A | 20220410 | L8 | 4 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 367.4 | 230.7 | 0.62 | 366.78 | 230.08 | 1.5951 | 0.65532 | 20220519 | 13-25 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/04
AR66A | 20220410 |  | 5 | 4 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 167.4 | 106.5 | 0.62 | 166.78 | 105.88 | 2.1298 | 0.97407 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/01; Gliders 388 and 380
AR66A | 20220410 |  | 5 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 164.8 | 104 | 0.62 | 164.18 | 103.38 | 2.1283 | 0.90513 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 05/02; Gliders 388 and 380
AR66A | 20220411 | L6 | 6 | 3 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 240.7 | 155.2 | 0.62 | 240.08 | 154.58 | 0.99669 | 0.5138 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/01
AR66A | 20220411 | L6 | 6 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 302 | 193.3 | 0.62 | 301.38 | 192.68 | 1.2683 | 0.61623 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/02
AR66A | 20220411 | L6 | 6 | 6 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 299.3 | 191.3 | 0.62 | 298.68 | 190.68 | 1.2602 | 0.60483 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/03
AR66A | 20220411 | L6 | 6 | 7 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 283.2 | 181.3 | 0.62 | 282.58 | 180.68 | 1.189 | 0.5782 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/04
AR66A | 20220411 | L6 | 6 | 8 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 312.1 | 199.7 | 0.62 | 311.48 | 199.08 | 1.3103 | 0.63505 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/05
AR66A | 20220412 |  | 7 | 4 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 351.6 | 222.6 | 0.62 | 350.98 | 221.98 | 1.5038 | 0.66531 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/01; AUV deploy AC-1
AR66A | 20220412 |  | 7 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 377.2 | 239.6 | 0.62 | 376.58 | 238.98 | 1.6056 | 0.73187 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 07/02; AUV deploy AC-1 flagged caution poor replication
AR66A | 20220412 |  | 7 | 5 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 405.8 | 253.1 | 0.62 | 405.18 | 252.48 | 1.7817 | 0.68772 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 07/03; AUV deploy AC-1 flagged caution poor replication
AR66A | 20220413 |  | 8 | 4 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 310.4 | 196.9 | 0.62 | 309.78 | 196.28 | 1.3231 | 0.59487 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 08/01; AUV recovery AC-4
AR66A | 20220413 |  | 8 | 5 |  |  | a | 1056 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 384.7 | 244.9 | 0.62 | 384.08 | 244.28 | 2.4468 | 1.1371 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 08/02; AUV recovery AC-4
AR66A | 20220413 |  | 8 | 5 |  |  | b | 1056 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 375.5 | 236.8 | 0.62 | 374.88 | 236.18 | 2.4276 | 1.0375 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 08/03; AUV recovery AC-4
AR66A | 20220416 |  | 19 | 3 |  |  | a | 1057 | 0 | 15 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 163.2 | 99.22 | 0.62 | 162.58 | 98.6 | 3.3562 | 0.97941 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 19/01; AUV deploy AL-1
AR66A | 20220416 |  | 19 | 4 |  |  | a | 1056 | 0 | 15 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 184.9 | 113.2 | 0.62 | 184.28 | 112.58 | 3.7648 | 1.1903 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 19/02; AUV deploy AL-1; extremely poor replication of N4 one should probably be tossed
AR66A | 20220416 |  | 19 | 4 |  |  | b | 1056 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 346 | 211.9 | 0.62 | 345.38 | 211.28 | 2.3471 | 0.75267 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 19/03; AUV deploy AL-1; extremely poor replication of N4 one should probably be tossed
AR66A | 20220416 |  | 20 | 3 |  |  | a | 1056 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 309.6 | 196.8 | 0.62 | 308.98 | 196.18 | 1.9743 | 0.90393 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 20/01; AUV recovery AL-2
AR66A | 20220416 |  | 20 | 4 |  |  | a | 1057 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 347.9 | 218.7 | 0.62 | 347.28 | 218.08 | 2.2592 | 0.93731 | 20220519 | 13-26 | 20230330 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 20/02; AUV recovery AL-2
AR66B | 20220420 | L2 | 1 | 2 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 142.1 | 103.5 | 0.71583 | 141.3842 | 102.7842 | 0.83442 | 1.0281 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  |  | phae > chl; depth 41m
AR66B | 20220420 | L2 | 1 | 2 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 161.8 | 118.2 | 0.6885 | 161.1115 | 117.5115 | 0.9425 | 1.1868 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  |  | phae > chl; depth 41m
AR66B | 20220420 | L2 | 1 | 4 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 116 | 78.61 | 0.63933 | 115.3607 | 77.9707 | 0.80826 | 0.60459 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L2 | 1 | 4 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 112.3 | 77.55 | 0.71583 | 111.5842 | 76.8342 | 0.75119 | 0.64106 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L2 | 1 | 6 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 131 | 89.07 | 0.71583 | 130.2842 | 88.3542 | 0.9064 | 0.6946 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L2 | 1 | 6 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 129 | 87.36 | 0.71583 | 128.2842 | 86.6442 | 0.90013 | 0.66988 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L2 | 1 | 9 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 128.8 | 87.82 | 0.71583 | 128.0842 | 87.1042 | 0.88587 | 0.69249 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L2 | 1 | 9 |  |  | b | 285 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 136.7 | 93.13 | 0.71583 | 135.9842 | 92.4142 | 0.94185 | 0.73272 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L2 | 1 | 9 |  |  | a | 285 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.97 | 34 | 0.71583 | 49.2542 | 33.2842 | 0.34522 | 0.25789 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L2 | 1 | 9 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.4 | 35.39 | 0.6885 | 51.7115 | 34.7015 | 0.36771 | 0.2611 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L3 | 2 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 186.5 | 114.3 | 0.71583 | 185.7842 | 113.5842 | 1.5338 | 0.48885 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L3 | 2 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 172.5 | 106.6 | 0.63933 | 171.8607 | 105.9607 | 1.4 | 0.48693 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L3 | 2 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 183.2 | 113 | 0.6885 | 182.5115 | 112.3115 | 1.4914 | 0.50868 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L3 | 2 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 183.6 | 113.5 | 0.71583 | 182.8842 | 112.7842 | 1.4892 | 0.51922 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L3 | 2 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 208.8 | 128.1 | 0.71583 | 208.0842 | 127.3842 | 1.7144 | 0.55402 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L3 | 2 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 187.3 | 116.7 | 0.71583 | 186.5842 | 115.9842 | 1.4998 | 0.56558 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L3 | 2 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 202.2 | 124.7 | 0.63933 | 201.5607 | 124.0607 | 1.6464 | 0.56282 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | Tissue capsule says C2N13
AR66B | 20220420 | L3 | 2 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 179.8 | 111 | 0.71583 | 179.0842 | 110.2842 | 1.4616 | 0.50232 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | Tissue capsule says C2N13
AR66B | 20220420 | L3 | 2 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 106.9 | 67.89 | 0.71583 | 106.1842 | 67.1742 | 0.82874 | 0.36749 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | C2N12<10a dropped on floor | Tissue capsule says C2N13
AR66B | 20220420 | L3 | 2 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 106.6 | 67.73 | 0.71583 | 105.8842 | 67.0142 | 0.82577 | 0.36761 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | Tissue capsule says C2N13
AR66B | 20220420 | L4 | 3 | 2 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 135.2 | 80.65 | 0.71583 | 134.4842 | 79.9342 | 2.3178 | 0.52916 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | lots of big cells seen on many filters
AR66B | 20220420 | L4 | 3 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 135.3 | 81.31 | 0.71583 | 134.5842 | 80.5942 | 2.294 | 0.57646 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | lots of big cells seen on many filters
AR66B | 20220420 | L4 | 3 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 316.6 | 188.5 | 0.63933 | 315.9607 | 187.8607 | 2.7214 | 0.624 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | lots of big cells seen on many filters
AR66B | 20220420 | L4 | 3 | 4 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 202.7 | 120.5 | 0.71583 | 201.9842 | 119.7842 | 3.4926 | 0.77363 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | lots of big cells seen on many filters
AR66B | 20220420 | L4 | 3 | 6 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 172.5 | 102.5 | 0.71583 | 171.7842 | 101.7842 | 2.9742 | 0.65091 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | lots of big cells seen on many filters
AR66B | 20220420 | L4 | 3 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 323.9 | 194.3 | 0.6885 | 323.2115 | 193.6115 | 2.7533 | 0.69454 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | lots of big cells seen on many filters
AR66B | 20220420 | L4 | 3 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 267.7 | 159.3 | 0.6885 | 267.0115 | 158.6115 | 2.3029 | 0.52165 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | lots of big cells seen on many filters
AR66B | 20220420 | L4 | 3 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 270.1 | 162.2 | 0.6885 | 269.4115 | 161.5115 | 2.2923 | 0.58391 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | lots of big cells seen on many filters
AR66B | 20220420 | L4 | 3 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 182.7 | 112.6 | 0.63933 | 182.0607 | 111.9607 | 1.4892 | 0.50455 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | lots of big cells seen on many filters
AR66B | 20220420 | L4 | 3 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 166.5 | 101.8 | 0.6885 | 165.8115 | 101.1115 | 1.3745 | 0.42607 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | lots of big cells seen on many filters
AR66B | 20220420 | L4 | 3 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 58.91 | 37.99 | 0.71583 | 58.1942 | 37.2742 | 0.44443 | 0.21934 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L4 | 3 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.62 | 40.84 | 0.71583 | 63.9042 | 40.1242 | 0.50519 | 0.20934 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L5 | 4 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.7 | 49.06 | 0.6885 | 70.0115 | 48.3715 | 0.45973 | 0.40167 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L5 | 4 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.63 | 45.15 | 0.6885 | 63.9415 | 44.4615 | 0.41384 | 0.37792 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | 2 tiss cap labeled C2N2 repb but these values clearly show C2 was typo and this sample matched rep b very well
AR66B | 20220420 | L5 | 4 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.48 | 44.03 | 0.71583 | 63.7642 | 43.3142 | 0.43445 | 0.33689 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | 2 tiss cap labeled C2N4 repa but these values clearly show C2 was typo and this sample matched rep b very well
AR66B | 20220420 | L5 | 4 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.71 | 43.74 | 0.6885 | 63.0215 | 43.0515 | 0.42425 | 0.34241 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L5 | 4 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 130.7 | 83.19 | 0.71583 | 129.9842 | 82.4742 | 1.0093 | 0.45937 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L5 | 4 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 136 | 85.52 | 0.6885 | 135.3115 | 84.8315 | 1.0724 | 0.43825 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L5 | 4 | 8 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 185.2 | 110.8 | 0.71583 | 184.4842 | 110.0842 | 3.1612 | 0.75957 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L5 | 4 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 162.8 | 97.72 | 0.6885 | 162.1115 | 97.0315 | 4.1477 | 1.036 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L5 | 4 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 268.3 | 162.7 | 0.63933 | 267.6607 | 162.0607 | 2.2434 | 0.64255 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | There are two C4N10a tissue capsules and the chlorophyll readings are very different from each other; 13-29
AR66B | 20220420 | L5 | 4 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 366.7 | 218.2 | 0.6885 | 366.0115 | 217.5115 | 3.1548 | 0.71863 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Tiss cap labeled C4N10 chla which already existed. Enter as chlb but not great matchup. Flag as 2
AR66B | 20220420 | L5 | 4 | 15 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 141.5 | 85.11 | 0.71583 | 140.7842 | 84.3942 | 2.3959 | 0.60982 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L5 | 4 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 226.9 | 137.8 | 0.63933 | 226.2607 | 137.1607 | 1.8929 | 0.54967 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L5 | 4 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.15 | 41.9 | 0.6885 | 64.4615 | 41.2115 | 0.49393 | 0.23996 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220420 | L5 | 4 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.31 | 43.38 | 0.71583 | 65.5942 | 42.6642 | 0.48713 | 0.27263 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220421 | L10 | 6 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22.67 | 16.15 | 0.71583 | 21.9542 | 15.4342 | 0.13851 | 0.13634 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220421 | L10 | 6 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23.6 | 16.44 | 0.6885 | 22.9115 | 15.7515 | 0.15211 | 0.12839 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220421 | L10 | 6 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 27.9 | 19.68 | 0.63933 | 27.2607 | 19.0407 | 0.17463 | 0.16445 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220421 | L10 | 6 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 29.48 | 20.25 | 0.71583 | 28.7642 | 19.5342 | 0.19609 | 0.15178 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220421 | L10 | 6 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 168.5 | 109.3 | 0.71583 | 167.7842 | 108.5842 | 1.2577 | 0.67599 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/01
AR66B | 20220421 | L10 | 6 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 195.2 | 122.4 | 0.63933 | 194.5607 | 121.7607 | 1.5466 | 0.62171 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/02
AR66B | 20220421 | L10 | 6 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 143.7 | 91.05 | 0.71583 | 142.9842 | 90.3342 | 1.1185 | 0.49014 | 20220519 | 13-31 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220421 | L10 | 6 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 136.5 | 86.52 | 0.63933 | 135.8607 | 85.8807 | 1.0618 | 0.46756 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220421 | L10 | 6 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 105.5 | 66.04 | 0.71583 | 104.7842 | 65.3242 | 0.8383 | 0.32498 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/03
AR66B | 20220421 | L10 | 6 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 111.9 | 70.94 | 0.63933 | 111.2607 | 70.3007 | 0.87017 | 0.38174 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 06/04
AR66B | 20220421 | L10 | 6 | 21 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 124.6 | 78.11 | 0.6885 | 123.9115 | 77.4215 | 0.98765 | 0.39106 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220421 | L10 | 6 | 21 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 117.9 | 74.51 | 0.6885 | 117.2115 | 73.8215 | 0.92179 | 0.39281 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220421 | L10 | 6 | 21 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 59.6 | 38.5 | 0.6885 | 58.9115 | 37.8115 | 0.44826 | 0.22509 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | sample not recorded on log but confident sampled and is 290ml
AR66B | 20220421 | L10 | 6 | 21 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 53.91 | 34.92 | 0.71583 | 53.1942 | 34.2042 | 0.40343 | 0.20567 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | sample not recorded on log but confident sampled and is 290ml
AR66B | 20220421 |  | 7 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 112.8 | 75.83 | 0.71583 | 112.0842 | 75.1142 | 0.7854 | 0.55222 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/01; glider recovery
AR66B | 20220421 |  | 7 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 155.9 | 100.3 | 0.63933 | 155.2607 | 99.6607 | 1.1812 | 0.59356 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/02; glider recovery
AR66B | 20220422 | L11 | 8 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23.1 | 16.51 | 0.63933 | 22.4607 | 15.8707 | 0.14 | 0.14262 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L11 | 8 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23.04 | 16.49 | 0.6885 | 22.3515 | 15.8015 | 0.13915 | 0.14224 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L11 | 8 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 82.85 | 53.79 | 0.71583 | 82.1342 | 53.0742 | 0.61736 | 0.32778 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L11 | 8 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 78.69 | 51.85 | 0.63933 | 78.0507 | 51.2107 | 0.5702 | 0.34175 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L11 | 8 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 185.5 | 116.8 | 0.71583 | 184.7842 | 116.0842 | 1.4595 | 0.60773 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L11 | 8 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 187.7 | 117.4 | 0.71583 | 186.9842 | 116.6842 | 1.4935 | 0.58442 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L11 | 8 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 206.5 | 129.6 | 0.63933 | 205.8607 | 128.9607 | 1.6337 | 0.66283 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L11 | 8 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 195 | 121.7 | 0.6885 | 194.3115 | 121.0115 | 1.5572 | 0.59775 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L11 | 8 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 199.4 | 125.6 | 0.6885 | 198.7115 | 124.9115 | 1.5678 | 0.65658 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L11 | 8 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 208.1 | 130.5 | 0.71583 | 207.3842 | 129.7842 | 1.6486 | 0.66262 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L11 | 8 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 224 | 141.1 | 0.63933 | 223.3607 | 140.4607 | 1.7612 | 0.74015 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L11 | 8 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 237.5 | 147.8 | 0.63933 | 236.8607 | 147.1607 | 1.9056 | 0.715 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L11 | 8 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.3 | 52.43 | 0.6885 | 79.6115 | 51.7415 | 0.59208 | 0.32933 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L11 | 8 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 75.64 | 49.28 | 0.6885 | 74.9515 | 48.5915 | 0.56 | 0.30531 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L12 | 11 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.52 | 43.56 | 0.6885 | 65.8315 | 42.8715 | 0.48777 | 0.27568 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 11/01
AR66B | 20220422 | L12 | 11 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 72.69 | 47.54 | 0.66617 | 72.0238 | 46.8738 | 0.53429 | 0.30043 | 20220519 | 13-44 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 11/02 | tissue capsule found at the bottom of a dewar
AR66B | 20220422 | L12 | 11 | 9 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 221.4 | 134.9 | 0.66617 | 220.7338 | 134.2338 | 3.6753 | 1.1056 | 20220519 | 13-44 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | tissue capsule found at the bottom of a dewar
AR66B | 20220422 | L12 | 11 | 9 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 204.7 | 126 | 0.6885 | 204.0115 | 125.3115 | 3.3439 | 1.1192 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L12 | 11 | 10 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 188.9 | 115.6 | 0.71583 | 188.1842 | 114.8842 | 3.1144 | 0.97727 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 11/03
AR66B | 20220422 | L12 | 11 | 10 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 190.1 | 115.6 | 0.63933 | 189.4607 | 114.9607 | 3.1654 | 0.929 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 11/04
AR66B | 20220422 | L9 | 15 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 72.28 | 49.7 | 0.6885 | 71.5915 | 49.0115 | 0.4797 | 0.39309 | 20220519 | 13-29 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L9 | 15 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.2 | 49.59 | 0.71583 | 69.4842 | 48.8742 | 0.43785 | 0.4325 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L9 | 15 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.82 | 35.41 | 0.63933 | 49.1807 | 34.7707 | 0.30613 | 0.31306 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  |  | phae>chl; depth 30m
AR66B | 20220422 | L9 | 15 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.82 | 33.17 | 0.63933 | 45.1807 | 32.5307 | 0.26874 | 0.31056 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  |  | phae>chl; depth 30m
AR66B | 20220422 | L9 | 15 | 12 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 171.5 | 106.3 | 0.6885 | 170.8115 | 105.6115 | 2.7703 | 0.99117 | 20220519 | 13-30 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 15/01
AR66B | 20220422 | L9 | 15 | 12 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 173.6 | 108.2 | 0.66617 | 172.9338 | 107.5338 | 2.7788 | 1.0511 | 20220519 | 13-44 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 15/02; N12chlb bad job trying to pick copepod off | tissue capsule found at the bottom of a dewar
AR66B | 20220422 | L9 | 15 | 15 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 178.6 | 112 | 0.6885 | 177.9115 | 111.3115 | 2.8297 | 1.1347 | 20220519 | 13-30 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L9 | 15 | 15 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 179.8 | 112.9 | 0.63933 | 179.1607 | 112.2607 | 2.8425 | 1.1558 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | N15 chlb refill half bottle later
AR66B | 20220422 | L9 | 15 | 18 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 188.6 | 115.9 | 0.63933 | 187.9607 | 115.2607 | 3.0889 | 1.0162 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 15/03; poor replication
AR66B | 20220422 | L9 | 15 | 18 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 208.2 | 125.6 | 0.63933 | 207.5607 | 124.9607 | 3.5096 | 0.941 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 15/04; poor replication
AR66B | 20220422 | L9 | 15 | 20 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 164.6 | 99.17 | 0.6885 | 163.9115 | 98.4815 | 2.78 | 0.72746 | 20220519 | 13-30 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L9 | 15 | 20 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 164.4 | 99.72 | 0.71583 | 163.6842 | 99.0042 | 2.7482 | 0.77794 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220422 | L9 | 15 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 159.5 | 98.51 | 0.66617 | 158.8338 | 97.8438 | 1.2957 | 0.4467 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | tissue capsule found at the bottom of a dewar
AR66B | 20220422 | L9 | 15 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 141.5 | 87.02 | 0.63933 | 140.8607 | 86.3807 | 1.1574 | 0.38086 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220424 | L7 | 19 | 7 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [9 empty] | OOI 19/01
AR66B | 20220424 | L7 | 19 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 117 | 79.38 | 0.71583 | 116.2842 | 78.6642 | 0.79921 | 0.60163 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 19/02
AR66B | 20220424 | L7 | 19 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 207.3 | 128.8 | 0.71583 | 206.5842 | 128.0842 | 1.6677 | 0.61323 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220424 | L7 | 19 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 193.8 | 123.3 | 0.66617 | 193.1338 | 122.6338 | 1.4977 | 0.68612 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | tissue capsule found at the bottom of a dewar
AR66B | 20220424 | L7 | 19 | 12 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 140.6 | 87.72 | 0.63933 | 139.9607 | 87.0807 | 2.2468 | 0.85464 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 19/03; large cells
AR66B | 20220424 | L7 | 19 | 12 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 124 | 77.67 | 0.63933 | 123.3607 | 77.0307 | 1.9685 | 0.775 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 19/04; large cells
AR66B | 20220424 | L7 | 19 | 14 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 167.5 | 104 | 0.63933 | 166.8607 | 103.3607 | 2.698 | 0.98324 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220424 | L7 | 19 | 14 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 174.4 | 107.7 | 0.66617 | 173.7338 | 107.0338 | 2.834 | 0.97809 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | tissue capsule found at the bottom of a dewar
AR66B | 20220424 | L7 | 19 | 18 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 148.2 | 92.07 | 0.63933 | 147.5607 | 91.4307 | 2.3849 | 0.87148 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 19/05
AR66B | 20220424 | L7 | 19 | 18 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 157.4 | 96.83 | 0.66617 | 156.7338 | 96.1638 | 2.5735 | 0.85141 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 19/06 | tissue capsule found at the bottom of a dewar
AR66B | 20220424 | L7 | 19 | 20 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 146.4 | 90.01 | 0.6885 | 145.7115 | 89.3215 | 2.3959 | 0.78532 | 20220519 | 13-30 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220424 | L7 | 19 | 20 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 148.5 | 90.77 | 0.66617 | 147.8338 | 90.1038 | 2.4529 | 0.75624 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | tissue capsule found at the bottom of a dewar
AR66B | 20220424 | L7 | 19 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 165.7 | 103.1 | 0.66617 | 165.0338 | 102.4338 | 1.3299 | 0.49423 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | sample not recorded on logsheet but likely sampled and would be 290ml | tissue capsule found at the bottom of a dewar
AR66B | 20220424 | L7 | 19 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 143 | 89.09 | 0.6885 | 142.3115 | 88.4015 | 1.1453 | 0.42896 | 20220519 | 13-30 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | sample not recorded on logsheet but likely sampled and would be 290ml
AR66B | 20220424 | L8 | 20 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 77.51 | 54.34 | 0.71583 | 76.7942 | 53.6242 | 0.49223 | 0.4627 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220424 | L8 | 20 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.43 | 45.87 | 0.6885 | 63.7415 | 45.1815 | 0.39429 | 0.41029 | 20220519 | 13-30 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | phae>chl; depth 60m
AR66B | 20220424 | L8 | 20 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 206.1 | 130.3 | 0.66617 | 205.4338 | 129.6338 | 1.6103 | 0.69818 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | tissue capsule found at the bottom of a dewar
AR66B | 20220424 | L8 | 20 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 205.7 | 131.7 | 0.66617 | 205.0338 | 131.0338 | 1.5721 | 0.76135 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | tissue capsule found at the bottom of a dewar
AR66B | 20220424 | L8 | 20 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 214.9 | 133.8 | 0.66617 | 214.2338 | 133.1338 | 1.7229 | 0.64791 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | tissue capsule found at the bottom of a dewar
AR66B | 20220424 | L8 | 20 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 212.9 | 134.9 | 0.66617 | 212.2338 | 134.2338 | 1.6571 | 0.73336 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | tissue capsule found at the bottom of a dewar
AR66B | 20220424 | L8 | 20 | 12 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 189.6 | 119 | 0.63933 | 188.9607 | 118.3607 | 2.9997 | 1.2158 | 20220519 | 13-27 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220424 | L8 | 20 | 12 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 178.3 | 110.8 | 0.6885 | 177.6115 | 110.1115 | 2.868 | 1.0537 | 20220519 | 13-30 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220424 | L8 | 20 | 15 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 262 | 158.8 | 0.63933 | 261.3607 | 158.1607 | 4.3848 | 1.2482 | 20220519 | 13-28 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220424 | L8 | 20 | 15 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR66B | 20220424 | L8 | 20 | 18 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 256.7 | 156.9 | 0.66617 | 256.0338 | 156.2338 | 4.2404 | 1.324 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | tissue capsule found at the bottom of a dewar
AR66B | 20220424 | L8 | 20 | 18 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 258.8 | 159 | 0.66617 | 258.1338 | 158.3338 | 4.2404 | 1.3988 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | tissue capsule found at the bottom of a dewar
AR66B | 20220424 | L8 | 20 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 190.5 | 118.8 | 0.71583 | 189.7842 | 118.0842 | 1.5232 | 0.57961 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220424 | L8 | 20 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 189.1 | 117.7 | 0.6885 | 188.4115 | 117.0115 | 1.5168 | 0.56688 | 20220519 | 13-30 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220425 | L13 | 21 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 118.9 | 75.86 | 0.66617 | 118.2338 | 75.1938 | 0.91436 | 0.42469 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 21/01
AR66B | 20220425 | L13 | 21 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 121.5 | 76.78 | 0.63933 | 120.8607 | 76.1407 | 0.95005 | 0.40586 | 20220519 | 13-28 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 21/02
AR66B | 20220425 | L13 | 21 | 6 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 203.9 | 129.3 | 0.66617 | 203.2338 | 128.6338 | 3.1697 | 1.4117 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 21/03; poor replication | poor replicate; tissue capsule was found at the bottom of a dewar
AR66B | 20220425 | L13 | 21 | 6 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 167.4 | 104.8 | 0.66617 | 166.7338 | 104.1338 | 2.6598 | 1.049 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 21/04; poor replication | poor replicate; tissue capsule was found at the bottom of a dewar
AR66B | 20220425 | L13 | 21 | 8 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR66B | 20220425 | L13 | 21 | 8 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 207.8 | 128.8 | 0.6885 | 207.1115 | 128.1115 | 3.3566 | 1.2062 | 20220519 | 13-30 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220425 | L13 | 21 | 9 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 142.3 | 89.44 | 0.66617 | 141.6338 | 88.7738 | 2.246 | 0.91579 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 21/05 | tissue capsule found at the bottom of a dewar
AR66B | 20220425 | L13 | 21 | 9 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 145.5 | 92.05 | 0.71583 | 144.7842 | 91.3342 | 2.271 | 0.98191 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 21/06
AR66B | 20220425 | L6 | 22 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 43.32 | 33.14 | 0.66617 | 42.6538 | 32.4738 | 0.21627 | 0.36202 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | I would usually flag this as a 2 | phae>chl; depth 87m; tissue capsule was found at the bottom of a dewar
AR66B | 20220425 | L6 | 22 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 44.03 | 32.58 | 0.6885 | 43.3415 | 31.8915 | 0.24325 | 0.32467 | 20220519 | 13-30 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | phae>chl; depth 87m
AR66B | 20220425 | L6 | 22 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 119.4 | 74.84 | 0.66617 | 118.7338 | 74.1738 | 0.94665 | 0.37423 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | tissue capsule found at the bottom of a dewar
AR66B | 20220425 | L6 | 22 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 101.1 | 63.95 | 0.6885 | 100.4115 | 63.2615 | 0.78923 | 0.33733 | 20220519 | 13-30 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220425 | L6 | 22 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 160.8 | 102.3 | 0.66617 | 160.1338 | 101.6338 | 1.2428 | 0.56709 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | tissue capsule found at the bottom of a dewar
AR66B | 20220425 | L6 | 22 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 172.7 | 107.2 | 0.71583 | 171.9842 | 106.4842 | 1.3915 | 0.50475 | 20220519 | 13-32 | 20230411 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220425 | L6 | 22 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 230 | 143.2 | 0.6885 | 229.3115 | 142.5115 | 1.844 | 0.69382 | 20220519 | 13-30 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220425 | L6 | 22 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 225.7 | 140.5 | 0.6885 | 225.0115 | 139.8115 | 1.81 | 0.67973 | 20220519 | 13-30 | 20230406 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220425 | L6 | 22 | 11 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 128.8 | 79.95 | 0.63933 | 128.1607 | 79.3107 | 2.0756 | 0.74914 | 20220519 | 13-28 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220425 | L6 | 22 | 11 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR66B | 20220425 | L6 | 22 | 14 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 151.4 | 94.76 | 0.63933 | 150.7607 | 94.1207 | 2.4066 | 0.94562 | 20220519 | 13-28 | 20230404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR66B | 20220425 | L6 | 22 | 14 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 157.4 | 98.46 | 0.66617 | 156.7338 | 97.7938 | 2.5043 | 0.97872 | 20220519 | 13-45 | 20230712 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | tissue capsule found at the bottom of a dewar
AR66B | 20220425 | L6 | 22 | 14 |  |  | a | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR66B | 20220425 | L6 | 22 | 14 |  |  | b | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN687 | 20220729 | L1 | 1 | 1 |  |  | a | 290 | 0 | 15 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 95.63 | 57.9 | 0.59517 | 95.0348 | 57.3048 | 2.4046 | 0.65678 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L1 | 1 | 1 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 164.4 | 100.4 | 0.59517 | 163.8048 | 99.8048 | 2.7193 | 0.83535 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L1 | 1 | 3 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 134.4 | 81.58 | 0.59517 | 133.8048 | 80.9848 | 2.2443 | 0.64008 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L1 | 1 | 3 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 137.6 | 83.62 | 0.59517 | 137.0048 | 83.0248 | 2.2935 | 0.66345 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L1 | 1 | 3 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 170.1 | 104 | 0.5395 | 169.5605 | 103.4605 | 1.4043 | 0.43816 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220729 | L1 | 1 | 3 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 176 | 108.8 | 0.5395 | 175.4605 | 108.2605 | 1.4276 | 0.50027 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220729 | L1 | 1 | 3 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 82.53 | 49.26 | 0.83 | 81.7 | 48.43 | 0.69956 | 0.15404 | 20220519 | 11-44 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR5
EN687 | 20220729 | L1 | 1 | 3 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 95.93 | 55.87 | 0.73067 | 95.1993 | 55.1393 | 0.83947 | 0.12909 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR6
EN687 | 20220729 | L1 | 1 | 3 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 84.65 | 49.05 | 0.83 | 83.82 | 48.22 | 0.39878 | 0.053991 | 20220519 | 11-44 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR11
EN687 | 20220729 | L1 | 1 | 3 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 94.49 | 54.98 | 0.83 | 93.66 | 54.15 | 0.44257 | 0.065874 | 20220519 | 11-44 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR12
EN687 | 20220729 | L1 | 1 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 144.1 | 88.48 | 0.59517 | 143.5048 | 87.8848 | 1.1816 | 0.38343 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L1 | 1 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 149.3 | 88.8 | 0.59517 | 148.7048 | 88.2048 | 1.2853 | 0.28546 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L1 | 1 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 87.28 | 54.11 | 0.5395 | 86.7405 | 53.5705 | 0.70467 | 0.2493 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220729 | L1 | 1 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.58 | 50.27 | 0.5395 | 80.0405 | 49.7305 | 0.64392 | 0.24168 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | filter a bit overly wet going into extraction. Possibly dilute reading. Flag as 2 since lower than rep a
EN687 | 20220729 | L1 | 1 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 35.33 | 19.85 | 0.83 | 34.5 | 19.02 | 0.3255 | 0.0097426 | 20220519 | 11-44 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR3
EN687 | 20220729 | L1 | 1 | 10 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.98 | 26.26 | 0.83 | 45.15 | 25.43 | 0.41465 | 0.033569 | 20220519 | 11-44 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR4; lost ~1ml cup leaked
EN687 | 20220729 | L1 | 1 | 10 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.16 | 40.49 | 0.63617 | 69.5238 | 39.8538 | 0.33235 | 0.041861 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR9; lost ~1ml
EN687 | 20220729 | L1 | 1 | 10 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 60.33 | 33.62 | 0.83 | 59.5 | 32.79 | 0.29865 | 0.0086757 | 20220519 | 11-44 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR10
EN687 | 20220729 | L1 | 1 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 104.1 | 62.08 | 0.59517 | 103.5048 | 61.4848 | 0.89269 | 0.20223 | 20220519 | 11-73 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L1 | 1 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 104.6 | 63.11 | 0.59517 | 104.0048 | 62.5148 | 0.88143 | 0.23183 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L1 | 1 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 84.07 | 49.61 | 0.5395 | 83.5305 | 49.0705 | 0.73208 | 0.14176 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | sort of poor chla replication but phaeo good match. This value a bit close to WSW value. Could <10 filter have leaked?
EN687 | 20220729 | L1 | 1 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 78.62 | 46.72 | 0.5395 | 78.0805 | 46.1805 | 0.67769 | 0.14468 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | sort of poor chla replication but phaeo good match
EN687 | 20220729 | L1 | 1 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.18 | 14.71 | 0.83 | 24.35 | 13.88 | 0.22015 | 0.024492 | 20220519 | 11-44 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR1
EN687 | 20220729 | L1 | 1 | 15 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23.57 | 14.09 | 0.83 | 22.74 | 13.26 | 0.19933 | 0.03438 | 20220519 | 11-44 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR2
EN687 | 20220729 | L1 | 1 | 15 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.76 | 11.49 | 0.63617 | 19.1238 | 10.8538 | 0.092637 | 0.0092763 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR7
EN687 | 20220729 | L1 | 1 | 15 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.01 | 9.963 | 0.63617 | 16.3738 | 9.3268 | 0.078651 | 0.0086066 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR8
EN687 | 20220729 | L2 | 2 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.25 | 47.53 | 0.59517 | 65.6548 | 46.9348 | 0.39769 | 0.43812 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L2 | 2 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.1 | 46.27 | 0.59517 | 62.5048 | 45.6748 | 0.35754 | 0.45583 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L2 | 2 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 67.59 | 48.83 | 0.59517 | 66.9948 | 48.2348 | 0.39854 | 0.46042 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L2 | 2 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 73.23 | 50.93 | 0.59517 | 72.6348 | 50.3348 | 0.47375 | 0.42261 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L2 | 2 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 147.1 | 96.93 | 0.59517 | 146.5048 | 96.3348 | 1.0658 | 0.64969 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L2 | 2 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 156.8 | 101.7 | 0.59517 | 156.2048 | 101.1048 | 1.1706 | 0.6299 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L2 | 2 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 108.3 | 71.57 | 0.5395 | 107.7605 | 71.0305 | 0.7803 | 0.4846 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220729 | L2 | 2 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 105 | 69.91 | 0.5395 | 104.4605 | 69.3705 | 0.74546 | 0.48988 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220729 | L2 | 2 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 46.38 | 29.92 | 0.63617 | 45.7438 | 29.2838 | 0.3461 | 0.17004 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR17; poor replication (31% range) but DF thinks within bio variation
EN687 | 20220729 | L2 | 2 | 8 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 62.75 | 40.15 | 0.63617 | 62.1138 | 39.5138 | 0.47521 | 0.22125 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR18; poor replication (31% range) but DF thinks within bio variation
EN687 | 20220729 | L2 | 2 | 8 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 27.8 | 16.75 | 0.83 | 26.97 | 15.92 | 0.12355 | 0.025659 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR23; poor replication (68% range) but low values and just within bio variation?
EN687 | 20220729 | L2 | 2 | 8 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.52 | 30.02 | 0.83 | 51.69 | 29.19 | 0.25158 | 0.022007 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR24; poor replication (68% range) but low values and just within bio variation?
EN687 | 20220729 | L2 | 2 | 10 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 324.8 | 201 | 0.59517 | 324.2048 | 200.4048 | 5.2601 | 1.8775 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | 17% range for replication but lots of chl sohigh variability
EN687 | 20220729 | L2 | 2 | 10 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 380.5 | 232.3 | 0.59517 | 379.9048 | 231.7048 | 6.2968 | 1.9555 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | 17% range for replication but lots of chl sohigh variability
EN687 | 20220729 | L2 | 2 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 245.9 | 158.4 | 0.5395 | 245.3605 | 157.8605 | 1.8589 | 0.95228 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220729 | L2 | 2 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 254.4 | 159.4 | 0.5395 | 253.8605 | 158.8605 | 2.0182 | 0.81076 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220729 | L2 | 2 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 176.4 | 101.3 | 0.63617 | 175.7638 | 100.6638 | 4.7373 | 0.58542 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR15
EN687 | 20220729 | L2 | 2 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 209.2 | 119 | 0.63617 | 208.5638 | 118.3638 | 5.6705 | 0.56689 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR16
EN687 | 20220729 | L2 | 2 | 10 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 286.4 | 160.1 | 0.83 | 285.57 | 159.27 | 4.2443 | 0.24218 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR21; poor replication (36% range); DF thinks something happened with this sample baed on size fraction patterns for rest of cast.
EN687 | 20220729 | L2 | 2 | 10 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 200.7 | 113.3 | 0.83 | 199.87 | 112.47 | 2.937 | 0.2311 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR22; poor replication (36 % range)
EN687 | 20220729 | L2 | 2 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.1 | 39.16 | 0.59517 | 62.5048 | 38.5648 | 0.50859 | 0.17817 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L2 | 2 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.88 | 40.15 | 0.59517 | 65.2848 | 39.5548 | 0.54662 | 0.15777 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220729 | L2 | 2 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 50.91 | 31.31 | 0.5395 | 50.3705 | 30.7705 | 0.41639 | 0.13157 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220729 | L2 | 2 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 53.16 | 32.63 | 0.5395 | 52.6205 | 32.0905 | 0.43615 | 0.13532 | 20220519 | 13-05 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220729 | L2 | 2 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.88 | 7.324 | 0.63617 | 11.2438 | 6.6878 | 0.095798 | 0.022078 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR13
EN687 | 20220729 | L2 | 2 | 16 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.09 | 7.369 | 0.63617 | 11.4538 | 6.7328 | 0.099268 | 0.019402 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR14
EN687 | 20220729 | L2 | 2 | 16 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.37 | 6.8 | 0.83 | 10.54 | 5.97 | 0.051191 | 0.0048648 | 20220519 | 11-44 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR19
EN687 | 20220729 | L2 | 2 | 16 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 9.095 | 5.508 | 0.83 | 8.265 | 4.678 | 0.040034 | 0.003731 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR20
EN687 | 20220730 | L4 | 5 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.62 | 22.14 | 0.59517 | 28.0248 | 21.5448 | 0.13766 | 0.246 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | poor replication (56%)
EN687 | 20220730 | L4 | 5 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.11 | 24.49 | 0.6235 | 27.4865 | 23.8665 | 0.076904 | 0.34811 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | RA still pretty high; poor replication (56%); chl/phaeo ratio is off. This one seems to be compromised. Toss and assume other rep is fine
EN687 | 20220730 | L4 | 5 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.67 | 45.65 | 0.59517 | 65.0748 | 45.0548 | 0.42531 | 0.37702 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L4 | 5 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.47 | 46.13 | 0.59517 | 64.8748 | 45.5348 | 0.41087 | 0.40001 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L4 | 5 | 5 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.71 | 9.648 | 0.83 | 12.88 | 8.818 | 0.085411 | 0.070011 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR31
EN687 | 20220730 | L4 | 5 | 5 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.67 | 10.59 | 0.83 | 13.84 | 9.76 | 0.085789 | 0.086236 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR32
EN687 | 20220730 | L4 | 5 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.221 | 5.985 | 0.63617 | 7.5848 | 5.3488 | 0.025001 | 0.025131 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR39
EN687 | 20220730 | L4 | 5 | 5 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.069 | 5.076 | 0.62633 | 6.4427 | 4.4497 | 0.022284 | 0.019421 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR40 | No blank reading after running, so only 3 values for verage
EN687 | 20220730 | L4 | 5 | 8 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 109.1 | 69.12 | 0.59517 | 108.5048 | 68.5248 | 1.6987 | 0.74187 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L4 | 5 | 8 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 107.6 | 69.51 | 0.59517 | 107.0048 | 68.9148 | 1.6184 | 0.83606 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L4 | 5 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 147.6 | 96 | 0.5395 | 147.0605 | 95.4605 | 1.0962 | 0.60374 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220730 | L4 | 5 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 149.6 | 96.18 | 0.5395 | 149.0605 | 95.6405 | 1.1349 | 0.56828 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220730 | L4 | 5 | 8 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 62.73 | 38.12 | 0.83 | 61.9 | 37.29 | 0.51571 | 0.13931 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR29
EN687 | 20220730 | L4 | 5 | 8 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.28 | 42.01 | 0.83 | 69.45 | 41.18 | 0.59443 | 0.13139 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR30
EN687 | 20220730 | L4 | 5 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 79.37 | 47.25 | 0.83 | 78.54 | 46.42 | 0.35979 | 0.076071 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR33
EN687 | 20220730 | L4 | 5 | 8 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 82.91 | 46.58 | 0.63617 | 82.2738 | 45.9438 | 0.40695 | 0.024442 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR38
EN687 | 20220730 | L4 | 5 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 83.86 | 51.92 | 0.59517 | 83.2648 | 51.3248 | 0.67854 | 0.23544 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L4 | 5 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 70.09 | 45.38 | 0.59517 | 69.4948 | 44.7848 | 0.52495 | 0.27257 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L4 | 5 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 53.31 | 33.44 | 0.5395 | 52.7705 | 32.9005 | 0.42212 | 0.16376 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | poor replication
EN687 | 20220730 | L4 | 5 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.05 | 31.93 | 0.5395 | 48.5105 | 31.3905 | 0.3637 | 0.19529 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | poor replication
EN687 | 20220730 | L4 | 5 | 14 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.97 | 10.25 | 0.63617 | 16.3338 | 9.6138 | 0.14082 | 0.028053 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR27
EN687 | 20220730 | L4 | 5 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22.12 | 13.85 | 0.83 | 21.29 | 13.02 | 0.1733 | 0.055404 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR28
EN687 | 20220730 | L4 | 5 | 14 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 26.86 | 15.82 | 0.63617 | 26.2238 | 15.1838 | 0.12344 | 0.018871 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR35
EN687 | 20220730 | L4 | 5 | 14 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22.34 | 13 | 0.63617 | 21.7038 | 12.3638 | 0.10424 | 0.011427 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR36
EN687 | 20220730 | L4 | 5 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 54.38 | 34.41 | 0.59517 | 53.7848 | 33.8148 | 0.42425 | 0.17792 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L4 | 5 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.95 | 31.91 | 0.59517 | 49.3548 | 31.3148 | 0.38325 | 0.1744 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L4 | 5 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 58.65 | 35.57 | 0.6235 | 58.0265 | 34.9465 | 0.49032 | 0.132 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L4 | 5 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 63.78 | 38.47 | 0.6235 | 63.1565 | 37.8465 | 0.53769 | 0.13627 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L4 | 5 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.72 | 27.55 | 0.5395 | 45.1805 | 27.0105 | 0.38601 | 0.09499 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220730 | L4 | 5 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 46.48 | 28.13 | 0.5395 | 45.9405 | 27.5905 | 0.38983 | 0.10149 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220730 | L4 | 5 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.18 | 7.26 | 0.63617 | 11.5438 | 6.6238 | 0.10345 | 0.013297 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR25
EN687 | 20220730 | L4 | 5 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.34 | 8.748 | 0.63617 | 14.7038 | 8.1118 | 0.13861 | 0.0043667 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR26
EN687 | 20220730 | L4 | 5 | 19 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.65 | 6.708 | 0.63617 | 11.0138 | 6.0718 | 0.055158 | 0.001648 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR34
EN687 | 20220730 | L4 | 5 | 19 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.2 | 6.133 | 0.63617 | 9.5638 | 5.4968 | 0.045557 | 0.0060564 | 20220519 | 11-36 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR37
EN687 | 20220730 | L5 | 6 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 41.82 | 30.4 | 0.59517 | 41.2248 | 29.8048 | 0.24261 | 0.28815 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L5 | 6 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 41.7 | 31.38 | 0.6235 | 41.0765 | 30.7565 | 0.21924 | 0.32847 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L5 | 6 | 5 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.226 | 6.242 | 0.83 | 7.396 | 5.412 | 0.041717 | 0.053672 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR45; poor replication (54% range) but low values and close to noise of flourometer
EN687 | 20220730 | L5 | 6 | 5 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.41 | 10.94 | 0.83 | 13.58 | 10.11 | 0.072963 | 0.10523 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR46; poor replication (54% range) but low values and close to noise of flourometer
EN687 | 20220730 | L5 | 6 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.515 | 5.526 | 0.63617 | 6.8788 | 4.8898 | 0.022239 | 0.023591 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR51
EN687 | 20220730 | L5 | 6 | 5 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.353 | 5.989 | 0.63617 | 7.7168 | 5.3528 | 0.026432 | 0.023737 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR52
EN687 | 20220730 | L5 | 6 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 141.1 | 95.09 | 0.6235 | 140.4765 | 94.4665 | 0.97745 | 0.7048 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | labNB says tiss cap said C1 but process of elimination shows C6N8 not run yet and fits with cast info
EN687 | 20220730 | L5 | 6 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 135.6 | 93.83 | 0.6235 | 134.9765 | 93.2065 | 0.88738 | 0.77243 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | labNB says tiss cap said C1 but process of elimination shows C6N8 not run yet and fits with cast info
EN687 | 20220730 | L5 | 6 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 32.65 | 23.79 | 0.63617 | 32.0138 | 23.1538 | 0.1863 | 0.2218 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR43; poor replication (67% range)
EN687 | 20220730 | L5 | 6 | 8 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 18.6 | 14.21 | 0.83 | 17.77 | 13.38 | 0.091994 | 0.14303 | 20220519 | 11-45 | 20220818 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR44; poor replication (67% range)
EN687 | 20220730 | L5 | 6 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 12.8 | 9.181 | 0.63617 | 12.1638 | 8.5448 | 0.040538 | 0.039694 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR49
EN687 | 20220730 | L5 | 6 | 8 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.43 | 12.24 | 0.63617 | 15.7938 | 11.6038 | 0.046935 | 0.062021 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR50
EN687 | 20220730 | L5 | 6 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 141 | 88.08 | 0.6235 | 140.3765 | 87.4565 | 1.1242 | 0.43316 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L5 | 6 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 145.9 | 89.8 | 0.6235 | 145.2765 | 89.1765 | 1.1918 | 0.39624 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L5 | 6 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 32.42 | 20.05 | 0.59517 | 31.8248 | 19.4548 | 0.26279 | 0.083657 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L5 | 6 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.88 | 20.33 | 0.59517 | 31.2848 | 19.7348 | 0.24537 | 0.10606 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L5 | 6 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 29.85 | 17.77 | 0.59517 | 29.2548 | 17.1748 | 0.25663 | 0.049216 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L5 | 6 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 29.76 | 18.56 | 0.59517 | 29.1648 | 17.9648 | 0.23794 | 0.081979 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220730 | L5 | 6 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 24.23 | 14.76 | 0.5395 | 23.6905 | 14.2205 | 0.20118 | 0.052053 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220730 | L5 | 6 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 24.18 | 15.26 | 0.5395 | 23.6405 | 14.7205 | 0.1895 | 0.072641 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220730 | L5 | 6 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.162 | 4.067 | 0.63617 | 5.5258 | 3.4308 | 0.044051 | 0.016419 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR41
EN687 | 20220730 | L5 | 6 | 15 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.911 | 3.204 | 0.63617 | 4.2748 | 2.5678 | 0.035893 | 0.0093667 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR42
EN687 | 20220730 | L5 | 6 | 15 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.132 | 1.6 | 0.63617 | 1.4958 | 0.96383 | 0.0059376 | 0.0030796 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR47; poor replication (71% range) but values are so low and within noise of flourometer
EN687 | 20220730 | L5 | 6 | 15 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.053 | 1.937 | 0.63617 | 2.4168 | 1.3008 | 0.012501 | -0.00028663 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR48; poor replication (71% range) but values are so low and within noise of flourometer
EN687 | 20220731 | L10 | 7 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 119.7 | 74.63 | 0.6235 | 119.0765 | 74.0065 | 0.95748 | 0.36042 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L10 | 7 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 115.5 | 73.76 | 0.6235 | 114.8765 | 73.1365 | 0.88674 | 0.41567 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L10 | 7 | 8 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 157.6 | 101.2 | 0.6235 | 156.9765 | 100.5765 | 2.3964 | 1.1857 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L10 | 7 | 8 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 155 | 99.37 | 0.59517 | 154.4048 | 98.7748 | 2.3636 | 1.1543 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L10 | 7 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 247.5 | 157 | 0.5395 | 246.9605 | 156.4605 | 1.9226 | 0.86362 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220731 | L10 | 7 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 249.5 | 161.4 | 0.5395 | 248.9605 | 160.8605 | 1.8716 | 0.99296 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220731 | L10 | 7 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 57.84 | 35.87 | 0.63617 | 57.2038 | 35.2338 | 0.46196 | 0.15906 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR55
EN687 | 20220731 | L10 | 7 | 8 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 47.64 | 30.18 | 0.63617 | 47.0038 | 29.5438 | 0.36588 | 0.15308 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR56
EN687 | 20220731 | L10 | 7 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.11 | 12.41 | 0.63617 | 18.4738 | 11.7738 | 0.07505 | 0.035501 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR61; poor replication (54% range) . DF thinks something is up with this sample based on ratio of size fractions for other depths .. Toss and assume othe rep is fine
EN687 | 20220731 | L10 | 7 | 8 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.93 | 20.21 | 0.63617 | 31.2938 | 19.5738 | 0.13128 | 0.052508 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR62; poor replication (54% range)
EN687 | 20220731 | L10 | 7 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 26.16 | 16.97 | 0.59517 | 25.5648 | 16.3748 | 0.19524 | 0.096365 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L10 | 7 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.2 | 15.92 | 0.6235 | 24.5765 | 15.2965 | 0.19715 | 0.075251 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L10 | 7 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.83 | 13.24 | 0.5395 | 20.2905 | 12.7005 | 0.16124 | 0.064924 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220731 | L10 | 7 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.57 | 13.62 | 0.81783 | 20.7522 | 12.8022 | 0.16889 | 0.059087 | 20220519 | 13-00 | 20230213 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220731 | L10 | 7 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.928 | 3.947 | 0.63617 | 5.2918 | 3.3108 | 0.041654 | 0.016701 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR57; poor replication (76% range) but values so low and close to noise of flourometer
EN687 | 20220731 | L10 | 7 | 15 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.698 | 1.816 | 0.63617 | 2.0618 | 1.1798 | 0.018546 | 0.0022495 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR58; poor replication (76% range) but values so low and close to noise of flourometer
EN687 | 20220731 | L10 | 7 | 15 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.002 | 2.036 | 0.63617 | 2.3658 | 1.3998 | 0.010801 | 0.002319 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR63; poor replication (63% range) but values are so low, bascially 0, within noise of flourometer
EN687 | 20220731 | L10 | 7 | 15 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 1.897 | 1.396 | 0.63617 | 1.2608 | 0.75983 | 0.0056018 | 0.0015198 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR64; poor replication (63% range) but values are so low, bascially 0, within noise of flourometer
EN687 | 20220731 | L10 | 7 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.81 | 10.65 | 0.59517 | 16.2148 | 10.0548 | 0.13087 | 0.04819 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L10 | 7 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.87 | 10.96 | 0.6235 | 16.2465 | 10.3365 | 0.12555 | 0.058517 | 20220519 | 11-71 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L10 | 7 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.8 | 7.777 | 0.5395 | 11.2605 | 7.2375 | 0.085466 | 0.043418 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220731 | L10 | 7 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.85 | 7.708 | 0.5395 | 11.3105 | 7.1685 | 0.087994 | 0.039662 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220731 | L10 | 7 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.03 | 3.084 | 0.63617 | 4.3938 | 2.4478 | 0.040918 | 0.0022262 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR53; somewhat poor replication (32% range) but values so low, within noise of flourometer
EN687 | 20220731 | L10 | 7 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.539 | 3.132 | 0.63617 | 3.9028 | 2.4958 | 0.029585 | 0.014406 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR54; somewhat poor replication (32% range) but values so low, within noise of flourometer
EN687 | 20220731 | L10 | 7 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 1.6 | 1.324 | 0.63617 | 0.96383 | 0.68783 | 0.0030916 | 0.0033669 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR59; poor replication (67% range) but values bascially 0, within noise of flourometer
EN687 | 20220731 | L10 | 7 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 1.849 | 1.288 | 0.63617 | 1.2128 | 0.65183 | 0.0062613 | -0.00016302 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR60; poor replication (67% range) but values bascially 0, within noise of flourometer
EN687 | 20220731 | L11 | 10 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.39 | 35.71 | 0.6235 | 51.7665 | 35.0865 | 0.35436 | 0.27046 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L11 | 10 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 50.88 | 34.71 | 0.6235 | 50.2565 | 34.0865 | 0.34352 | 0.26349 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L11 | 10 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.01 | 7.183 | 0.63617 | 10.3738 | 6.5468 | 0.08047 | 0.034922 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR69
EN687 | 20220731 | L11 | 10 | 7 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.3 | 8.716 | 0.63617 | 12.6638 | 8.0798 | 0.096387 | 0.046024 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR70
EN687 | 20220731 | L11 | 10 | 7 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.38 | 3.241 | 0.63617 | 3.7438 | 2.6048 | 0.012735 | 0.011679 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR75; poor replication (75% range) but values very low, within noise of flourometer
EN687 | 20220731 | L11 | 10 | 7 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.877 | 5.353 | 0.63617 | 7.2408 | 4.7168 | 0.028221 | 0.015987 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR76; poor replication (75% range) but values very low, within noise of flourometer
EN687 | 20220731 | L11 | 10 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 170.6 | 110.5 | 0.6235 | 169.9765 | 109.8765 | 1.2768 | 0.67988 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L11 | 10 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 167.6 | 112.9 | 0.6235 | 166.9765 | 112.2765 | 1.1621 | 0.83734 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L11 | 10 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 127.5 | 85.37 | 0.5395 | 126.9605 | 84.8305 | 0.89502 | 0.61563 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220731 | L11 | 10 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 127.6 | 85.46 | 0.5395 | 127.0605 | 84.9205 | 0.89524 | 0.61702 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220731 | L11 | 10 | 9 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 47.34 | 28.88 | 0.63617 | 46.7038 | 28.2438 | 0.38683 | 0.10928 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR67
EN687 | 20220731 | L11 | 10 | 9 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 51.12 | 31.62 | 0.63617 | 50.4838 | 30.9838 | 0.41002 | 0.13608 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR68
EN687 | 20220731 | L11 | 10 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.29 | 10.57 | 0.63617 | 16.6538 | 9.9338 | 0.075274 | 0.018 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR73
EN687 | 20220731 | L11 | 10 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 18.5 | 11.74 | 0.63617 | 17.8638 | 11.1038 | 0.075723 | 0.028538 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR74
EN687 | 20220731 | L11 | 10 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 77.04 | 50.5 | 0.6235 | 76.4165 | 49.8765 | 0.56382 | 0.32437 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L11 | 10 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 75.14 | 49.98 | 0.6235 | 74.5165 | 49.3565 | 0.53451 | 0.34443 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L11 | 10 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 59.78 | 39.68 | 0.5395 | 59.2405 | 39.1405 | 0.42701 | 0.27 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220731 | L11 | 10 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 62.39 | 41.05 | 0.5395 | 61.8505 | 40.5105 | 0.45335 | 0.26805 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220731 | L11 | 10 | 13 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.44 | 7.572 | 0.62633 | 10.8137 | 6.9457 | 0.081055 | 0.04095 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR77; poor replication (57% range). Something seems off with this sample based on size fraction ratios of other depths | No blank reading after running, so only 3 values for verage
EN687 | 20220731 | L11 | 10 | 13 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.9 | 13.92 | 0.62633 | 20.2737 | 13.2937 | 0.14627 | 0.087243 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR78; poor replication (57% range) | No blank reading after running, so only 3 values for verage
EN687 | 20220731 | L11 | 10 | 13 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.182 | 3.993 | 0.62633 | 5.5557 | 3.3667 | 0.024476 | 0.0070786 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR79 | No blank reading after running, so only 3 values for verage
EN687 | 20220731 | L11 | 10 | 13 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.429 | 4.456 | 0.62633 | 5.8027 | 3.8297 | 0.022021 | 0.013808 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR80 | No blank reading after running, so only 3 values for verage
EN687 | 20220731 | L11 | 10 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.12 | 13.04 | 0.59517 | 19.5248 | 12.4448 | 0.15041 | 0.071206 | 20220519 | 11-74 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L11 | 10 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.15 | 12.85 | 0.6235 | 19.5265 | 12.2265 | 0.15508 | 0.062644 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L11 | 10 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.45 | 9.069 | 0.6235 | 12.8265 | 8.4455 | 0.093071 | 0.057325 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L11 | 10 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.03 | 8.686 | 0.6235 | 13.4065 | 8.0625 | 0.11353 | 0.030046 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220731 | L11 | 10 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.21 | 6.516 | 0.5395 | 9.6705 | 5.9765 | 0.078477 | 0.027952 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220731 | L11 | 10 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.85 | 7.146 | 0.5395 | 10.3105 | 6.6065 | 0.078689 | 0.038959 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220731 | L11 | 10 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.286 | 2.296 | 0.63617 | 2.6498 | 1.6598 | 0.020817 | 0.0084389 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR65; poor replication (50% range) but low values and within noise of flourometer
EN687 | 20220731 | L11 | 10 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.324 | 2.661 | 0.63617 | 3.6878 | 2.0248 | 0.034968 | 0.00072118 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR66; poor replication (50% range) but low values and within noise of flourometer
EN687 | 20220731 | L11 | 10 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 1.544 | 1.247 | 0.63617 | 0.90783 | 0.61083 | 0.0033269 | 0.0024086 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR71
EN687 | 20220731 | L11 | 10 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 1.524 | 1.122 | 0.63617 | 0.88783 | 0.48583 | 0.0044867 | 5.8544e-05 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR72
EN687 | 20220801 | L9 | 13 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.74 | 23.88 | 0.6235 | 34.1165 | 23.2565 | 0.23071 | 0.18344 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L9 | 13 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 35.18 | 24.18 | 0.6235 | 34.5565 | 23.5565 | 0.23369 | 0.1858 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L9 | 13 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 82.15 | 54.45 | 0.6235 | 81.5265 | 53.8265 | 0.58847 | 0.37007 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L9 | 13 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 81.62 | 53.33 | 0.6235 | 80.9965 | 52.7065 | 0.601 | 0.33759 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L9 | 13 | 12 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 234.6 | 151.4 | 0.6235 | 233.9765 | 150.7765 | 3.5351 | 1.835 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L9 | 13 | 12 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 237.9 | 149.4 | 0.6235 | 237.2765 | 148.7765 | 3.7602 | 1.5385 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L9 | 13 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 276.7 | 165.6 | 0.63617 | 276.0638 | 164.9638 | 2.3361 | 0.5715 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR85
EN687 | 20220801 | L9 | 13 | 12 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 273.9 | 163.4 | 0.63617 | 273.2638 | 162.7638 | 2.3156 | 0.54348 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR86
EN687 | 20220801 | L9 | 13 | 12 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 265.1 | 155.5 | 0.63617 | 264.4638 | 154.8638 | 1.2277 | 0.22642 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR91
EN687 | 20220801 | L9 | 13 | 12 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 268.1 | 155.8 | 0.63617 | 267.4638 | 155.1638 | 1.2579 | 0.19899 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR92
EN687 | 20220801 | L9 | 13 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 64.64 | 41.34 | 0.6235 | 64.0165 | 40.7165 | 0.49499 | 0.23008 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | labNb said C19N16a but that DNE and once all samples run process of elimination says it was C13
EN687 | 20220801 | L9 | 13 | 16 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN687 | 20220801 | L9 | 13 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.6 | 9.471 | 0.63617 | 14.9638 | 8.8348 | 0.12887 | 0.026845 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR83
EN687 | 20220801 | L9 | 13 | 16 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.01 | 10.26 | 0.63617 | 16.3738 | 9.6238 | 0.14193 | 0.027694 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR84
EN687 | 20220801 | L9 | 13 | 16 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.884 | 3.346 | 0.63617 | 4.2478 | 2.7098 | 0.017197 | 0.0082014 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR89
EN687 | 20220801 | L9 | 13 | 16 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.942 | 3.901 | 0.63617 | 5.3058 | 3.2648 | 0.022821 | 0.007779 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR90
EN687 | 20220801 | L9 | 13 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.92 | 13.59 | 0.6235 | 20.2965 | 12.9665 | 0.15572 | 0.075185 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220801 | L9 | 13 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.81 | 13.61 | 0.6235 | 20.1865 | 12.9865 | 0.15296 | 0.078303 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | had to partially refill bottle to redo filter much later, particles might have settled in CTD niskin
EN687 | 20220801 | L9 | 13 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.56 | 11.48 | 0.5395 | 17.0205 | 10.9405 | 0.12917 | 0.065661 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220801 | L9 | 13 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.95 | 11.35 | 0.6235 | 16.3265 | 10.7265 | 0.11897 | 0.072048 | 20220519 | 11-72 | 20221106 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220801 | L9 | 13 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.554 | 3.788 | 0.63617 | 4.9178 | 3.1518 | 0.037133 | 0.018419 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR81
EN687 | 20220801 | L9 | 13 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.783 | 3.136 | 0.63617 | 4.1468 | 2.4998 | 0.034631 | 0.0094297 | 20220519 | 11-37 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR82
EN687 | 20220801 | L9 | 13 | 19 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 1.534 | 1.112 | 0.63617 | 0.89783 | 0.47583 | 0.0047099 | -0.00025823 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR87; poor replication (39% range) but such low value that bascially 0 and within noise of flourometer
EN687 | 20220801 | L9 | 13 | 19 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.036 | 1.338 | 0.63617 | 1.3998 | 0.70183 | 0.0078187 | -0.0012287 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR88; poor replication (39% range) and negatvie phaeo but such low value that bascially 0 and within noise of flourometer
EN687 | 20220801 | L8 | 16 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.12 | 14.46 | 0.5195 | 19.6005 | 13.9405 | 0.12024 | 0.12801 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L8 | 16 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22 | 15.79 | 0.5195 | 21.4805 | 15.2705 | 0.13193 | 0.14001 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L8 | 16 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 51.77 | 35.96 | 0.5195 | 51.2505 | 35.4405 | 0.33587 | 0.29525 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L8 | 16 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.23 | 35.58 | 0.5195 | 51.7105 | 35.0605 | 0.35372 | 0.27063 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L8 | 16 | 3 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.27 | 14.39 | 0.63617 | 20.6338 | 13.7538 | 0.14466 | 0.097754 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR97
EN687 | 20220801 | L8 | 16 | 3 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 22.84 | 14.87 | 0.63617 | 22.2038 | 14.2338 | 0.16758 | 0.083295 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR98
EN687 | 20220801 | L8 | 16 | 3 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.88 | 7.042 | 0.63617 | 10.2438 | 6.4058 | 0.042914 | 0.017126 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR103
EN687 | 20220801 | L8 | 16 | 3 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.581 | 5.751 | 0.63617 | 7.9448 | 5.1148 | 0.031643 | 0.016296 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR104
EN687 | 20220801 | L8 | 16 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 116.9 | 80.09 | 0.5195 | 116.3805 | 79.5705 | 0.782 | 0.63498 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L8 | 16 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 103 | 70.09 | 0.5195 | 102.4805 | 69.5705 | 0.69915 | 0.53975 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L8 | 16 | 9 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 187.5 | 122.6 | 0.5195 | 186.9805 | 122.0805 | 2.7575 | 1.5905 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L8 | 16 | 9 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 184 | 121 | 0.5195 | 183.4805 | 120.4805 | 2.6768 | 1.6142 | 20220519 | 11-63 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L8 | 16 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 234.1 | 155.1 | 0.5395 | 233.5605 | 154.5605 | 1.6783 | 1.0741 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220801 | L8 | 16 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 253.4 | 168.5 | 0.5395 | 252.8605 | 167.9605 | 1.8036 | 1.1874 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220801 | L8 | 16 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 194.2 | 118.7 | 0.63617 | 193.5638 | 118.0638 | 1.5875 | 0.49341 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR95
EN687 | 20220801 | L8 | 16 | 9 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 194.7 | 118.9 | 0.63617 | 194.0638 | 118.2638 | 1.5884 | 0.48896 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR96
EN687 | 20220801 | L8 | 16 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 89.29 | 52.9 | 0.63617 | 88.6538 | 52.2638 | 0.40762 | 0.083112 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR101
EN687 | 20220801 | L8 | 16 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 89.26 | 52.48 | 0.63617 | 88.6238 | 51.8438 | 0.41199 | 0.0748 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR102
EN687 | 20220801 | L8 | 16 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 88.86 | 61.03 | 0.5195 | 88.3405 | 60.5105 | 0.59123 | 0.48633 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L8 | 16 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 101.7 | 69.96 | 0.5195 | 101.1805 | 69.4405 | 0.6743 | 0.56229 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L8 | 16 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 75.89 | 52.83 | 0.5395 | 75.3505 | 52.2905 | 0.48989 | 0.44129 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220801 | L8 | 16 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 79.72 | 54.81 | 0.5395 | 79.1805 | 54.2705 | 0.5292 | 0.43725 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220801 | L8 | 16 | 13 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 27.29 | 19 | 0.62633 | 26.6637 | 18.3737 | 0.17372 | 0.14902 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR105 | No blank reading after running, so only 3 values for verage
EN687 | 20220801 | L8 | 16 | 13 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 26.55 | 17.79 | 0.62633 | 25.9237 | 17.1637 | 0.18357 | 0.11792 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR106 | No blank reading after running, so only 3 values for verage
EN687 | 20220801 | L8 | 16 | 13 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.479 | 3.219 | 0.62633 | 3.8527 | 2.5927 | 0.014088 | 0.010212 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR107 | No blank reading after running, so only 3 values for verage
EN687 | 20220801 | L8 | 16 | 13 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.194 | 3.596 | 0.62633 | 4.5677 | 2.9697 | 0.017835 | 0.0099477 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR108 | No blank reading after running, so only 3 values for verage
EN687 | 20220801 | L8 | 16 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.57 | 25.43 | 0.5195 | 38.0505 | 24.9105 | 0.27915 | 0.16445 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L8 | 16 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.21 | 24.1 | 0.5195 | 35.6905 | 23.5805 | 0.25727 | 0.16265 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220801 | L8 | 16 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.28 | 19.25 | 0.5395 | 27.7405 | 18.7105 | 0.19184 | 0.14136 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220801 | L8 | 16 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 27.58 | 19.05 | 0.5395 | 27.0405 | 18.5105 | 0.18121 | 0.14842 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220801 | L8 | 16 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.22 | 11.62 | 0.63617 | 16.5838 | 10.9838 | 0.11775 | 0.075846 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR93
EN687 | 20220801 | L8 | 16 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.93 | 10.44 | 0.63617 | 15.2938 | 9.8038 | 0.11544 | 0.057361 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR94
EN687 | 20220801 | L8 | 16 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.458 | 2.407 | 0.63617 | 2.8218 | 1.7708 | 0.011773 | 0.0048546 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR99
EN687 | 20220801 | L8 | 16 | 19 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.83 | 1.986 | 0.63617 | 2.1938 | 1.3498 | 0.0094541 | 0.0032203 | 20220519 | 11-38 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR100
EN687 | 20220802 | L6 | 18 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.92 | 21.86 | 0.5395 | 31.3805 | 21.3205 | 0.21372 | 0.16595 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220802 | L6 | 18 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.51 | 22.93 | 0.5195 | 32.9905 | 22.4105 | 0.22477 | 0.17432 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220802 | L6 | 18 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 40.85 | 28.3 | 0.5195 | 40.3305 | 27.7805 | 0.26662 | 0.2281 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L6 | 18 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.06 | 25.98 | 0.5195 | 37.5405 | 25.4605 | 0.25663 | 0.19677 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L6 | 18 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 131.3 | 90.15 | 0.5195 | 130.7805 | 89.6305 | 0.8742 | 0.72192 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L6 | 18 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 135.1 | 91.91 | 0.5195 | 134.5805 | 91.3905 | 0.91754 | 0.70993 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L6 | 18 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.04 | 15.63 | 0.62633 | 24.4137 | 15.0037 | 0.19786 | 0.066585 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR113 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 6 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 23.75 | 15.33 | 0.62633 | 23.1237 | 14.7037 | 0.17705 | 0.082114 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR114 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 6 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.712 | 3.215 | 0.62633 | 4.0857 | 2.5887 | 0.016738 | 0.0075241 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR123; poor replication (47% range) but low values | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 6 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.478 | 6.044 | 0.62633 | 7.8517 | 5.4177 | 0.027215 | 0.023562 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR124; poor replication (47% range) but low values | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 96.98 | 62.34 | 0.5195 | 96.4605 | 61.8205 | 0.7359 | 0.36499 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L6 | 18 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 94.35 | 60.91 | 0.5195 | 93.8305 | 60.3905 | 0.71041 | 0.36502 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L6 | 18 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.85 | 51.28 | 0.5395 | 80.3105 | 50.7405 | 0.62819 | 0.27539 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220802 | L6 | 18 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 77.58 | 49.68 | 0.5395 | 77.0405 | 49.1405 | 0.59272 | 0.28237 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220802 | L6 | 18 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.86 | 10.69 | 0.62633 | 15.2337 | 10.0637 | 0.10871 | 0.068669 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR111 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 14.62 | 10.31 | 0.62633 | 13.9937 | 9.6837 | 0.090317 | 0.079782 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR112 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 10 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.048 | 4.825 | 0.62633 | 6.4217 | 4.1987 | 0.024901 | 0.014523 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR119 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 10 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 6.161 | 4.363 | 0.62633 | 5.5347 | 3.7367 | 0.02014 | 0.014945 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR120 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 42.91 | 27.69 | 0.5195 | 42.3905 | 27.1705 | 0.32334 | 0.16051 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L6 | 18 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 43.92 | 28.23 | 0.5195 | 43.4005 | 27.7105 | 0.33332 | 0.16014 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L6 | 18 | 17 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.543 | 6.067 | 0.62633 | 7.9167 | 5.4407 | 0.051885 | 0.043683 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR115; poor replication (59% range) but low values | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 17 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 11.73 | 7.173 | 0.62633 | 11.1037 | 6.5467 | 0.095493 | 0.019503 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR116; poor replication (59% range) but low values | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 17 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.241 | 2.133 | 0.62633 | 2.6147 | 1.5067 | 0.012389 | 0.0017325 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR121 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 17 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.034 | 2.117 | 0.62633 | 2.4077 | 1.4907 | 0.010235 | 0.0037114 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR122 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.11 | 16.09 | 0.5195 | 24.5905 | 15.5705 | 0.19162 | 0.085654 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L6 | 18 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.7 | 16.26 | 0.5195 | 25.1805 | 15.7405 | 0.20055 | 0.079758 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L6 | 18 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.68 | 13.32 | 0.5395 | 20.1405 | 12.7805 | 0.15636 | 0.071235 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220802 | L6 | 18 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.57 | 13.43 | 0.5395 | 20.0305 | 12.8905 | 0.15168 | 0.077868 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220802 | L6 | 18 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 5.554 | 3.281 | 0.62633 | 4.9277 | 2.6547 | 0.047794 | -0.001004 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | DF2 | TAR109; negative phaeo; poor replication (63% range) but low values | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.5 | 6.12 | 0.62633 | 9.8737 | 5.4937 | 0.092097 | 0.0047314 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR110; poor replication (63% range) but low values; note when analyzing not sure if enough volume to fill tube for fluorometer reading | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 2.786 | 1.053 | 0.62633 | 2.1597 | 0.42667 | 0.019412 | -0.015406 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR117; negative phaeo; very poor replication (161% range) | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L6 | 18 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 1.566 | 1.379 | 0.62633 | 0.93967 | 0.75267 | 0.0020871 | 0.0049545 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 | 1 | TAR118; could Rb accidentally be an Ra reading? Seems like chl would be this low relative to other size frac from same niksin; very poor replication (161% range) | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L7 | 21 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 79.76 | 53.72 | 0.5195 | 79.2405 | 53.2005 | 0.5532 | 0.39419 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L7 | 21 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 80.56 | 52.91 | 0.5195 | 80.0405 | 52.3905 | 0.58741 | 0.34556 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L7 | 21 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 115.7 | 77.84 | 0.5195 | 115.1805 | 77.3205 | 0.80431 | 0.5726 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L7 | 21 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 115.8 | 77.89 | 0.5195 | 115.2805 | 77.3705 | 0.80537 | 0.57243 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L7 | 21 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 24.96 | 15.8 | 0.62633 | 24.3337 | 15.1737 | 0.19261 | 0.074838 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR127 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L7 | 21 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.64 | 13.43 | 0.62633 | 21.0137 | 12.8037 | 0.17204 | 0.052861 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR128 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L7 | 21 | 10 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.04 | 13.53 | 0.62633 | 20.4137 | 12.9037 | 0.084124 | 0.037037 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR133; somewhat poor replication (36% range) but within bio variation | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L7 | 21 | 10 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 29.04 | 18.19 | 0.62633 | 28.4137 | 17.5637 | 0.12154 | 0.043379 | 20220519 | 11-40 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR134; somewhat poor replication (36% range) but within bio variation | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L7 | 21 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 71.07 | 46.63 | 0.5195 | 70.5505 | 46.1105 | 0.51921 | 0.30192 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L7 | 21 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 69.98 | 45.81 | 0.5195 | 69.4605 | 45.2905 | 0.51348 | 0.29305 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L7 | 21 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.04 | 13.19 | 0.62633 | 20.4137 | 12.5637 | 0.16506 | 0.056381 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR129 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L7 | 21 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.67 | 9.664 | 0.62633 | 15.0437 | 9.0377 | 0.12629 | 0.033007 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR130 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L7 | 21 | 14 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.51 | 6.768 | 0.62633 | 9.8837 | 6.1417 | 0.04184 | 0.015723 | 20220519 | 11-40 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR135
EN687 | 20220802 | L7 | 21 | 14 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 10.44 | 6.555 | 0.62633 | 9.8137 | 5.9287 | 0.043439 | 0.012128 | 20220519 | 11-40 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR136
EN687 | 20220802 | L7 | 21 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.68 | 16.71 | 0.5195 | 25.1605 | 16.1905 | 0.19056 | 0.097757 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L7 | 21 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 25.09 | 16.98 | 0.5195 | 24.5705 | 16.4605 | 0.17229 | 0.12083 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L7 | 21 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 32 | 21.4 | 0.5395 | 31.4605 | 20.8605 | 0.22519 | 0.14629 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220802 | L7 | 21 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.06 | 22 | 0.5195 | 33.5405 | 21.4805 | 0.25621 | 0.12632 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220802 | L7 | 21 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.46 | 5.061 | 0.62633 | 6.8337 | 4.4347 | 0.050443 | 0.02772 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR125 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L7 | 21 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 9.338 | 6.242 | 0.62633 | 8.7117 | 5.6157 | 0.065099 | 0.03388 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR126 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L7 | 21 | 18 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 4.593 | 3.315 | 0.62633 | 3.9667 | 2.6887 | 0.014264 | 0.01089 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR131; poor replication (42% range) but values so low basically within noise of flourometer | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L7 | 21 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 3.83 | 2.998 | 0.62633 | 3.2037 | 2.3717 | 0.0093197 | 0.012949 | 20220519 | 11-39 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR132; poor replication (42% range) but values so low basically within noise of flourometer | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L3 | 22 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 133.7 | 91.06 | 0.5195 | 133.1805 | 90.5405 | 0.90586 | 0.70647 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L3 | 22 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 132.3 | 87.85 | 0.5195 | 131.7805 | 87.3305 | 0.94431 | 0.61086 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L3 | 22 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 164.3 | 107.8 | 0.5195 | 163.7805 | 107.2805 | 1.2003 | 0.71013 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L3 | 22 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 165.7 | 107.1 | 0.5195 | 165.1805 | 106.5805 | 1.2449 | 0.65305 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L3 | 22 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 60.13 | 39.43 | 0.62633 | 59.5037 | 38.8037 | 0.43525 | 0.24868 | 20220519 | 11-40 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR139; poor replication (52 % range) | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L3 | 22 | 7 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 33.12 | 21 | 0.62633 | 32.4937 | 20.3737 | 0.25398 | 0.1039 | 20220519 | 11-40 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 | DF2 | TAR140; poor replication (52% range). This rep seems low, compared to other size fractions in this cast so toss and assume other rep is fine | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L3 | 22 | 7 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 42.03 | 25.81 | 0.62633 | 41.4037 | 25.1837 | 0.18169 | 0.054775 | 20220519 | 11-40 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR143 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L3 | 22 | 7 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 46.97 | 28.62 | 0.62633 | 46.3437 | 27.9937 | 0.20555 | 0.057301 | 20220519 | 11-40 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR144 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L3 | 22 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 118.6 | 75.25 | 0.5195 | 118.0805 | 74.7305 | 0.92094 | 0.40985 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L3 | 22 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 119.3 | 74.79 | 0.5195 | 118.7805 | 74.2705 | 0.94559 | 0.37701 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L3 | 22 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.28 | 24.48 | 0.5195 | 37.7605 | 23.9605 | 0.29317 | 0.13351 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L3 | 22 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.98 | 24.13 | 0.5195 | 38.4605 | 23.6105 | 0.31548 | 0.10497 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220802 | L3 | 22 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.22 | 17.77 | 0.5395 | 27.6805 | 17.2305 | 0.222 | 0.084835 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220802 | L3 | 22 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 29.4 | 18.39 | 0.81783 | 28.5822 | 17.5722 | 0.2339 | 0.079023 | 20220519 | 13-00 | 20230214 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220802 | L3 | 22 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 8.217 | 5.223 | 0.62633 | 7.5907 | 4.5967 | 0.062954 | 0.018064 | 20220519 | 11-40 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR137 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L3 | 22 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.624 | 4.683 | 0.62633 | 6.9977 | 4.0567 | 0.06184 | 0.0096611 | 20220519 | 11-40 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR138 | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L3 | 22 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 7.301 | 4.385 | 0.62633 | 6.6747 | 3.7587 | 0.032664 | 0.0026286 | 20220519 | 11-40 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR141; poor replication (62% range) but low values | No blank reading after running, so only 3 values for verage
EN687 | 20220802 | L3 | 22 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 13.65 | 7.983 | 0.62633 | 13.0237 | 7.3567 | 0.063479 | 0.005597 | 20220519 | 11-40 | 20220816 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2 | TAR142;poor replication (62% range) but low values | No blank reading after running, so only 3 values for verage
EN687 | 20220803 | MVCO | 26 | 1 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 245.4 | 150.3 | 0.5195 | 244.8805 | 149.7805 | 4.0407 | 1.2939 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220803 | MVCO | 26 | 1 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 233.1 | 142 | 0.5195 | 232.5805 | 141.4805 | 3.8707 | 1.1682 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220803 | MVCO | 26 | 3 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 150 | 92.23 | 0.5195 | 149.4805 | 91.7105 | 2.4546 | 0.81177 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220803 | MVCO | 26 | 3 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 149.7 | 91.99 | 0.5195 | 149.1805 | 91.4705 | 2.452 | 0.80577 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220803 | MVCO | 26 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 106.2 | 64.96 | 0.5195 | 105.6805 | 64.4405 | 0.87612 | 0.27143 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220803 | MVCO | 26 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 106.9 | 66.57 | 0.5195 | 106.3805 | 66.0505 | 0.85678 | 0.31943 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 | DF2
EN687 | 20220803 | MVCO | 26 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 60.99 | 38.8 | 0.5195 | 60.4705 | 38.2805 | 0.47141 | 0.21028 | 20220519 | 11-64 | 20221105 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN687 | 20220803 | MVCO | 26 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 61.98 | 39.72 | 0.5395 | 61.4405 | 39.1805 | 0.4729 | 0.22482 | 20220519 | 13-06 | 20230221 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230111 | L1 | 1 | 2 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 158.3 | 96.25 | 0.70967 | 157.5903 | 95.5403 | 2.6364 | 0.76632 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230111 | L1 | 1 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 140.8 | 86.76 | 0.70967 | 140.0903 | 86.0503 | 2.2961 | 0.76866 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230111 | L1 | 1 | 2 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 95.55 | 62.76 | 0.72233 | 94.8277 | 62.0377 | 0.6966 | 0.40816 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230111 | L1 | 1 | 2 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 243.4 | 143.3 | 0.7955 | 242.6045 | 142.5045 | 2.1048 | 0.40693 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR5
EN695 | 20230111 | L1 | 1 | 2 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 242.8 | 143.3 | 0.7955 | 242.0045 | 142.5045 | 2.0922 | 0.41955 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR6
EN695 | 20230111 | L1 | 1 | 2 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 361.6 | 210.3 | 0.7955 | 360.8045 | 209.5045 | 1.6917 | 0.27187 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR11
EN695 | 20230111 | L1 | 1 | 2 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 361.4 | 210 | 0.7955 | 360.6045 | 209.2045 | 1.6928 | 0.26794 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR12
EN695 | 20230111 | L1 | 1 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 256.8 | 158.8 | 0.72233 | 256.0777 | 158.0777 | 2.0819 | 0.73308 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230111 | L1 | 1 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 295.2 | 178.7 | 0.3865 | 294.8135 | 178.3135 | 2.475 | 0.70042 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230111 | L1 | 1 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 100.5 | 65.01 | 0.3865 | 100.1135 | 64.6235 | 0.75396 | 0.39685 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230111 | L1 | 1 | 8 |  |  | b | 215 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 67.84 | 44.32 | 0.3865 | 67.4535 | 43.9335 | 0.67397 | 0.38131 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230111 | L1 | 1 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 223.9 | 132.3 | 0.7955 | 223.1045 | 131.5045 | 1.9261 | 0.39178 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR3; logsheet says small leak but they did not compensate on sheet for vol filt - maybe we should reduce a couple ml?? But how many….
EN695 | 20230111 | L1 | 1 | 8 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 226.9 | 134.2 | 0.7955 | 226.1045 | 133.4045 | 1.9426 | 0.40077 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR4
EN695 | 20230111 | L1 | 1 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 349.7 | 201.7 | 0.7955 | 348.9045 | 200.9045 | 1.6578 | 0.22858 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR9
EN695 | 20230111 | L1 | 1 | 8 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 347.5 | 200.1 | 0.7955 | 346.7045 | 199.3045 | 1.6511 | 0.22028 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR10
EN695 | 20230111 | L1 | 1 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 311.1 | 190 | 0.70967 | 310.3903 | 189.2903 | 2.5727 | 0.79817 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  |  | poor replication
EN695 | 20230111 | L1 | 1 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 228.9 | 141.8 | 0.70967 | 228.1903 | 141.0903 | 1.8504 | 0.66214 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  |  | poor replication
EN695 | 20230111 | L1 | 1 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 83.95 | 54.51 | 0.70967 | 83.2403 | 53.8003 | 0.62543 | 0.33264 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230111 | L1 | 1 | 15 |  |  | b | 190 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 61.82 | 40.83 | 0.3865 | 61.4335 | 40.4435 | 0.68061 | 0.41866 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230111 | L1 | 1 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 192.6 | 115.8 | 0.7955 | 191.8045 | 115.0045 | 1.6149 | 0.41216 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR1 | poor replication
EN695 | 20230111 | L1 | 1 | 15 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 252.1 | 150.2 | 0.7955 | 251.3045 | 149.4045 | 2.1426 | 0.4907 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR2 | poor replication
EN695 | 20230111 | L1 | 1 | 15 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 331.4 | 191.5 | 0.7955 | 330.6045 | 190.7045 | 1.5671 | 0.22354 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR7
EN695 | 20230111 | L1 | 1 | 15 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 326.8 | 188.8 | 0.7955 | 326.0045 | 188.0045 | 1.5402 | 0.21868 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR8
EN695 | 20230112 | L6 | 2 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 48.53 | 36.01 | 0.72233 | 47.8077 | 35.2877 | 0.26598 | 0.36242 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae>chl, depth = 86m
EN695 | 20230112 | L6 | 2 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 54.71 | 38.93 | 0.70967 | 54.0003 | 38.2203 | 0.33524 | 0.34539 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae>chl, depth = 86m
EN695 | 20230112 | L6 | 2 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 62.3 | 43.76 | 0.3865 | 61.9135 | 43.3735 | 0.39387 | 0.37852 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L6 | 2 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 68.46 | 47.2 | 0.3865 | 68.0735 | 46.8135 | 0.45165 | 0.38199 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L6 | 2 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 73.15 | 49.71 | 0.70967 | 72.4403 | 49.0003 | 0.49797 | 0.37462 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L6 | 2 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 69.62 | 47.4 | 0.72233 | 68.8977 | 46.6777 | 0.47205 | 0.35918 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L6 | 2 | 4 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 29.23 | 21.65 | 0.70967 | 28.5203 | 20.9403 | 0.16103 | 0.21187 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae>chl, depth = 50m
EN695 | 20230112 | L6 | 2 | 4 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 27.55 | 20.23 | 0.3865 | 27.1635 | 19.8435 | 0.15551 | 0.19786 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae>chl, depth = 50m
EN695 | 20230112 | L6 | 2 | 4 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.79 | 41.93 | 0.7955 | 65.9945 | 41.1345 | 0.52273 | 0.20229 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR17
EN695 | 20230112 | L6 | 2 | 4 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 66.73 | 42.74 | 0.7955 | 65.9345 | 41.9445 | 0.50443 | 0.23486 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR18
EN695 | 20230112 | L6 | 2 | 4 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 95.53 | 56.99 | 0.7955 | 94.7345 | 56.1945 | 0.43092 | 0.095762 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR23
EN695 | 20230112 | L6 | 2 | 4 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 102.2 | 61.34 | 0.7955 | 101.4045 | 60.5445 | 0.45686 | 0.11059 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR24
EN695 | 20230112 | L6 | 2 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 193.4 | 120.5 | 0.70967 | 192.6903 | 119.7903 | 1.5487 | 0.5845 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L6 | 2 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 215.7 | 133.1 | 0.72233 | 214.9777 | 132.3777 | 1.7548 | 0.60258 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L6 | 2 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 305.8 | 188.1 | 0.3865 | 305.4135 | 187.7135 | 2.5005 | 0.84232 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L6 | 2 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 307.1 | 187.5 | 0.70967 | 306.3903 | 186.7903 | 2.5408 | 0.78552 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L6 | 2 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 94.79 | 63.51 | 0.3865 | 94.4035 | 63.1235 | 0.66452 | 0.45957 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L6 | 2 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 102.1 | 66.78 | 0.3865 | 101.7135 | 66.3935 | 0.75035 | 0.43198 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L6 | 2 | 11 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 225.5 | 133.7 | 0.7955 | 224.7045 | 132.9045 | 1.9237 | 0.41085 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR15
EN695 | 20230112 | L6 | 2 | 11 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 242.7 | 143.5 | 0.7955 | 241.9045 | 142.7045 | 2.0859 | 0.42938 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR16
EN695 | 20230112 | L6 | 2 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 374.5 | 217.7 | 0.7955 | 373.7045 | 216.9045 | 1.7564 | 0.28024 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR21
EN695 | 20230112 | L6 | 2 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 390.9 | 226.3 | 0.7955 | 390.1045 | 225.5045 | 1.8438 | 0.27362 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR22
EN695 | 20230112 | L6 | 2 | 18 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 153 | 94.76 | 0.72233 | 152.2777 | 94.0377 | 2.4745 | 0.87468 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  |  | poor replication
EN695 | 20230112 | L6 | 2 | 18 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 199.9 | 121.2 | 0.72233 | 199.1777 | 120.4777 | 3.3439 | 0.94704 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  |  | poor replication
EN695 | 20230112 | L6 | 2 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 93.4 | 61.4 | 0.70967 | 92.6903 | 60.6903 | 0.67982 | 0.40095 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L6 | 2 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 96.92 | 63.85 | 0.70967 | 96.2103 | 63.1403 | 0.70255 | 0.42184 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L6 | 2 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 250 | 147.9 | 0.7955 | 249.2045 | 147.1045 | 2.1468 | 0.44596 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR13 | poor replication
EN695 | 20230112 | L6 | 2 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 200.4 | 120 | 0.7955 | 199.6045 | 119.2045 | 1.6906 | 0.41049 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR14 | poor replication
EN695 | 20230112 | L6 | 2 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 341 | 199.1 | 0.7955 | 340.2045 | 198.3045 | 1.5895 | 0.2725 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR19
EN695 | 20230112 | L6 | 2 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 337.8 | 198.1 | 0.7955 | 337.0045 | 197.3045 | 1.5592 | 0.28671 | 20220519 | 13-35 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR20
EN695 | 20230112 | L9 | 3 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 17.49 | 13.12 | 0.70967 | 16.7803 | 12.4103 | 0.092838 | 0.12816 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae>chl, depth = 93m
EN695 | 20230112 | L9 | 3 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 18.48 | 13.61 | 0.72233 | 17.7577 | 12.8877 | 0.10346 | 0.12604 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae>chl, depth = 93m
EN695 | 20230112 | L9 | 3 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 96.46 | 62.2 | 0.3865 | 96.0735 | 61.8135 | 0.72783 | 0.37294 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L9 | 3 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 100.6 | 66.69 | 0.3865 | 100.2135 | 66.3035 | 0.7204 | 0.46033 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L9 | 3 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 132.3 | 86.26 | 0.70967 | 131.5903 | 85.5503 | 0.97809 | 0.54538 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L9 | 3 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 144.4 | 92.12 | 0.70967 | 143.6903 | 91.4103 | 1.1107 | 0.51717 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L9 | 3 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.67 | 35.51 | 0.70967 | 51.9603 | 34.8003 | 0.36455 | 0.25517 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L9 | 3 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 52.54 | 35.81 | 0.3865 | 52.1535 | 35.4235 | 0.35542 | 0.2754 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L9 | 3 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 87.27 | 56.17 | 0.7955 | 86.4745 | 55.3745 | 0.65393 | 0.32207 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR27
EN695 | 20230112 | L9 | 3 | 16 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 93.49 | 58.77 | 0.7955 | 92.6945 | 57.9745 | 0.72757 | 0.29079 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR28
EN695 | 20230112 | L9 | 3 | 16 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 115.3 | 69.97 | 0.7955 | 114.5045 | 69.1745 | 0.50777 | 0.14175 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR33
EN695 | 20230112 | L9 | 3 | 16 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 109.7 | 67.01 | 0.7955 | 108.9045 | 66.2145 | 0.47819 | 0.14353 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR34
EN695 | 20230112 | L9 | 3 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 132.3 | 85.81 | 0.3865 | 131.9135 | 85.4235 | 0.98765 | 0.53356 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L9 | 3 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 137 | 87.46 | 0.72233 | 136.2777 | 86.7377 | 1.0524 | 0.49217 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L9 | 3 | 17 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 140.6 | 86.35 | 0.7955 | 139.8045 | 85.5545 | 1.1407 | 0.36724 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR29
EN695 | 20230112 | L9 | 3 | 17 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 122.9 | 76.57 | 0.7955 | 122.1045 | 75.7745 | 0.97417 | 0.36139 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR30
EN695 | 20230112 | L9 | 3 | 17 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 183.4 | 117.1 | 0.7955 | 182.6045 | 116.3045 | 0.74132 | 0.34876 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR35
EN695 | 20230112 | L9 | 3 | 17 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 161.1 | 97.46 | 0.7955 | 160.3045 | 96.6645 | 0.71287 | 0.19477 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR36
EN695 | 20230112 | L9 | 3 | 21 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 99.29 | 65.93 | 0.72233 | 98.5677 | 65.2077 | 0.70871 | 0.4525 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L9 | 3 | 21 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 103.6 | 69.2 | 0.70967 | 102.8903 | 68.4903 | 0.7308 | 0.48886 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L9 | 3 | 21 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 41.99 | 28.54 | 0.70967 | 41.2803 | 27.8303 | 0.28574 | 0.20986 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L9 | 3 | 21 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 41.29 | 28.3 | 0.3865 | 40.9035 | 27.9135 | 0.27596 | 0.22112 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L9 | 3 | 21 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 76.58 | 48.51 | 0.7955 | 75.7845 | 47.7145 | 0.59022 | 0.25077 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR25
EN695 | 20230112 | L9 | 3 | 21 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 94.14 | 58.75 | 0.7955 | 93.3445 | 57.9545 | 0.74414 | 0.27734 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR26
EN695 | 20230112 | L9 | 3 | 21 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 162.2 | 98.26 | 0.7955 | 161.4045 | 97.4645 | 0.71623 | 0.19893 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR31
EN695 | 20230112 | L9 | 3 | 21 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 141.6 | 86.5 | 0.7955 | 140.8045 | 85.7045 | 0.61721 | 0.18753 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR32
EN695 | 20230112 | L10 | 4 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 55.45 | 36.03 | 0.70967 | 54.7403 | 35.3203 | 0.41256 | 0.21641 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L10 | 4 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 74.15 | 47.17 | 0.3865 | 73.7635 | 46.7835 | 0.57317 | 0.25994 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L10 | 4 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 136.3 | 84.67 | 0.72233 | 135.5777 | 83.9477 | 1.0968 | 0.39808 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  |  | poor replication
EN695 | 20230112 | L10 | 4 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 194.6 | 119.4 | 0.3865 | 194.2135 | 119.0135 | 1.5976 | 0.5218 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  |  | poor replication
EN695 | 20230112 | L10 | 4 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 46.65 | 30.66 | 0.70967 | 45.9403 | 29.9503 | 0.3397 | 0.19365 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L10 | 4 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 28.87 | 19.67 | 0.3865 | 28.4835 | 19.2835 | 0.19545 | 0.14795 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L10 | 4 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 86.72 | 53.63 | 0.7955 | 85.9245 | 52.8345 | 0.69578 | 0.23546 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR41
EN695 | 20230112 | L10 | 4 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 95.56 | 58.53 | 0.7955 | 94.7645 | 57.7345 | 0.77862 | 0.23898 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR42
EN695 | 20230112 | L10 | 4 | 12 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 304.2 | 180.9 | 0.7955 | 303.4045 | 180.1045 | 1.3786 | 0.3094 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR47
EN695 | 20230112 | L10 | 4 | 12 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 253 | 150.5 | 0.7955 | 252.2045 | 149.7045 | 1.1461 | 0.25704 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR48
EN695 | 20230112 | L10 | 4 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 142.4 | 88.99 | 0.70967 | 141.6903 | 88.2803 | 1.1347 | 0.43742 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L10 | 4 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 138.7 | 85.91 | 0.70967 | 137.9903 | 85.2003 | 1.1215 | 0.39575 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | dropped on floor, looks ok
EN695 | 20230112 | L10 | 4 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 35.87 | 23.84 | 0.72233 | 35.1477 | 23.1177 | 0.25557 | 0.15611 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L10 | 4 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 30.15 | 20.51 | 0.70967 | 29.4403 | 19.8003 | 0.2048 | 0.14781 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L10 | 4 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 121.5 | 74.26 | 0.7955 | 120.7045 | 73.4645 | 0.99331 | 0.30154 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR39
EN695 | 20230112 | L10 | 4 | 15 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 115 | 70.82 | 0.7955 | 114.2045 | 70.0245 | 0.9258 | 0.30422 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR40
EN695 | 20230112 | L10 | 4 | 15 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 189.4 | 113.1 | 0.7955 | 188.6045 | 112.3045 | 0.85468 | 0.19982 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR45
EN695 | 20230112 | L10 | 4 | 15 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 210.5 | 126.5 | 0.7955 | 209.7045 | 125.7045 | 0.94093 | 0.23938 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR46
EN695 | 20230112 | L10 | 4 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 122.5 | 77.01 | 0.3865 | 122.1135 | 76.6235 | 0.9664 | 0.3981 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L10 | 4 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 163.1 | 100.4 | 0.3865 | 162.7135 | 100.0135 | 1.332 | 0.44901 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L10 | 4 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.13 | 20.79 | 0.70967 | 30.4203 | 20.0803 | 0.21967 | 0.13792 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L10 | 4 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 27.01 | 18.34 | 0.70967 | 26.3003 | 17.6303 | 0.18419 | 0.12977 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L10 | 4 | 20 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 103.7 | 64.09 | 0.7955 | 102.9045 | 63.2945 | 0.83287 | 0.28273 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR37 | poor replication; could 5's and 20's gotten mixed up?
EN695 | 20230112 | L10 | 4 | 20 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 136.6 | 82.33 | 0.7955 | 135.8045 | 81.5345 | 1.1411 | 0.29596 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR38 | poor replication; could 5's and 20's gotten mixed up?
EN695 | 20230112 | L10 | 4 | 20 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 200.6 | 120.9 | 0.7955 | 199.8045 | 120.1045 | 0.89276 | 0.23497 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR43 | could 5's and 20's gotten mixed up?
EN695 | 20230112 | L10 | 4 | 20 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 238.1 | 143.9 | 0.7955 | 237.3045 | 143.1045 | 1.0514 | 0.28746 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR44 | could 5's and 20's gotten mixed up?
EN695 | 20230112 | L11 | 5 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 15.6 | 12.51 | 0.70967 | 14.8903 | 11.8003 | 0.065645 | 0.14449 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L11 | 5 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 21.05 | 16.16 | 0.70967 | 20.3403 | 15.4503 | 0.10388 | 0.17125 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L11 | 5 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 138.4 | 86.44 | 0.3865 | 138.0135 | 86.0535 | 1.1039 | 0.42857 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L11 | 5 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 129.1 | 79.91 | 0.70967 | 128.3903 | 79.2003 | 1.045 | 0.36538 | 20220519 | 13-42 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L11 | 5 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 214 | 134.6 | 0.70967 | 213.2903 | 133.8903 | 1.6868 | 0.6975 | 20220519 | 13-43 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L11 | 5 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 242.6 | 153.3 | 0.70967 | 241.8903 | 152.5903 | 1.8971 | 0.82019 | 20220519 | 13-43 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L11 | 5 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 82.26 | 53.18 | 0.70967 | 81.5503 | 52.4703 | 0.61779 | 0.3166 | 20220519 | 13-43 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L11 | 5 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.44 | 23.75 | 0.3865 | 34.0535 | 23.3635 | 0.2271 | 0.18895 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L11 | 5 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 184.4 | 115.9 | 0.7955 | 183.6045 | 115.1045 | 1.4403 | 0.58844 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR53
EN695 | 20230112 | L11 | 5 | 9 |  |  | b | 291 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 183.4 | 116.3 | 0.7955 | 182.6045 | 115.5045 | 1.4206 | 0.62922 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR54; logsheet tiny leak 2ml
EN695 | 20230112 | L11 | 5 | 9 |  |  | a | 545 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 266.4 | 161.4 | 0.7955 | 265.6045 | 160.6045 | 1.187 | 0.33489 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR59; small leak ~6ml
EN695 | 20230112 | L11 | 5 | 9 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 284 | 173.9 | 0.7955 | 283.2045 | 173.1045 | 1.2311 | 0.39138 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR60
EN695 | 20230112 | L11 | 5 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 274.3 | 168.6 | 0.70967 | 273.5903 | 167.8903 | 2.2455 | 0.74424 | 20220519 | 13-43 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L11 | 5 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 231 | 143.1 | 0.3865 | 230.6135 | 142.7135 | 1.8674 | 0.67405 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L11 | 5 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 38.15 | 25.99 | 0.3865 | 37.7635 | 25.6035 | 0.25833 | 0.19761 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L11 | 5 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 37.28 | 25.13 | 0.70967 | 36.5703 | 24.4203 | 0.25812 | 0.17676 | 20220519 | 13-43 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L11 | 5 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 147.6 | 93.89 | 0.7955 | 146.8045 | 93.0945 | 1.1294 | 0.51149 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR51
EN695 | 20230112 | L11 | 5 | 12 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 132.9 | 84.93 | 0.7955 | 132.1045 | 84.1345 | 1.0052 | 0.47265 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR52
EN695 | 20230112 | L11 | 5 | 12 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 252.9 | 155.8 | 0.7955 | 252.1045 | 155.0045 | 1.0877 | 0.36776 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR57 | poor replication
EN695 | 20230112 | L11 | 5 | 12 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 324 | 197.4 | 0.7955 | 323.2045 | 196.6045 | 1.4181 | 0.42792 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR58 | poor replication; higher value than >5s
EN695 | 20230112 | L11 | 5 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 283 | 173.2 | 0.70967 | 282.2903 | 172.4903 | 2.3326 | 0.73906 | 20220519 | 13-43 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L11 | 5 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 255.5 | 158.2 | 0.72233 | 254.7777 | 157.4777 | 2.0671 | 0.73727 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L11 | 5 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 31.24 | 21.37 | 0.70967 | 30.5303 | 20.6603 | 0.20968 | 0.15823 | 20220519 | 13-43 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  |  | poor replication
EN695 | 20230112 | L11 | 5 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 87.66 | 56.58 | 0.70967 | 86.9503 | 55.8703 | 0.66027 | 0.33466 | 20220519 | 13-43 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  |  | tissue capsule said C5N19 chl b, but there is another tissue capsule for that same sample, but the chl reading is similar to to replication a
EN695 | 20230112 | L11 | 5 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 226.3 | 138.3 | 0.7955 | 225.5045 | 137.5045 | 1.8504 | 0.57323 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR49 | poor replication
EN695 | 20230112 | L11 | 5 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 173.2 | 108.6 | 0.7955 | 172.4045 | 107.8045 | 1.3583 | 0.54178 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR50 | poor replication
EN695 | 20230112 | L11 | 5 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 150.3 | 90.93 | 0.7955 | 149.5045 | 90.1345 | 1.9951 | 0.54387 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR55
EN695 | 20230112 | L11 | 5 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 165.1 | 98.9 | 0.7955 | 164.3045 | 98.1045 | 2.2166 | 0.53691 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR56
EN695 | 20230112 | L8 | 6 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 32.51 | 24.86 | 0.72233 | 31.7877 | 24.1377 | 0.16252 | 0.26732 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L8 | 6 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 34.21 | 24.87 | 0.72233 | 33.4877 | 24.1477 | 0.19842 | 0.2316 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L8 | 6 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 39.78 | 29.77 | 0.3865 | 39.3935 | 29.3835 | 0.21266 | 0.3106 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae>chl, depth = 50m
EN695 | 20230112 | L8 | 6 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 47.65 | 33.98 | 0.72233 | 46.9277 | 33.2577 | 0.29041 | 0.30184 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | phae>chl, depth = 50m
EN695 | 20230112 | L8 | 6 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 19.7 | 15.49 | 0.72233 | 18.9777 | 14.7677 | 0.089439 | 0.17354 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | phae>chl, depth = 50m
EN695 | 20230112 | L8 | 6 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 20.14 | 15.19 | 0.3865 | 19.7535 | 14.8035 | 0.10516 | 0.15846 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | phae>chl, depth = 50m
EN695 | 20230112 | L8 | 6 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 36.02 | 25.89 | 0.7955 | 35.2245 | 25.0945 | 0.213 | 0.2293 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR65
EN695 | 20230112 | L8 | 6 | 8 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 45.44 | 32.12 | 0.7955 | 44.6445 | 31.3245 | 0.28008 | 0.27203 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR66
EN695 | 20230112 | L8 | 6 | 8 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 49.43 | 33.44 | 0.7955 | 48.6345 | 32.6445 | 0.17879 | 0.12717 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR71
EN695 | 20230112 | L8 | 6 | 8 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 43.46 | 30.88 | 0.7955 | 42.6645 | 30.0845 | 0.14066 | 0.14131 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR72
EN695 | 20230112 | L8 | 6 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 189.9 | 122.9 | 0.72233 | 189.1777 | 122.1777 | 1.4234 | 0.75235 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L8 | 6 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 186.2 | 118.8 | 0.72233 | 185.4777 | 118.0777 | 1.4319 | 0.67084 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L8 | 6 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 309.7 | 192.2 | 0.72233 | 308.9777 | 191.4777 | 2.4962 | 0.9136 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L8 | 6 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 281 | 177 | 0.72233 | 280.2777 | 176.2777 | 2.2094 | 0.92972 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L8 | 6 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 117.9 | 76.95 | 0.70967 | 117.1903 | 76.2403 | 0.86996 | 0.48772 | 20220519 | 13-43 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L8 | 6 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 105 | 69.85 | 0.72233 | 104.2777 | 69.1277 | 0.74674 | 0.48428 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L8 | 6 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 207.4 | 125 | 0.7955 | 206.6045 | 124.2045 | 1.7326 | 0.45656 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR63 | poor replication
EN695 | 20230112 | L8 | 6 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 258.4 | 155.5 | 0.7955 | 257.6045 | 154.7045 | 2.1563 | 0.56118 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR64 | poor replication
EN695 | 20230112 | L8 | 6 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 387.1 | 226.7 | 0.7955 | 386.3045 | 225.9045 | 1.7967 | 0.32442 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR69
EN695 | 20230112 | L8 | 6 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 387.3 | 228.6 | 0.7955 | 386.5045 | 227.8045 | 1.7777 | 0.36131 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR70
EN695 | 20230112 | L8 | 6 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 269.7 | 170.4 | 0.72233 | 268.9777 | 169.6777 | 2.1096 | 0.91204 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L8 | 6 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 295.5 | 185.7 | 0.72233 | 294.7777 | 184.9777 | 2.3326 | 0.96143 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L8 | 6 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 124.2 | 81.2 | 0.70967 | 123.4903 | 80.4903 | 0.91351 | 0.51985 | 20220519 | 13-43 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L8 | 6 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 150.5 | 98.27 | 0.3865 | 150.1135 | 97.8835 | 1.1096 | 0.6335 | 20220519 | 13-46 | 20230720 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L8 | 6 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 245 | 147.7 | 0.7955 | 244.2045 | 146.9045 | 2.0459 | 0.54336 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR61 | poor replication
EN695 | 20230112 | L8 | 6 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 199.2 | 120.4 | 0.7955 | 198.4045 | 119.6045 | 1.6569 | 0.45118 | 20220519 | 13-36 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR62 | poor replication
EN695 | 20230112 | L8 | 6 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 385.9 | 226.2 | 0.7955 | 385.1045 | 225.4045 | 1.7889 | 0.32757 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR67
EN695 | 20230112 | L8 | 6 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 385.1 | 229.2 | 0.7955 | 384.3045 | 228.4045 | 1.74 | 0.39686 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR68
EN695 | 20230112 | L7 | 7 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 109.9 | 71.39 | 0.70967 | 109.1903 | 70.6803 | 0.81812 | 0.44055 | 20220519 | 13-43 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L7 | 7 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 112 | 74.38 | 0.72233 | 111.2777 | 73.6577 | 0.79921 | 0.51247 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L7 | 7 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 245.8 | 153.5 | 0.72233 | 245.0777 | 152.7777 | 1.9609 | 0.75979 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L7 | 7 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 264 | 166.1 | 0.72233 | 263.2777 | 165.3777 | 2.0798 | 0.8652 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L7 | 7 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 182.8 | 111.3 | 0.7955 | 182.0045 | 110.5045 | 1.5034 | 0.44428 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR77
EN695 | 20230112 | L7 | 7 | 10 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 184.3 | 111 | 0.7955 | 183.5045 | 110.2045 | 1.5413 | 0.40115 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR78
EN695 | 20230112 | L7 | 7 | 10 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 167.2 | 98.76 | 0.7955 | 166.4045 | 97.9645 | 0.76524 | 0.15294 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR84 | poor replication
EN695 | 20230112 | L7 | 7 | 10 |  |  | b | 547 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 329.3 | 193.7 | 0.7955 | 328.5045 | 192.9045 | 1.5273 | 0.29397 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR85;logsheet -4mL | poor replication
EN695 | 20230112 | L7 | 7 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 293.8 | 183.9 | 0.72233 | 293.0777 | 183.1777 | 2.3348 | 0.92725 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L7 | 7 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 305.6 | 188.5 | 0.64233 | 304.9577 | 187.8577 | 2.4877 | 0.85763 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L7 | 7 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 107.8 | 70.71 | 0.70967 | 107.0903 | 70.0003 | 0.78795 | 0.4586 | 20220519 | 13-43 | 20230627 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L7 | 7 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 101.2 | 67.97 | 0.72233 | 100.4777 | 67.2477 | 0.70595 | 0.49159 | 20220519 | 13-41 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L7 | 7 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 161.3 | 98.74 | 0.7955 | 160.5045 | 97.9445 | 1.3154 | 0.41089 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR75
EN695 | 20230112 | L7 | 7 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 160.3 | 96.91 | 0.7955 | 159.5045 | 96.1145 | 1.3284 | 0.35995 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR76
EN695 | 20230112 | L7 | 7 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 283.7 | 167.1 | 0.7955 | 282.9045 | 166.3045 | 1.3061 | 0.25543 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR81
EN695 | 20230112 | L7 | 7 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 270.4 | 159.5 | 0.7955 | 269.6045 | 158.7045 | 1.2423 | 0.24792 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR83; logsheet no tube 82
EN695 | 20230112 | L7 | 7 | 19 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 164 | 102.3 | 0.72233 | 163.2777 | 101.5777 | 2.6216 | 0.99621 | 20220519 | 13-41 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L7 | 7 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 310.8 | 193.1 | 0.64233 | 310.1577 | 192.4577 | 2.5005 | 0.9268 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L7 | 7 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 123.3 | 80.49 | 0.72233 | 122.5777 | 79.7677 | 0.90947 | 0.51102 | 20220519 | 13-41 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L7 | 7 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 133.6 | 86.86 | 0.72233 | 132.8777 | 86.1377 | 0.99296 | 0.54097 | 20220519 | 13-41 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230112 | L7 | 7 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 163.9 | 99.63 | 0.7955 | 163.1045 | 98.8345 | 1.3514 | 0.39062 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR73; logsheet lost 3mL
EN695 | 20230112 | L7 | 7 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 163.4 | 99.81 | 0.7955 | 162.6045 | 99.0145 | 1.3371 | 0.40809 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR74
EN695 | 20230112 | L7 | 7 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 277.6 | 163.7 | 0.7955 | 276.8045 | 162.9045 | 1.2759 | 0.25375 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR79
EN695 | 20230112 | L7 | 7 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 235.7 | 139.5 | 0.7955 | 234.9045 | 138.7045 | 1.0737 | 0.22398 | 20220519 | 13-37 | 20230607 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR80
EN695 | 20230113 | L5 | 8 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 95.93 | 64.57 | 0.72233 | 95.2077 | 63.8477 | 0.66622 | 0.47077 | 20220519 | 13-41 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L5 | 8 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 103.3 | 67.82 | 0.64233 | 102.6577 | 67.1777 | 0.75375 | 0.44254 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L5 | 8 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 205.7 | 128.4 | 0.72233 | 204.9777 | 127.6777 | 1.6422 | 0.63148 | 20220519 | 13-41 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L5 | 8 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 222.2 | 137.9 | 0.64233 | 221.5577 | 137.2577 | 1.7909 | 0.65337 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L5 | 8 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 191.1 | 113.8 | 0.71583 | 190.3842 | 113.0842 | 1.6254 | 0.3678 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR88
EN695 | 20230113 | L5 | 8 | 7 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 177.7 | 106.7 | 0.71583 | 176.9842 | 105.9842 | 1.4929 | 0.37512 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR89
EN695 | 20230113 | L5 | 8 | 7 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 316.3 | 181.9 | 0.71583 | 315.5842 | 181.1842 | 1.5028 | 0.1954 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR94
EN695 | 20230113 | L5 | 8 | 7 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 312 | 181.9 | 0.71583 | 311.2842 | 181.1842 | 1.4547 | 0.24348 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR95
EN695 | 20230113 | L5 | 8 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 240.8 | 147.9 | 0.64233 | 240.1577 | 147.2577 | 1.9736 | 0.64875 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L5 | 8 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 265 | 161.6 | 0.72233 | 264.2777 | 160.8777 | 2.1967 | 0.66822 | 20220519 | 13-41 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L5 | 8 | 11 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 200.4 | 118.1 | 0.71583 | 199.6842 | 117.3842 | 1.7246 | 0.3373 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR90
EN695 | 20230113 | L5 | 8 | 11 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 178.7 | 105.8 | 0.71583 | 177.9842 | 105.0842 | 1.5329 | 0.31931 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR91
EN695 | 20230113 | L5 | 8 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 273.5 | 159.3 | 0.71583 | 272.7842 | 158.5842 | 1.2792 | 0.20982 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR96
EN695 | 20230113 | L5 | 8 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 306.4 | 177.6 | 0.71583 | 305.6842 | 176.8842 | 1.4428 | 0.21811 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR97
EN695 | 20230113 | L5 | 8 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 233.5 | 145.2 | 0.64233 | 232.8577 | 144.5577 | 1.8759 | 0.69839 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L5 | 8 | 19 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 123.6 | 76.61 | 0.72233 | 122.8777 | 75.8877 | 1.9965 | 0.70625 | 20220519 | 13-41 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L5 | 8 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 88.86 | 57.25 | 0.64233 | 88.2177 | 56.6077 | 0.67153 | 0.33653 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L5 | 8 | 19 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 77.67 | 50.66 | 0.72233 | 76.9477 | 49.9377 | 0.58388 | 0.32101 | 20220519 | 13-41 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L5 | 8 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 199.8 | 116.3 | 0.71583 | 199.0842 | 115.5842 | 1.7557 | 0.28149 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR86
EN695 | 20230113 | L5 | 8 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 204.9 | 120.8 | 0.71583 | 204.1842 | 120.0842 | 1.7684 | 0.34819 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR87
EN695 | 20230113 | L5 | 8 | 19 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 318.9 | 185.2 | 0.71583 | 318.1842 | 184.4842 | 1.4922 | 0.23374 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR92
EN695 | 20230113 | L5 | 8 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 325.4 | 188.1 | 0.71583 | 324.6842 | 187.3842 | 1.5324 | 0.22069 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR93
EN695 | 20230113 | L4 | 9 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 167 | 111 | 0.64233 | 166.3577 | 110.3577 | 1.1897 | 0.77555 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L4 | 9 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 163.9 | 107.5 | 0.64233 | 163.2577 | 106.8577 | 1.1982 | 0.70473 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L4 | 9 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 157.6 | 101.3 | 0.72233 | 156.8777 | 100.5777 | 1.1961 | 0.59502 | 20220519 | 13-41 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L4 | 9 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 181 | 114.7 | 0.64233 | 180.3577 | 114.0577 | 1.4085 | 0.62262 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L4 | 9 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 246.5 | 154 | 0.72233 | 245.7777 | 153.2777 | 1.9651 | 0.76445 | 20220519 | 13-41 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L4 | 9 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 243.5 | 151.6 | 0.64233 | 242.8577 | 150.9577 | 1.9524 | 0.73588 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L4 | 9 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 108.8 | 71.76 | 0.64233 | 108.1577 | 71.1177 | 0.78689 | 0.47956 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  |  | poor replication
EN695 | 20230113 | L4 | 9 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 179.8 | 113.1 | 0.64233 | 179.1577 | 112.4577 | 1.417 | 0.58563 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | filter might not have been completely been on frit | poor replication
EN695 | 20230113 | L4 | 9 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 187.7 | 110.9 | 0.71583 | 186.9842 | 110.1842 | 1.6149 | 0.3272 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR102
EN695 | 20230113 | L4 | 9 | 6 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 188.1 | 113.7 | 0.71583 | 187.3842 | 112.9842 | 1.5644 | 0.42701 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR103
EN695 | 20230113 | L4 | 9 | 6 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 324.4 | 188.1 | 0.71583 | 323.6842 | 187.3842 | 1.524 | 0.23227 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR108
EN695 | 20230113 | L4 | 9 | 6 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 347.4 | 202.4 | 0.71583 | 346.6842 | 201.6842 | 1.6213 | 0.26902 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR109
EN695 | 20230113 | L4 | 9 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 268.4 | 166.1 | 0.64233 | 267.7577 | 165.4577 | 2.1733 | 0.77315 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L4 | 9 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 281.8 | 174.7 | 0.64233 | 281.1577 | 174.0577 | 2.2753 | 0.82433 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L4 | 9 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 92.36 | 61.53 | 0.64233 | 91.7177 | 60.8877 | 0.65496 | 0.42932 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  |  | poor replication
EN695 | 20230113 | L4 | 9 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 154.9 | 99.35 | 0.72233 | 154.1777 | 98.6277 | 1.1801 | 0.57623 | 20220519 | 13-40 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  |  | poor replication; cast number and niskin number was not written on tissue capsule, but this was the only missing data, so I assume it is this one
EN695 | 20230113 | L4 | 9 | 11 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 197.6 | 119.4 | 0.71583 | 196.8842 | 118.6842 | 1.6387 | 0.44605 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR100
EN695 | 20230113 | L4 | 9 | 11 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 190 | 112.3 | 0.71583 | 189.2842 | 111.5842 | 1.6338 | 0.33295 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR101
EN695 | 20230113 | L4 | 9 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 323.7 | 186.8 | 0.71583 | 322.9842 | 186.0842 | 1.5335 | 0.21376 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR106
EN695 | 20230113 | L4 | 9 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 367.8 | 213.9 | 0.71583 | 367.0842 | 213.1842 | 1.7239 | 0.2778 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR107
EN695 | 20230113 | L4 | 9 | 24 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 228.8 | 144.2 | 0.72233 | 228.0777 | 143.4777 | 1.7973 | 0.75776 | 20220519 | 13-41 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L4 | 9 | 24 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 235.4 | 148.5 | 0.72233 | 234.6777 | 147.7777 | 1.8461 | 0.78547 | 20220519 | 13-41 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L4 | 9 | 24 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 88.56 | 59.01 | 0.64233 | 87.9177 | 58.3677 | 0.62777 | 0.41163 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L4 | 9 | 24 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 85.24 | 57 | 0.64233 | 84.5977 | 56.3577 | 0.59994 | 0.40367 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L4 | 9 | 24 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 183.9 | 107.9 | 0.71583 | 183.1842 | 107.1842 | 1.598 | 0.29114 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR98
EN695 | 20230113 | L4 | 9 | 24 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 192.7 | 114 | 0.71583 | 191.9842 | 113.2842 | 1.6548 | 0.34188 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR99
EN695 | 20230113 | L4 | 9 | 24 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 273.7 | 163.4 | 0.71583 | 272.9842 | 162.6842 | 1.2355 | 0.29201 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR104
EN695 | 20230113 | L4 | 9 | 24 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 237.9 | 139.6 | 0.71583 | 237.1842 | 138.8842 | 1.0971 | 0.20222 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR105
EN695 | 20230113 | L3 | 10 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 218.1 | 136.4 | 0.64233 | 217.4577 | 135.7577 | 1.7357 | 0.68189 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | filter might be off frit
EN695 | 20230113 | L3 | 10 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 201.1 | 126 | 0.64233 | 200.4577 | 125.3577 | 1.5954 | 0.6369 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L3 | 10 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 227.4 | 140.9 | 0.64233 | 226.7577 | 140.2577 | 1.8376 | 0.66006 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L3 | 10 | 4 |  |  | b | 240 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 225.2 | 138.5 | 0.64233 | 224.5577 | 137.8577 | 2.2256 | 0.74079 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L3 | 10 | 4 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 177.6 | 105.1 | 0.71583 | 176.8842 | 104.3842 | 1.5244 | 0.31538 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR112
EN695 | 20230113 | L3 | 10 | 4 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 172.1 | 102.7 | 0.71583 | 171.3842 | 101.9842 | 1.4543 | 0.33711 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR113
EN695 | 20230113 | L3 | 10 | 4 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 232.8 | 135 | 0.71583 | 232.0842 | 134.2842 | 1.0955 | 0.16536 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR118 | poor replication
EN695 | 20230113 | L3 | 10 | 4 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 186.5 | 108.3 | 0.71583 | 185.7842 | 107.5842 | 0.87596 | 0.13421 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR119 | poor replication
EN695 | 20230113 | L3 | 10 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 242.3 | 148.5 | 0.64233 | 241.6577 | 147.8577 | 1.9927 | 0.64031 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L3 | 10 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 226.9 | 140.6 | 0.72233 | 226.1777 | 139.8777 | 1.8334 | 0.65754 | 20220519 | 13-41 | 20230622 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L3 | 10 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 173.2 | 102.2 | 0.71583 | 172.4842 | 101.4842 | 1.4929 | 0.29581 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR114
EN695 | 20230113 | L3 | 10 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 164.8 | 97.61 | 0.71583 | 164.0842 | 96.8942 | 1.4128 | 0.29502 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR115
EN695 | 20230113 | L3 | 10 | 12 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 239.1 | 138.2 | 0.71583 | 238.3842 | 137.4842 | 1.1282 | 0.16039 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR120
EN695 | 20230113 | L3 | 10 | 12 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 268.5 | 157.1 | 0.71583 | 267.7842 | 156.3842 | 1.2456 | 0.22013 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR121
EN695 | 20230113 | L3 | 10 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 198 | 123 | 0.64233 | 197.3577 | 122.3577 | 1.5933 | 0.5856 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L3 | 10 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 216.5 | 133.5 | 0.64233 | 215.8577 | 132.8577 | 1.7633 | 0.60263 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L3 | 10 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 69.58 | 45.96 | 0.64233 | 68.9377 | 45.3177 | 0.50179 | 0.30522 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L3 | 10 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 65.29 | 43.76 | 0.64233 | 64.6477 | 43.1177 | 0.45739 | 0.31044 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L3 | 10 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 196.1 | 115.4 | 0.71583 | 195.3842 | 114.6842 | 1.6969 | 0.32451 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR110
EN695 | 20230113 | L3 | 10 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 201.6 | 119.1 | 0.71583 | 200.8842 | 118.3842 | 1.7347 | 0.35187 | 20220519 | 13-33 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR111
EN695 | 20230113 | L3 | 10 | 19 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 332.1 | 195 | 0.71583 | 331.3842 | 194.2842 | 1.5302 | 0.28747 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR116 | poor replication
EN695 | 20230113 | L3 | 10 | 19 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 204 | 119.8 | 0.71583 | 203.2842 | 119.0842 | 0.94317 | 0.17498 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR117 | poor replication
EN695 | 20230113 | L2 | 11 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 233.5 | 145 | 0.64233 | 232.8577 | 144.3577 | 1.8801 | 0.69058 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L2 | 11 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 209.4 | 131.4 | 0.64233 | 208.7577 | 130.7577 | 1.6571 | 0.67146 | 20220519 | 13-38 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L2 | 11 | 2 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 186.1 | 110.5 | 0.71583 | 185.3842 | 109.7842 | 1.5896 | 0.34538 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR126
EN695 | 20230113 | L2 | 11 | 2 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 186.7 | 110.6 | 0.71583 | 185.9842 | 109.8842 | 1.6001 | 0.33663 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR127
EN695 | 20230113 | L2 | 11 | 2 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 249 | 145.1 | 0.71583 | 248.2842 | 144.3842 | 1.1617 | 0.19152 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR132
EN695 | 20230113 | L2 | 11 | 2 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 245.5 | 142.6 | 0.71583 | 244.7842 | 141.8842 | 1.1505 | 0.17927 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR133
EN695 | 20230113 | L2 | 11 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 200.7 | 126.6 | 0.64233 | 200.0577 | 125.9577 | 1.5742 | 0.66883 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L2 | 11 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 203.5 | 127 | 0.64233 | 202.8577 | 126.3577 | 1.6252 | 0.62497 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L2 | 11 | 6 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 169.3 | 101.8 | 0.71583 | 168.5842 | 101.0842 | 1.4145 | 0.36112 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR124
EN695 | 20230113 | L2 | 11 | 6 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 178.5 | 106.2 | 0.71583 | 177.7842 | 105.4842 | 1.5202 | 0.33898 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR125; logsheet small leak
EN695 | 20230113 | L2 | 11 | 6 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 276 | 157.6 | 0.71583 | 275.2842 | 156.8842 | 1.3263 | 0.14682 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR130 | poor replication
EN695 | 20230113 | L2 | 11 | 6 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 197.8 | 115.4 | 0.71583 | 197.0842 | 114.6842 | 0.92301 | 0.15383 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR131 | poor replication
EN695 | 20230113 | L2 | 11 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 217.3 | 134.4 | 0.64233 | 216.6577 | 133.7577 | 1.7612 | 0.62078 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L2 | 11 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 209 | 128.7 | 0.64233 | 208.3577 | 128.0577 | 1.7059 | 0.57451 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L2 | 11 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 82.18 | 54.42 | 0.64233 | 81.5377 | 53.7777 | 0.58974 | 0.36792 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L2 | 11 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 79.61 | 53.78 | 0.64233 | 78.9677 | 53.1377 | 0.54874 | 0.39753 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230113 | L2 | 11 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 151.3 | 90.94 | 0.71583 | 150.5842 | 90.2242 | 1.2692 | 0.32107 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR122
EN695 | 20230113 | L2 | 11 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 162.1 | 95.74 | 0.71583 | 161.3842 | 95.0242 | 1.3953 | 0.27951 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR123
EN695 | 20230113 | L2 | 11 | 12 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 222 | 129.3 | 0.71583 | 221.2842 | 128.5842 | 1.0346 | 0.16836 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR128
EN695 | 20230113 | L2 | 11 | 12 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 194.5 | 113.1 | 0.71583 | 193.7842 | 112.3842 | 0.91181 | 0.14344 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR129
EN695 | 20230114 | L1 | 15 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 369.3 | 256.3 | 0.64233 | 368.6577 | 255.6577 | 2.4006 | 2.1521 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230114 | L1 | 15 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 396.4 | 268.6 | 0.64233 | 395.7577 | 267.9577 | 2.715 | 2.0567 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230114 | L1 | 15 | 4 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 331.9 | 221.4 | 0.71583 | 331.1842 | 220.6842 | 2.3156 | 1.5609 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR136
EN695 | 20230114 | L1 | 15 | 4 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 298.2 | 199.8 | 0.71583 | 297.4842 | 199.0842 | 2.069 | 1.4399 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR137
EN695 | 20230114 | L1 | 15 | 4 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 160.9 | 101.6 | 0.71583 | 160.1842 | 100.8842 | 1.9928 | 0.84903 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR140
EN695 | 20230114 | L1 | 15 | 4 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 153.3 | 95.63 | 0.71583 | 152.5842 | 94.9142 | 1.938 | 0.73564 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR141
EN695 | 20230114 | L1 | 15 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 399 | 272.2 | 0.64233 | 398.3577 | 271.5577 | 2.6938 | 2.1421 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230114 | L1 | 15 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 406.9 | 278.2 | 0.64233 | 406.2577 | 277.5577 | 2.7341 | 2.2086 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN695 | 20230114 | L1 | 15 | 12 |  |  | a | 190 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 202.5 | 137.3 | 0.71583 | 201.7842 | 136.5842 | 2.1141 | 1.5983 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR134; logsheet took too long to filter, so didn't filter whole volume | poor replication
EN695 | 20230114 | L1 | 15 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 360.5 | 241 | 0.71583 | 359.7842 | 240.2842 | 2.5127 | 1.7224 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR135 | poor replication
EN695 | 20230114 | L1 | 15 | 12 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 175.4 | 110.3 | 0.71583 | 174.6842 | 109.5842 | 2.1877 | 0.89919 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR138
EN695 | 20230114 | L1 | 15 | 12 |  |  | b | 552 | 20 | 5 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 152.9 | 96.86 | 0.71583 | 152.1842 | 96.1442 | 1.8764 | 0.82207 | 20220519 | 13-34 | 20230413 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR139
EN695 | 20230114 | MVCO | 16 | 2 |  |  | a | 290 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 204.5 | 132.8 | 0.64233 | 203.8577 | 132.1577 | 9.1393 | 4.9814 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Chl and Phaeo numbers seem really high
EN695 | 20230114 | MVCO | 16 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 210.4 | 136.6 | 0.64233 | 209.7577 | 135.9577 | 9.407 | 5.1197 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Chl and Phaeo numbers seem really high
EN695 | 20230114 | MVCO | 16 | 5 |  |  | a | 290 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 148.2 | 95.56 | 0.64233 | 147.5577 | 94.9177 | 6.7098 | 3.4319 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Chl and Phaeo numbers seem really high
EN695 | 20230114 | MVCO | 16 | 5 |  |  | b | 290 | 0 | 10 | 1 | 0 | 0.33333 | Cal_051922a | 1.8382 | 0.56187 | 152.3 | 97.62 | 0.64233 | 151.6577 | 96.9777 | 6.9698 | 3.392 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Chl and Phaeo numbers seem really high
EN695 | 20230114 | MVCO | 16 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 94.45 | 70.53 | 0.64233 | 93.8077 | 69.8877 | 0.50816 | 0.73639 | 20220519 | 13-39 | 20230615 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl, depth = 5 m
EN695 | 20230114 | MVCO | 16 | 5 |  |  | b | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | 9 |  | only <10a done, no <10b
AR70A | 20221110 |  | 3 | 3 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 78.4 | 59.45 | 0.50833 | 77.8917 | 58.9417 | 0.11045 | 0.17752 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 03/01; Glider 564 recovery | Phaeo>Chl, depth = 63m
AR70A | 20221110 |  | 3 | 4 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 203.6 | 135.9 | 0.50833 | 203.0917 | 135.3917 | 0.39497 | 0.26715 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 03/02; Glider 564 recovery
AR70A | 20221110 |  | 4 | 7 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 191.2 | 132.6 | 0.50833 | 190.6917 | 132.0917 | 0.34156 | 0.30381 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/01; OSSM
AR70A | 20221110 |  | 4 | 8 |  |  | b | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 195.2 | 135.5 | 0.50833 | 194.6917 | 134.9917 | 0.3483 | 0.31187 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/02; OSSM
AR70A | 20221111 |  | 7 | 5 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 16.99 | 13.29 | 0.50833 | 16.4817 | 12.7817 | 0.021566 | 0.040883 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 07/01; Glider 559 recovery | poor replication
AR70A | 20221111 |  | 7 | 5 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 159.5 | 104.5 | 0.50833 | 158.9917 | 103.9917 | 0.64176 | 0.37537 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | OOI 07/02; Glider 559 recovery | poor replication
AR70A | 20221111 |  | 11 | 5 |  |  | a | 1057 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 170.9 | 109.2 | 0.50833 | 170.3917 | 108.6917 | 0.35963 | 0.17142 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 11/01; CNSM
AR70A | 20221111 |  | 11 | 6 |  |  | b | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 166.9 | 108.2 | 0.50833 | 166.3917 | 107.6917 | 0.34246 | 0.18419 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 11/02; CNSM
AR70A | 20221111 |  | 11 | 7 |  |  | a | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 193.5 | 124.3 | 0.50833 | 192.9917 | 123.7917 | 0.40372 | 0.20167 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 11/03; CNSM
AR70A | 20221111 |  | 11 | 8 |  |  | b | 1056 | 0 | 5 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 196.1 | 124.9 | 0.50833 | 195.5917 | 124.3917 | 0.41539 | 0.19294 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 11/04; CNSM
AR70A | 20221113 |  | 28 | 3 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 169.8 | 111.6 | 0.50833 | 169.2917 | 111.0917 | 0.67845 | 0.40709 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 28/01; ISSM
AR70A | 20221113 |  | 28 | 4 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 161.7 | 107.9 | 0.50833 | 161.1917 | 107.3917 | 0.62775 | 0.42263 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 28/02; ISSM
AR70A | 20221113 |  | 28 | 5 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 159.4 | 103 | 0.50833 | 158.8917 | 102.4917 | 0.65809 | 0.34436 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 28/03; ISSM
AR70A | 20221113 |  | 28 | 6 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 148.3 | 97.62 | 0.50833 | 147.7917 | 97.1117 | 0.59135 | 0.35848 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 28/04; ISSM
AR70A | 20221113 |  | 28 | 7 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 278.8 | 180.2 | 0.50833 | 278.2917 | 179.6917 | 1.1494 | 0.60647 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 28/05; ISSM
AR70A | 20221113 |  | 28 | 8 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 284.2 | 184.4 | 0.50833 | 283.6917 | 183.8917 | 1.1645 | 0.63412 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 28/06; ISSM
AR70B | 20221118 |  | 4 | 7 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 160.1 | 105.6 | 0.50833 | 159.5917 | 105.0917 | 0.63532 | 0.39159 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/01; PMUO
AR70B | 20221118 |  | 4 | 8 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 165.1 | 110.1 | 0.50833 | 164.5917 | 109.5917 | 0.64176 | 0.43014 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 04/02; PMUO
AR70B | 20221118 |  | 7 | 6 |  |  | a | 1057 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 375.2 | 241.9 | 0.50833 | 374.6917 | 241.3917 | 1.5539 | 0.80487 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/01; PMUI
AR70B | 20221118 |  | 7 | 7 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 355.2 | 231.7 | 0.50833 | 354.6917 | 231.1917 | 1.441 | 0.82021 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/02; PMUI
AR70B | 20221118 |  | 7 | 8 |  |  | a | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 390.1 | 250.9 | 0.50833 | 389.5917 | 250.3917 | 1.6242 | 0.82481 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/03; PMUI
AR70B | 20221118 |  | 7 | 9 |  |  | b | 1056 | 0 | 10 | 1 | 0 | 1 | Cal_051922a | 1.8382 | 0.56187 | 377.4 | 245.6 | 0.50833 | 376.8917 | 245.0917 | 1.5379 | 0.85932 | 20220519 | 13-47 | 20230727 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | OOI 07/04; PMUI
EN706 | 20230807 | L1 | 1 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 270.2 | 177.5 | 0.624 | 269.576 | 176.876 | 2.0532 | 1.1338 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230807 | L1 | 1 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 245.9 | 162.1 | 0.624 | 245.276 | 161.476 | 1.8561 | 1.0535 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | I think the filter got switched with <10a
EN706 | 20230807 | L1 | 1 | 2 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 282.2 | 181.9 | 0.624 | 281.576 | 181.276 | 2.2215 | 1.0448 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | I think the filter got switched with b
EN706 | 20230807 | L1 | 1 | 2 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 233.9 | 154.2 | 0.624 | 233.276 | 153.576 | 1.7653 | 1.0019 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230807 | L1 | 1 | 2 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 66.7 | 44.04 | 0.69667 | 66.0033 | 43.3433 | 0.49675 | 0.27623 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR5; slightly <, leaky
EN706 | 20230807 | L1 | 1 | 2 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 71.85 | 47.72 | 0.69667 | 71.1533 | 47.0233 | 0.52898 | 0.30963 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR6
EN706 | 20230807 | L1 | 1 | 2 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 70.55 | 42.9 | 0.69667 | 69.8533 | 42.2033 | 0.32232 | 0.077905 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR11
EN706 | 20230807 | L1 | 1 | 2 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 70.45 | 44.13 | 0.69667 | 69.7533 | 43.4333 | 0.30682 | 0.10507 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR12
EN706 | 20230807 | L1 | 1 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 278.6 | 182.1 | 0.624 | 277.976 | 181.476 | 2.1374 | 1.1325 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230807 | L1 | 1 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 284.1 | 183.9 | 0.624 | 283.476 | 183.276 | 2.2193 | 1.083 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230807 | L1 | 1 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 290.3 | 189.3 | 0.624 | 289.676 | 188.676 | 2.237 | 1.1626 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | The <10 chl shouldn't be a higher value than regular chl sample
EN706 | 20230807 | L1 | 1 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 230.8 | 149.1 | 0.624 | 230.176 | 148.476 | 1.8096 | 0.86574 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230807 | L1 | 1 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 67.2 | 44.79 | 0.69667 | 66.5033 | 44.0933 | 0.49127 | 0.29508 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR3
EN706 | 20230807 | L1 | 1 | 7 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 62.33 | 40.58 | 0.69667 | 61.6333 | 39.8833 | 0.47518 | 0.23367 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR4; slightly <, leaky
EN706 | 20230807 | L1 | 1 | 7 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 73.23 | 44.35 | 0.69667 | 72.5333 | 43.6533 | 0.33728 | 0.077458 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR9
EN706 | 20230807 | L1 | 1 | 7 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 74.35 | 44.8 | 0.69667 | 73.6533 | 44.1033 | 0.3451 | 0.073909 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR10
EN706 | 20230807 | L1 | 1 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 130.7 | 79.71 | 0.624 | 130.076 | 79.086 | 1.1294 | 0.29563 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230807 | L1 | 1 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 146.4 | 89.06 | 0.624 | 145.776 | 88.436 | 1.27 | 0.32346 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230807 | L1 | 1 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 129.7 | 78.9 | 0.624 | 129.076 | 78.276 | 1.1252 | 0.28525 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230807 | L1 | 1 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 117.1 | 71.93 | 0.624 | 116.476 | 71.306 | 1.0005 | 0.28436 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230807 | L1 | 1 | 13 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 37.3 | 21.8 | 0.69667 | 36.6033 | 21.1033 | 0.33979 | 0.036563 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR1; slightly <, leaky
EN706 | 20230807 | L1 | 1 | 13 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 31 | 18.23 | 0.69667 | 30.3033 | 17.5333 | 0.27995 | 0.032743 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR2; slightly <, learky
EN706 | 20230807 | L1 | 1 | 13 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 46.87 | 25.89 | 0.69667 | 46.1733 | 25.1933 | 0.24501 | -0.0056628 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR7
EN706 | 20230807 | L1 | 1 | 13 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 58.49 | 33.75 | 0.69667 | 57.7933 | 33.0533 | 0.28788 | 0.02501 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR8
EN706 | 20230808 | L2 | 2 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 76.82 | 53.25 | 0.804 | 76.016 | 52.446 | 0.52205 | 0.42294 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L2 | 2 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 74.16 | 53.23 | 0.624 | 73.536 | 52.606 | 0.46358 | 0.4843 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo > Chl; depth = 41m
EN706 | 20230808 | L2 | 2 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 111.5 | 77.17 | 0.624 | 110.876 | 76.546 | 0.76037 | 0.61887 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L2 | 2 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 113.2 | 76.49 | 0.804 | 112.396 | 75.686 | 0.81309 | 0.55066 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L2 | 2 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 71.53 | 50.89 | 0.624 | 70.906 | 50.266 | 0.45715 | 0.44856 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L2 | 2 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 66.33 | 47.93 | 0.624 | 65.706 | 47.306 | 0.40754 | 0.44484 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo Chl; depth = 30m
EN706 | 20230808 | L2 | 2 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 64.94 | 42.06 | 0.69667 | 64.2433 | 41.3633 | 0.50158 | 0.23609 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR17; leaked
EN706 | 20230808 | L2 | 2 | 6 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 68.52 | 44.53 | 0.69667 | 67.8233 | 43.8333 | 0.52591 | 0.25581 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR18
EN706 | 20230808 | L2 | 2 | 6 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 75.73 | 44.15 | 0.69667 | 75.0333 | 43.4533 | 0.36814 | 0.043946 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR23
EN706 | 20230808 | L2 | 2 | 6 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 95.1 | 58.57 | 0.69667 | 94.4033 | 57.8733 | 0.42584 | 0.12299 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR24
EN706 | 20230808 | L2 | 2 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 320.5 | 213 | 0.624 | 319.876 | 212.376 | 2.381 | 1.4457 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L2 | 2 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 350.7 | 230.3 | 0.624 | 350.076 | 229.676 | 2.6667 | 1.4717 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L2 | 2 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 275.3 | 179.6 | 0.624 | 274.676 | 178.976 | 2.1196 | 1.1052 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replication
EN706 | 20230808 | L2 | 2 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 235.4 | 154.3 | 0.624 | 234.776 | 153.676 | 1.7963 | 0.97272 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replication
EN706 | 20230808 | L2 | 2 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 83.38 | 54.98 | 0.69667 | 82.6833 | 54.2833 | 0.62259 | 0.3455 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR15
EN706 | 20230808 | L2 | 2 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 85.27 | 56.2 | 0.69667 | 84.5733 | 55.5033 | 0.63511 | 0.35137 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR16
EN706 | 20230808 | L2 | 2 | 10 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 119.8 | 72.62 | 0.69667 | 119.1033 | 71.9233 | 0.55099 | 0.13232 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR21
EN706 | 20230808 | L2 | 2 | 10 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 105.2 | 62.08 | 0.69667 | 104.5033 | 61.3833 | 0.50358 | 0.079602 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR22
EN706 | 20230808 | L2 | 2 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 182.8 | 118.2 | 0.624 | 182.176 | 117.576 | 1.4308 | 0.68771 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L2 | 2 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 184.9 | 119 | 0.624 | 184.276 | 118.376 | 1.4596 | 0.67334 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L2 | 2 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 169 | 106.2 | 0.624 | 168.376 | 105.576 | 1.3909 | 0.51136 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L2 | 2 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 189.6 | 117.6 | 0.624 | 188.976 | 116.976 | 1.5947 | 0.513 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L2 | 2 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 153.5 | 96.18 | 0.624 | 152.876 | 95.556 | 1.2696 | 0.45219 | 20240307 | 13-73 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L2 | 2 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 145.4 | 91.9 | 0.624 | 144.776 | 91.276 | 1.185 | 0.45968 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L2 | 2 | 20 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 42.72 | 26.1 | 0.69667 | 42.0233 | 25.4033 | 0.36435 | 0.088696 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR13
EN706 | 20230808 | L2 | 2 | 20 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 43.67 | 26.42 | 0.69667 | 42.9733 | 25.7233 | 0.37816 | 0.080592 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR14
EN706 | 20230808 | L2 | 2 | 20 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 63.35 | 36.59 | 0.69667 | 62.6533 | 35.8933 | 0.31252 | 0.028492 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR19
EN706 | 20230808 | L2 | 2 | 20 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 56.17 | 33.12 | 0.69667 | 55.4733 | 32.4233 | 0.26821 | 0.038711 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR20
EN706 | 20230808 | L4 | 5 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 23.21 | 17.45 | 0.804 | 22.406 | 16.646 | 0.12758 | 0.17236 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo > Chl; depth = 60m
EN706 | 20230808 | L4 | 5 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.05 | 18.54 | 0.624 | 24.426 | 17.916 | 0.14419 | 0.17863 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo > Chl; depth = 60m
EN706 | 20230808 | L4 | 5 | 1 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13.75 | 10.96 | 0.69667 | 13.0533 | 10.2633 | 0.060955 | 0.12146 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR37
EN706 | 20230808 | L4 | 5 | 1 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.21 | 9.245 | 0.69667 | 10.5133 | 8.5483 | 0.04293 | 0.109 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR38
EN706 | 20230808 | L4 | 5 | 1 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 10.38 | 7.241 | 0.69667 | 9.6833 | 6.5443 | 0.036592 | 0.02547 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR39
EN706 | 20230808 | L4 | 5 | 1 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.59 | 8.589 | 0.69667 | 11.8933 | 7.8923 | 0.046556 | 0.028154 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR40
EN706 | 20230808 | L4 | 5 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 72.06 | 52.79 | 0.624 | 71.436 | 52.166 | 0.42681 | 0.51314 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo > Chl; depth = 47m
EN706 | 20230808 | L4 | 5 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 75.46 | 55.3 | 0.624 | 74.836 | 54.676 | 0.44652 | 0.53865 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | There were two tissue capsules C5N16b and for one of those samples, the number did not match with C5N16a; it matched better with C5N4a, and I never came across C5N4b.
EN706 | 20230808 | L4 | 5 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 225.2 | 147.1 | 0.624 | 224.576 | 146.476 | 1.7298 | 0.90944 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replication
EN706 | 20230808 | L4 | 5 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 197.4 | 130.7 | 0.804 | 196.596 | 129.896 | 1.4773 | 0.86319 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replication
EN706 | 20230808 | L4 | 5 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 153 | 102.7 | 0.804 | 152.196 | 101.896 | 1.1141 | 0.72191 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L4 | 5 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 147.9 | 99.15 | 0.804 | 147.096 | 98.346 | 1.0798 | 0.69228 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L4 | 5 | 5 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 82.45 | 52.39 | 0.69667 | 81.7533 | 51.6933 | 0.65898 | 0.26292 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR29
EN706 | 20230808 | L4 | 5 | 5 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 78.27 | 50.94 | 0.69667 | 77.5733 | 50.2433 | 0.59913 | 0.2969 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR30
EN706 | 20230808 | L4 | 5 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 126.4 | 75.3 | 0.69667 | 125.7033 | 74.6033 | 0.59569 | 0.1118 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR35
EN706 | 20230808 | L4 | 5 | 5 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 101.6 | 61.11 | 0.69667 | 100.9033 | 60.4133 | 0.472 | 0.10092 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR36
EN706 | 20230808 | L4 | 5 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 193.4 | 124.7 | 0.624 | 192.776 | 124.076 | 1.5216 | 0.71402 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L4 | 5 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 192.2 | 123.7 | 0.624 | 191.576 | 123.076 | 1.5172 | 0.70043 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L4 | 5 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 171.7 | 109.9 | 0.624 | 171.076 | 109.276 | 1.3688 | 0.60018 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L4 | 5 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 176.3 | 112.2 | 0.804 | 175.496 | 111.396 | 1.4197 | 0.58744 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L4 | 5 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 58.7 | 37.57 | 0.69667 | 58.0033 | 36.8733 | 0.46321 | 0.19438 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR27
EN706 | 20230808 | L4 | 5 | 9 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 53.21 | 34.09 | 0.69667 | 52.5133 | 33.3933 | 0.41772 | 0.17578 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR28
EN706 | 20230808 | L4 | 5 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 35.6 | 23.32 | 0.69667 | 34.9033 | 22.6233 | 0.14341 | 0.071524 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR33
EN706 | 20230808 | L4 | 5 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 37.77 | 23.75 | 0.69667 | 37.0733 | 23.0533 | 0.16373 | 0.055288 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR34
EN706 | 20230808 | L4 | 5 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 113.5 | 73.12 | 0.624 | 112.876 | 72.496 | 0.89437 | 0.41189 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L4 | 5 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 113.6 | 74.65 | 0.624 | 112.976 | 74.026 | 0.8627 | 0.47113 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L4 | 5 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 34.25 | 21.23 | 0.624 | 33.626 | 20.606 | 0.28838 | 0.082909 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L4 | 5 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 31.67 | 20.11 | 0.624 | 31.046 | 19.486 | 0.25604 | 0.095066 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L4 | 5 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.37 | 16.22 | 0.624 | 24.746 | 15.596 | 0.20266 | 0.078353 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L4 | 5 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 23.67 | 15.11 | 0.624 | 23.046 | 14.486 | 0.18959 | 0.07142 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L4 | 5 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.612 | 4.85 | 0.69667 | 6.9153 | 4.1533 | 0.060549 | 0.013521 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR25
EN706 | 20230808 | L4 | 5 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.951 | 5.055 | 0.69667 | 7.2543 | 4.3583 | 0.063486 | 0.01424 | 20240307 | 13-66 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR26
EN706 | 20230808 | L4 | 5 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.338 | 4.214 | 0.69667 | 5.6413 | 3.5173 | 0.024805 | 0.0086117 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR31
EN706 | 20230808 | L4 | 5 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.051 | 4.836 | 0.69667 | 6.3543 | 4.1393 | 0.025774 | 0.01341 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR32
EN706 | 20230808 | L5 | 6 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 58.81 | 39.89 | 0.624 | 58.186 | 39.266 | 0.41906 | 0.28845 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L5 | 6 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 60.59 | 41.52 | 0.624 | 59.966 | 40.896 | 0.42238 | 0.3145 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L5 | 6 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 30.72 | 22.4 | 0.624 | 30.096 | 21.776 | 0.18428 | 0.20809 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo > Chl; depth = 44m
EN706 | 20230808 | L5 | 6 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 28.62 | 21.19 | 0.624 | 27.996 | 20.566 | 0.16457 | 0.206 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo > Chl; depth = 44m
EN706 | 20230808 | L5 | 6 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 26.05 | 17.28 | 0.69667 | 25.3533 | 16.5833 | 0.19226 | 0.10349 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR45; need to compare to WSW chl once those are run
EN706 | 20230808 | L5 | 6 | 6 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 27.26 | 17.63 | 0.69667 | 26.5633 | 16.9333 | 0.21111 | 0.090878 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR46; need to compare to WSW chl once those are run
EN706 | 20230808 | L5 | 6 | 6 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.258 | 5.365 | 0.69667 | 7.5613 | 4.6683 | 0.033725 | 0.010547 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR51; need to compare to WSW chl once those are run
EN706 | 20230808 | L5 | 6 | 6 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.322 | 5.273 | 0.69667 | 7.6253 | 4.5763 | 0.035543 | 0.0078559 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR52; need to compare to WSW chl once those are run
EN706 | 20230808 | L5 | 6 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 153.1 | 102.4 | 0.624 | 152.476 | 101.776 | 1.1229 | 0.71089 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L5 | 6 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 148.1 | 97.73 | 0.624 | 147.476 | 97.106 | 1.1156 | 0.63406 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L5 | 6 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 95.67 | 61.54 | 0.624 | 95.046 | 60.916 | 0.75594 | 0.34167 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L5 | 6 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 94.32 | 61.26 | 0.624 | 93.696 | 60.636 | 0.73224 | 0.36032 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Two tissue capsules were marked as C6N10a, so one of them is C6N10b
EN706 | 20230808 | L5 | 6 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 38.7 | 24.89 | 0.5755 | 38.1245 | 24.3145 | 0.30274 | 0.13088 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR43; need to compare to WSW chl once those are run
EN706 | 20230808 | L5 | 6 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 39.95 | 25.81 | 0.69667 | 39.2533 | 25.1133 | 0.30892 | 0.13742 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR44; need to compare to WSW chl once those are run
EN706 | 20230808 | L5 | 6 | 10 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 19.71 | 12.06 | 0.5755 | 19.1345 | 11.4845 | 0.089341 | 0.019769 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR49; need to compare to WSW chl once those are run
EN706 | 20230808 | L5 | 6 | 10 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 17.11 | 10.73 | 0.69667 | 16.4133 | 10.0333 | 0.074509 | 0.020814 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR50; need to compare to WSW chl once those are run
EN706 | 20230808 | L5 | 6 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 77.73 | 48.59 | 0.804 | 76.926 | 47.786 | 0.64542 | 0.21561 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L5 | 6 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 70.48 | 44.51 | 0.804 | 69.676 | 43.706 | 0.57521 | 0.21231 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L5 | 6 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 36.4 | 22.32 | 0.804 | 35.596 | 21.516 | 0.31186 | 0.075828 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L5 | 6 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 33.91 | 21.34 | 0.624 | 33.286 | 20.716 | 0.27841 | 0.094858 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | I think the filter got switched with <10a
EN706 | 20230808 | L5 | 6 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 35.88 | 22.74 | 0.624 | 35.256 | 22.116 | 0.29104 | 0.10746 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | I think the filter got switched with b
EN706 | 20230808 | L5 | 6 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 32.44 | 20.25 | 0.624 | 31.816 | 19.626 | 0.26999 | 0.083634 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230808 | L5 | 6 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.859 | 5.084 | 0.5755 | 7.2835 | 4.5085 | 0.061463 | 0.019773 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR41
EN706 | 20230808 | L5 | 6 | 16 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.731 | 4.362 | 0.5755 | 6.1555 | 3.7865 | 0.052471 | 0.015756 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR42
EN706 | 20230808 | L5 | 6 | 16 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.659 | 3.874 | 0.5755 | 5.0835 | 3.2985 | 0.020771 | 0.010454 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR47
EN706 | 20230808 | L5 | 6 | 16 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.854 | 4.32 | 0.5755 | 6.2785 | 3.7445 | 0.029593 | 0.0059817 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR48
EN706 | 20230809 | L11 | 9 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 54.22 | 39.34 | 0.624 | 53.596 | 38.716 | 0.32958 | 0.36803 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phae> Chl; depth = 62m
EN706 | 20230809 | L11 | 9 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 49.9 | 36.47 | 0.804 | 49.096 | 35.666 | 0.29746 | 0.34518 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phae> Chl; depth = 62m
EN706 | 20230809 | L11 | 9 | 6 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.13 | 10.28 | 0.5755 | 13.5545 | 9.7045 | 0.084113 | 0.088368 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR65
EN706 | 20230809 | L11 | 9 | 6 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.16 | 10.19 | 0.5755 | 13.5845 | 9.6145 | 0.086735 | 0.084146 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR66
EN706 | 20230809 | L11 | 9 | 6 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.43 | 4.673 | 0.5755 | 5.8545 | 4.0975 | 0.020482 | 0.018376 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR67
EN706 | 20230809 | L11 | 9 | 6 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.021 | 5.895 | 0.5755 | 7.4455 | 5.3195 | 0.024739 | 0.025617 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR68
EN706 | 20230809 | L11 | 9 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 119.6 | 84.19 | 0.804 | 118.796 | 83.386 | 0.78429 | 0.71819 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L11 | 9 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 131.1 | 90.87 | 0.804 | 130.296 | 90.066 | 0.89105 | 0.7318 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L11 | 9 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 110.7 | 77.18 | 0.624 | 110.076 | 76.556 | 0.74243 | 0.63699 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L11 | 9 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 93.78 | 65.95 | 0.804 | 92.976 | 65.146 | 0.6164 | 0.55742 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L11 | 9 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 27.92 | 19.24 | 0.5755 | 27.3445 | 18.6645 | 0.19028 | 0.14258 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR57
EN706 | 20230809 | L11 | 9 | 9 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 28.44 | 19.91 | 0.5755 | 27.8645 | 19.3345 | 0.187 | 0.15781 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR58
EN706 | 20230809 | L11 | 9 | 9 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.68 | 8.643 | 0.5755 | 12.1045 | 8.0675 | 0.047061 | 0.029447 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR63
EN706 | 20230809 | L11 | 9 | 9 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.821 | 7.197 | 0.5755 | 9.2455 | 6.6215 | 0.030589 | 0.032205 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR64
EN706 | 20230809 | L11 | 9 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 77.96 | 52.68 | 0.804 | 77.156 | 51.876 | 0.55992 | 0.3748 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L11 | 9 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 78.74 | 52.67 | 0.804 | 77.936 | 51.866 | 0.57742 | 0.35712 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L11 | 9 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 33.35 | 21.77 | 0.804 | 32.546 | 20.966 | 0.25648 | 0.12129 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L11 | 9 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 32.56 | 21.27 | 0.804 | 31.756 | 20.466 | 0.25006 | 0.1187 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L11 | 9 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 32.3 | 21.21 | 0.624 | 31.676 | 20.586 | 0.24563 | 0.1253 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L11 | 9 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 28.95 | 18.73 | 0.804 | 28.146 | 17.926 | 0.22636 | 0.096636 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L11 | 9 | 15 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.727 | 4.624 | 0.5755 | 6.1515 | 4.0485 | 0.045945 | 0.02601 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR55
EN706 | 20230809 | L11 | 9 | 15 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.656 | 4.823 | 0.5755 | 7.0805 | 4.2475 | 0.062105 | 0.013644 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR56
EN706 | 20230809 | L11 | 9 | 15 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.319 | 1.917 | 0.5755 | 1.7435 | 1.3415 | 0.0046948 | 0.0080503 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR61
EN706 | 20230809 | L11 | 9 | 15 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.455 | 1.815 | 0.5755 | 1.8795 | 1.2395 | 0.0074742 | 0.0043018 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR62
EN706 | 20230809 | L11 | 9 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.2 | 8.262 | 0.804 | 11.396 | 7.458 | 0.087222 | 0.047159 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | I think the filter got switched with <10b
EN706 | 20230809 | L11 | 9 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13.95 | 9.134 | 0.804 | 13.146 | 8.33 | 0.10667 | 0.043424 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | filter was dropped on the ground when extracting sample
EN706 | 20230809 | L11 | 9 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 10.71 | 7.106 | 0.624 | 10.086 | 6.482 | 0.079825 | 0.036971 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L11 | 9 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13.21 | 8.841 | 0.804 | 12.406 | 8.037 | 0.096768 | 0.048045 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | I think the filter got switched with a
EN706 | 20230809 | L11 | 9 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.083 | 1.905 | 0.5755 | 2.5075 | 1.3295 | 0.025824 | -0.002114 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR53; TAR chl log has Niskin 21 written
EN706 | 20230809 | L11 | 9 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.819 | 1.989 | 0.5755 | 2.2435 | 1.4135 | 0.018195 | 0.0070129 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR54; TAR chl log has Niskin 21 written
EN706 | 20230809 | L11 | 9 | 19 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.109 | 1.367 | 0.5755 | 1.5335 | 0.7915 | 0.008634 | -0.0011416 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR59; TAR chl log has Niskin 21 written
EN706 | 20230809 | L11 | 9 | 19 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.383 | 1.721 | 0.5755 | 1.8075 | 1.1455 | 0.0077312 | 0.0031518 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR60; TAR chl log has Niskin 21 written
EN706 | 20230809 | L10 | 10 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.07 | 12.21 | 0.804 | 15.266 | 11.406 | 0.085495 | 0.12002 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phae > Chl; depht = 80m
EN706 | 20230809 | L10 | 10 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 17.29 | 12.77 | 0.804 | 16.486 | 11.966 | 0.10011 | 0.1155 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phae > Chl; depht = 80m
EN706 | 20230809 | L10 | 10 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.87 | 10.31 | 0.804 | 15.066 | 9.506 | 0.12315 | 0.048135 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L10 | 10 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.64 | 10.88 | 0.624 | 16.016 | 10.256 | 0.12758 | 0.057219 | 20240307 | 13-74 | 20240424 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L10 | 10 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.87 | 9.682 | 0.804 | 14.066 | 8.878 | 0.11491 | 0.045059 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L10 | 10 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.01 | 10.01 | 0.804 | 14.206 | 9.206 | 0.11074 | 0.055133 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L10 | 10 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.53 | 11.87 | 0.69667 | 14.8333 | 11.1733 | 0.080235 | 0.11903 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR73; TAR chl log has Niskin 8 written
EN706 | 20230809 | L10 | 10 | 7 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.91 | 11.68 | 0.69667 | 15.2133 | 10.9833 | 0.09273 | 0.10315 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR74; TAR chl log has Niskin 8 written
EN706 | 20230809 | L10 | 10 | 7 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.284 | 4.95 | 0.69667 | 5.5873 | 4.2533 | 0.015551 | 0.024785 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR79; TAR chl log has Niskin 8 written
EN706 | 20230809 | L10 | 10 | 7 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.656 | 4.483 | 0.69667 | 4.9593 | 3.7863 | 0.013674 | 0.022233 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR80; TAR chl log has Niskin 8 written
EN706 | 20230809 | L10 | 10 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 178.3 | 124.5 | 0.804 | 177.496 | 123.696 | 1.1916 | 1.0372 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L10 | 10 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 171.1 | 118.9 | 0.804 | 170.296 | 118.096 | 1.1562 | 0.97173 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L10 | 10 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 159.4 | 109.9 | 0.804 | 158.596 | 109.096 | 1.0964 | 0.86937 | 20240307 | 13-71 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L10 | 10 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 144.2 | 100 | 0.804 | 143.396 | 99.196 | 0.97898 | 0.80837 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L10 | 10 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 53.47 | 34.76 | 0.804 | 52.666 | 33.956 | 0.41441 | 0.19743 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L10 | 10 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 54.33 | 35.02 | 0.804 | 53.526 | 34.216 | 0.42769 | 0.18882 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L10 | 10 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 19.82 | 13.09 | 0.69667 | 19.1233 | 12.3933 | 0.14754 | 0.073486 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR71; TAR chl log has Niskin 16 written
EN706 | 20230809 | L10 | 10 | 15 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.21 | 9.298 | 0.69667 | 13.5133 | 8.6013 | 0.10732 | 0.045559 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR72; TAR chl log has Niskin 16 written
EN706 | 20230809 | L10 | 10 | 15 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.67 | 2.68 | 0.69667 | 2.9733 | 1.9833 | 0.011562 | 0.0072812 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR77; TAR chl log has Niskin 16 written
EN706 | 20230809 | L10 | 10 | 15 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.256 | 3.541 | 0.69667 | 4.5593 | 2.8443 | 0.020029 | 0.0069943 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR78; TAR chl log has Niskin 16 written
EN706 | 20230809 | L10 | 10 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.06 | 10.59 | 0.804 | 15.256 | 9.786 | 0.12115 | 0.055174 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L10 | 10 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 17.17 | 11.29 | 0.804 | 16.366 | 10.486 | 0.13024 | 0.058706 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L10 | 10 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.74 | 9.999 | 0.804 | 13.936 | 9.195 | 0.10501 | 0.060671 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | <10 c value | n19<10a not a complete sample so filtered a <10c should probably discard <10a; once analyzed and make <10c into <10a
EN706 | 20230809 | L10 | 10 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13.99 | 9.392 | 0.804 | 13.186 | 8.588 | 0.10184 | 0.052902 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230809 | L10 | 10 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.796 | 2.648 | 0.69667 | 3.0993 | 1.9513 | 0.025167 | 0.0096333 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR69; TAR chl log has Niskin 21 written | CHL20s and 5s from N21
EN706 | 20230809 | L10 | 10 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.556 | 2.798 | 0.69667 | 2.8593 | 2.1013 | 0.016617 | 0.020858 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR70; TAR chl log has Niskin 21 written
EN706 | 20230809 | L10 | 10 | 19 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.872 | 1.916 | 0.69667 | 2.1753 | 1.2193 | 0.011124 | 0.00041825 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR75; TAR chl log has Niskin 21 written
EN706 | 20230809 | L10 | 10 | 19 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.197 | 5.096 | 0.69667 | 6.5003 | 4.3993 | 0.024537 | 0.01726 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR76; TAR chl log has Niskin 21 written
EN706 | 20230810 | L9 | 13 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 146.4 | 102.2 | 0.804 | 145.596 | 101.396 | 0.97898 | 0.84801 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L9 | 13 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 166.3 | 113.8 | 0.804 | 165.496 | 112.996 | 1.1628 | 0.87319 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L9 | 13 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 133.7 | 92.69 | 0.804 | 132.896 | 91.886 | 0.90833 | 0.74731 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L9 | 13 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 136.1 | 93.4 | 0.804 | 135.296 | 92.596 | 0.94576 | 0.72268 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L9 | 13 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 53.16 | 35.86 | 0.5755 | 52.5845 | 35.2845 | 0.37925 | 0.25001 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR85
EN706 | 20230810 | L9 | 13 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 55.23 | 35.91 | 0.5755 | 54.6545 | 35.3345 | 0.42353 | 0.20662 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR86
EN706 | 20230810 | L9 | 13 | 12 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.16 | 8.383 | 0.5755 | 11.5845 | 7.8075 | 0.04403 | 0.030012 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR91
EN706 | 20230810 | L9 | 13 | 12 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.2 | 7.889 | 0.5755 | 10.6245 | 7.3135 | 0.038597 | 0.030759 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR92
EN706 | 20230810 | L9 | 13 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 42.83 | 28.4 | 0.73617 | 42.0938 | 27.6638 | 0.31961 | 0.17885 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L9 | 13 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 38.03 | 25.51 | 0.804 | 37.226 | 24.706 | 0.2773 | 0.16786 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L9 | 13 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.35 | 7.344 | 0.5755 | 10.7745 | 6.7685 | 0.08782 | 0.032889 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR83
EN706 | 20230810 | L9 | 13 | 18 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.91 | 8.231 | 0.5755 | 12.3345 | 7.6555 | 0.10222 | 0.033839 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR84
EN706 | 20230810 | L9 | 13 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.113 | 2.341 | 0.5755 | 2.5375 | 1.7655 | 0.0090158 | 0.0077575 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR89
EN706 | 20230810 | L9 | 13 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.489 | 2.552 | 0.5755 | 2.9135 | 1.9765 | 0.010943 | 0.0078352 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR90
EN706 | 20230810 | L9 | 13 | 23 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.3 | 10.92 | 0.804 | 15.496 | 10.116 | 0.11916 | 0.063113 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L9 | 13 | 23 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.54 | 10.86 | 0.804 | 15.736 | 10.056 | 0.12581 | 0.055387 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L9 | 13 | 23 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.93 | 8.093 | 0.73617 | 11.1938 | 7.3568 | 0.084985 | 0.047573 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L9 | 13 | 23 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.29 | 10.31 | 0.804 | 14.486 | 9.506 | 0.1103 | 0.060982 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L9 | 13 | 23 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.612 | 2.445 | 0.5755 | 3.0365 | 1.8695 | 0.025583 | 0.0077574 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR81
EN706 | 20230810 | L9 | 13 | 23 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.048 | 2.438 | 0.5755 | 2.4725 | 1.8625 | 0.013372 | 0.019843 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR82
EN706 | 20230810 | L9 | 13 | 23 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.641 | 2.035 | 0.5755 | 2.0655 | 1.4595 | 0.0070772 | 0.006789 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR87
EN706 | 20230810 | L9 | 13 | 23 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.74 | 2.079 | 0.5755 | 2.1645 | 1.5035 | 0.0076915 | 0.0065409 | 20240307 | 13-54 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR88
EN706 | 20230810 | L7 | 14 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 29.95 | 23.28 | 0.804 | 29.146 | 22.476 | 0.14773 | 0.25725 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phae > Chl; depth = 58m
EN706 | 20230810 | L7 | 14 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 29.04 | 22.26 | 0.73617 | 28.3038 | 21.5238 | 0.15017 | 0.23766 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phae > Chl; depth = 58m
EN706 | 20230810 | L7 | 14 | 5 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.07 | 8.439 | 0.5755 | 10.4945 | 7.8635 | 0.057677 | 0.08256 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR97 | Phae > Chl; depth = 58m
EN706 | 20230810 | L7 | 14 | 5 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 10.66 | 7.949 | 0.5755 | 10.0845 | 7.3735 | 0.059431 | 0.072068 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR98 | Phae > Chl; depth = 58m
EN706 | 20230810 | L7 | 14 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.948 | 3.017 | 0.5755 | 3.3725 | 2.4415 | 0.010853 | 0.012301 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR103 | Phae > Chl; depth = 58m
EN706 | 20230810 | L7 | 14 | 5 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.885 | 3.778 | 0.5755 | 4.3095 | 3.2025 | 0.012905 | 0.017466 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR104 | Phae > Chl; depth = 58m
EN706 | 20230810 | L7 | 14 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 101.8 | 68.91 | 0.73617 | 101.0638 | 68.1738 | 0.72848 | 0.49991 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L7 | 14 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 98.19 | 68.38 | 0.804 | 97.386 | 67.576 | 0.66026 | 0.55735 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L7 | 14 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 77.9 | 54.97 | 0.73617 | 77.1638 | 54.2338 | 0.50787 | 0.46933 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L7 | 14 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 79.96 | 56.38 | 0.804 | 79.156 | 55.576 | 0.52227 | 0.47912 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L7 | 14 | 10 |  |  | a | 290 | 10 |  |  |  |  | Cal_030724 | 1.8135 | 0.57627 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | 2 |  | tissue capsule missing
EN706 | 20230810 | L7 | 14 | 10 |  |  | b | 290 | 10 |  |  |  |  | Cal_030724 | 1.8135 | 0.57627 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | 2 |  | tissue capsule missing
EN706 | 20230810 | L7 | 14 | 10 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 21.18 | 14.38 | 0.5755 | 20.6045 | 13.8045 | 0.14856 | 0.096788 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR95
EN706 | 20230810 | L7 | 14 | 10 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 19.38 | 13.43 | 0.5755 | 18.8045 | 12.8545 | 0.13044 | 0.09881 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR96
EN706 | 20230810 | L7 | 14 | 10 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.007 | 4.132 | 0.5755 | 5.4315 | 3.5565 | 0.021897 | 0.011892 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR101
EN706 | 20230810 | L7 | 14 | 10 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.692 | 3.022 | 0.5755 | 4.1165 | 2.4465 | 0.019503 | 0.0037402 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR102
EN706 | 20230810 | L7 | 14 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 68.61 | 45.88 | 0.73617 | 67.8738 | 45.1438 | 0.50344 | 0.30998 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L7 | 14 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 69.51 | 47.15 | 0.804 | 68.706 | 46.346 | 0.49525 | 0.33983 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L7 | 14 | 15 |  |  | a | 285 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 46.59 | 30.23 | 0.73617 | 45.8538 | 29.4938 | 0.36871 | 0.17204 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L7 | 14 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 48.79 | 31.69 | 0.804 | 47.986 | 30.886 | 0.37875 | 0.17777 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L7 | 14 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.39 | 16.28 | 0.73617 | 24.6538 | 15.5438 | 0.20178 | 0.078299 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L7 | 14 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 24.7 | 15.49 | 0.73617 | 23.9638 | 14.7538 | 0.20399 | 0.061849 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L7 | 14 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 18.11 | 11.95 | 0.804 | 17.306 | 11.146 | 0.13644 | 0.064396 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L7 | 14 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 19.49 | 12.44 | 0.804 | 18.686 | 11.636 | 0.15615 | 0.053513 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L7 | 14 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.25 | 8.166 | 0.5755 | 11.6745 | 7.5905 | 0.08953 | 0.045838 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR93 | poor replication
EN706 | 20230810 | L7 | 14 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.51 | 5.192 | 0.5755 | 6.9345 | 4.6165 | 0.050815 | 0.031515 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR94 | poor replication
EN706 | 20230810 | L7 | 14 | 18 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.994 | 3.189 | 0.5755 | 4.4185 | 2.6135 | 0.021003 | 0.0037366 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR99
EN706 | 20230810 | L7 | 14 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.805 | 3.581 | 0.5755 | 4.2295 | 3.0055 | 0.014294 | 0.01426 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR100
EN706 | 20230810 | L8 | 15 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 59.21 | 43.3 | 0.73617 | 58.4738 | 42.5638 | 0.35239 | 0.41454 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L8 | 15 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 62.77 | 45.65 | 0.73617 | 62.0338 | 44.9138 | 0.37919 | 0.43009 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L8 | 15 | 5 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.78 | 10.78 | 0.5755 | 14.2045 | 10.2045 | 0.08739 | 0.093977 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR117 | Phae > Chl; depth = 57m
EN706 | 20230810 | L8 | 15 | 5 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.71 | 11.53 | 0.5755 | 15.1345 | 10.9545 | 0.091323 | 0.10337 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR118 | Phae > Chl; depth = 57m
EN706 | 20230810 | L8 | 15 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.856 | 6.147 | 0.5755 | 8.2805 | 5.5715 | 0.03158 | 0.021257 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR119
EN706 | 20230810 | L8 | 15 | 5 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.474 | 5.814 | 0.5755 | 7.8985 | 5.2385 | 0.030952 | 0.018636 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR120
EN706 | 20230810 | L8 | 15 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 158.8 | 115.2 | 0.73617 | 158.0638 | 114.4638 | 0.96569 | 1.0968 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phae > Chl; depth = 40m
EN706 | 20230810 | L8 | 15 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 163.6 | 118.3 | 0.73617 | 162.8638 | 117.5638 | 1.0033 | 1.115 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phae > Chl; depth = 40m
EN706 | 20230810 | L8 | 15 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 136.1 | 98.23 | 0.73617 | 135.3638 | 97.4938 | 0.83878 | 0.91791 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 9 |  | volume filtered not on logsheet
EN706 | 20230810 | L8 | 15 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 105.5 | 75.46 | 0.73617 | 104.7638 | 74.7238 | 0.66535 | 0.68105 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 9 |  | volume filtered not on logsheet
EN706 | 20230810 | L8 | 15 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 36.26 | 25.5 | 0.5755 | 35.6845 | 24.9245 | 0.23588 | 0.20862 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR109; TAR chl log has Niskin 9 written
EN706 | 20230810 | L8 | 15 | 8 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 42.65 | 29.64 | 0.5755 | 42.0745 | 29.0645 | 0.28521 | 0.23313 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR110; TAR chl log has Niskin 9 written
EN706 | 20230810 | L8 | 15 | 8 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.56 | 9.918 | 0.5755 | 13.9845 | 9.3425 | 0.054113 | 0.034485 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR114; TAR chl log has Niskin 9 written
EN706 | 20230810 | L8 | 15 | 8 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.1 | 11 | 0.5755 | 15.5245 | 10.4245 | 0.059452 | 0.039407 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR116; TAR chl log has Niskin 9 written
EN706 | 20230810 | L8 | 15 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 57.94 | 39.34 | 0.73617 | 57.2038 | 38.6038 | 0.41197 | 0.28361 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L8 | 15 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 51.9 | 35.36 | 0.73617 | 51.1638 | 34.6238 | 0.36634 | 0.25752 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L8 | 15 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 47.07 | 32.22 | 0.73617 | 46.3338 | 31.4838 | 0.32891 | 0.23838 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 9 |  | volume filtered not on logsheet
EN706 | 20230810 | L8 | 15 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 46.41 | 31.73 | 0.73617 | 45.6738 | 30.9938 | 0.32515 | 0.23331 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 9 |  | volume filtered not on logsheet
EN706 | 20230810 | L8 | 15 | 14 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13 | 8.857 | 0.5755 | 12.4245 | 8.2815 | 0.090514 | 0.056675 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR107
EN706 | 20230810 | L8 | 15 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.65 | 8.715 | 0.5755 | 12.0745 | 8.1395 | 0.086263 | 0.058896 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR108
EN706 | 20230810 | L8 | 15 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.08 | 3.13 | 0.5755 | 3.5045 | 2.5545 | 0.011095 | 0.013175 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR113
EN706 | 20230810 | L8 | 15 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.979 | 3.035 | 0.5755 | 3.4035 | 2.4595 | 0.011025 | 0.012342 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR115
EN706 | 20230810 | L8 | 15 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 29.58 | 19.65 | 0.73617 | 28.8438 | 18.9138 | 0.21994 | 0.12086 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L8 | 15 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 34.86 | 23.47 | 0.73617 | 34.1238 | 22.7338 | 0.25228 | 0.15735 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L8 | 15 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 20.05 | 13.39 | 0.804 | 19.246 | 12.586 | 0.14751 | 0.079268 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L8 | 15 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 19.02 | 12.55 | 0.804 | 18.216 | 11.746 | 0.1433 | 0.068341 | 20240307 | 13-72 | 20240410 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L8 | 15 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.51 | 9.547 | 0.73617 | 13.7738 | 8.8108 | 0.10992 | 0.048832 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L8 | 15 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.99 | 8.74 | 0.73617 | 12.2538 | 8.0038 | 0.094133 | 0.050084 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230810 | L8 | 15 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.606 | 4.489 | 0.5755 | 6.0305 | 3.9135 | 0.046409 | 0.023384 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR105
EN706 | 20230810 | L8 | 15 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.977 | 4.546 | 0.5755 | 6.4015 | 3.9705 | 0.053293 | 0.017517 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR106
EN706 | 20230810 | L8 | 15 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.804 | 3.959 | 0.5755 | 5.2285 | 3.3835 | 0.021547 | 0.010598 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR111
EN706 | 20230810 | L8 | 15 | 19 |  |  | b | 553 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.029 | 4.567 | 0.5755 | 6.4535 | 3.9915 | 0.028596 | 0.0091195 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR112
EN706 | 20230811 | L6 | 19 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 28.73 | 22 | 0.73617 | 27.9938 | 21.2638 | 0.14906 | 0.23408 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phae > Chl; depth = 75m
EN706 | 20230811 | L6 | 19 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 32.04 | 24.29 | 0.73617 | 31.3038 | 23.5538 | 0.17165 | 0.25275 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phae > Chl; depth = 75m
EN706 | 20230811 | L6 | 19 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 17.57 | 12.62 | 0.73617 | 16.8338 | 11.8838 | 0.10964 | 0.10449 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.84 | 11.47 | 0.73617 | 15.1038 | 10.7338 | 0.096791 | 0.096616 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 85.89 | 54.92 | 0.73617 | 85.1538 | 54.1838 | 0.68595 | 0.29036 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 84.3 | 55.68 | 0.73617 | 83.5638 | 54.9438 | 0.6339 | 0.3561 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 73.28 | 47.62 | 0.73617 | 72.5438 | 46.8838 | 0.56834 | 0.27643 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 76.44 | 50.26 | 0.73617 | 75.7038 | 49.5238 | 0.57986 | 0.31248 | 20240307 | 13-69 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 5 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.49 | 11.44 | 0.5755 | 15.9145 | 10.8645 | 0.11033 | 0.082767 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR133
EN706 | 20230811 | L6 | 19 | 5 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 21.99 | 14.2 | 0.5755 | 21.4145 | 13.6245 | 0.17019 | 0.071959 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR134
EN706 | 20230811 | L6 | 19 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.655 | 3.012 | 0.5755 | 4.0795 | 2.4365 | 0.019153 | 0.0039533 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR135
EN706 | 20230811 | L6 | 19 | 5 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.705 | 3.17 | 0.5755 | 4.1295 | 2.5945 | 0.017862 | 0.0066985 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR136
EN706 | 20230811 | L6 | 19 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 111.6 | 73.03 | 0.73617 | 110.8638 | 72.2938 | 0.85428 | 0.44834 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 138.3 | 90.07 | 0.73617 | 137.5638 | 89.3338 | 1.0682 | 0.54141 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 103.7 | 67.02 | 0.73617 | 102.9638 | 66.2838 | 0.81242 | 0.38191 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 113.1 | 73.79 | 0.73617 | 112.3638 | 73.0538 | 0.87067 | 0.44564 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 23.64 | 15.48 | 0.5755 | 23.0645 | 14.9045 | 0.17888 | 0.086921 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR125
EN706 | 20230811 | L6 | 19 | 9 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 26.09 | 17.09 | 0.5755 | 25.5145 | 16.5145 | 0.1973 | 0.097219 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR126
EN706 | 20230811 | L6 | 19 | 9 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.453 | 3.976 | 0.5755 | 4.8775 | 3.4005 | 0.017218 | 0.01503 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR131
EN706 | 20230811 | L6 | 19 | 9 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.84 | 4.09 | 0.5755 | 5.2645 | 3.5145 | 0.0204 | 0.012929 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR132
EN706 | 20230811 | L6 | 19 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 51.76 | 34.19 | 0.73617 | 51.0238 | 33.4538 | 0.38916 | 0.21363 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 56.96 | 37.39 | 0.73617 | 56.2238 | 36.6538 | 0.43345 | 0.22699 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.46 | 9.427 | 0.5755 | 13.8845 | 8.8515 | 0.11033 | 0.047523 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR123
EN706 | 20230811 | L6 | 19 | 15 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.8 | 10.19 | 0.5755 | 15.2245 | 9.6145 | 0.12256 | 0.048316 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR124
EN706 | 20230811 | L6 | 19 | 15 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.04 | 3.799 | 0.5755 | 4.4645 | 3.2235 | 0.014493 | 0.016132 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR129
EN706 | 20230811 | L6 | 19 | 15 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.915 | 3.314 | 0.5755 | 4.3395 | 2.7385 | 0.018697 | 0.0073202 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR130
EN706 | 20230811 | L6 | 19 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 36.5 | 23.48 | 0.73617 | 35.7638 | 22.7438 | 0.28838 | 0.12143 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 37.1 | 23.65 | 0.73617 | 36.3638 | 22.9138 | 0.2979 | 0.11497 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.53 | 16.18 | 0.73617 | 24.7938 | 15.4438 | 0.20709 | 0.071181 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.91 | 16.55 | 0.73617 | 25.1738 | 15.8138 | 0.20731 | 0.077627 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L6 | 19 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.36 | 7.277 | 0.5755 | 10.7845 | 6.7015 | 0.089508 | 0.030006 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR121
EN706 | 20230811 | L6 | 19 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13.03 | 8.424 | 0.5755 | 12.4545 | 7.8485 | 0.10097 | 0.038996 | 20240307 | 13-55 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR122
EN706 | 20230811 | L6 | 19 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.33 | 7.28 | 0.5755 | 10.7545 | 6.7045 | 0.047298 | 0.016399 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR127
EN706 | 20230811 | L6 | 19 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.331 | 6.026 | 0.5755 | 8.7555 | 5.4505 | 0.038458 | 0.013138 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR128
EN706 | 20230811 | L3 | 20 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 146.2 | 93.85 | 0.73617 | 145.4638 | 93.1138 | 1.1595 | 0.51827 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L3 | 20 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 141.4 | 91.83 | 0.73617 | 140.6638 | 91.0938 | 1.0979 | 0.54345 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L3 | 20 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 256.2 | 160.4 | 0.73617 | 255.4638 | 159.6638 | 2.1219 | 0.75502 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L3 | 20 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 246.8 | 155.5 | 0.73617 | 246.0638 | 154.7638 | 2.0222 | 0.7664 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L3 | 20 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 228.9 | 135.1 | 0.5755 | 228.3245 | 134.5245 | 2.0563 | 0.34281 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR141
EN706 | 20230811 | L3 | 20 | 9 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 227.2 | 134.4 | 0.5755 | 226.6245 | 133.8245 | 2.0344 | 0.35224 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR142
EN706 | 20230811 | L3 | 20 | 9 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 391.6 | 228.8 | 0.69667 | 390.9033 | 228.1033 | 1.8978 | 0.26538 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR147
EN706 | 20230811 | L3 | 20 | 9 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 376.7 | 219.9 | 0.69667 | 376.0033 | 219.2033 | 1.8279 | 0.25092 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR148
EN706 | 20230811 | L3 | 20 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 170.5 | 112.4 | 0.73617 | 169.7638 | 111.6638 | 1.2868 | 0.72515 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L3 | 20 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 181.7 | 119.4 | 0.73617 | 180.9638 | 118.6638 | 1.3799 | 0.75826 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L3 | 20 | 11 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 92.38 | 56.7 | 0.5755 | 91.8045 | 56.1245 | 0.78218 | 0.21874 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR139
EN706 | 20230811 | L3 | 20 | 11 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 91.7 | 56.18 | 0.5755 | 91.1245 | 55.6045 | 0.77602 | 0.21225 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR140
EN706 | 20230811 | L3 | 20 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 153.8 | 91.07 | 0.69667 | 153.1033 | 90.3733 | 0.73259 | 0.12601 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR145; log sheet says spilled about 10mL
EN706 | 20230811 | L3 | 20 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 161.7 | 94.98 | 0.69667 | 161.0033 | 94.2833 | 0.77919 | 0.11656 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR146
EN706 | 20230811 | L3 | 20 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 143.1 | 98.9 | 0.73617 | 142.3638 | 98.1638 | 0.97898 | 0.78978 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L3 | 20 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 145.5 | 99.23 | 0.73617 | 144.7638 | 98.4938 | 1.0248 | 0.74987 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L3 | 20 | 22 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 98.97 | 61.28 | 0.73617 | 98.2338 | 60.5438 | 0.83479 | 0.25611 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L3 | 20 | 22 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 113.9 | 69.67 | 0.73617 | 113.1638 | 68.9338 | 0.97964 | 0.26243 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L3 | 20 | 22 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 50.24 | 31.77 | 0.73617 | 49.5038 | 31.0338 | 0.40909 | 0.15009 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L3 | 20 | 22 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 60.45 | 37.62 | 0.73617 | 59.7138 | 36.8838 | 0.50566 | 0.15893 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN706 | 20230811 | L3 | 20 | 22 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 63.09 | 37.04 | 0.5755 | 62.5145 | 36.4645 | 0.57107 | 0.079234 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR137
EN706 | 20230811 | L3 | 20 | 22 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 61.17 | 35.76 | 0.5755 | 60.5945 | 35.1845 | 0.55704 | 0.070437 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR138
EN706 | 20230811 | L3 | 20 | 22 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 110.5 | 65.79 | 0.69667 | 109.8033 | 65.0933 | 0.52215 | 0.096281 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR143
EN706 | 20230811 | L3 | 20 | 22 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 87.06 | 51.52 | 0.69667 | 86.3633 | 50.8233 | 0.41355 | 0.067553 | 20240307 | 13-67 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR144
EN706 | 20230812 | MVCO | 24 | 2 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 174.2 | 113.1 | 0.73617 | 173.4638 | 112.3638 | 2.7066 | 1.3426 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | cast number on tissue capsule says 23
EN706 | 20230812 | MVCO | 24 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 181.2 | 116.3 | 0.73617 | 180.4638 | 115.5638 | 2.8749 | 1.2896 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | cast number on tissue capsule says 23
EN706 | 20230812 | MVCO | 24 | 4 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 164.4 | 107.8 | 0.73617 | 163.6638 | 107.0638 | 2.5073 | 1.351 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | cast number on tissue capsule says 23
EN706 | 20230812 | MVCO | 24 | 4 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 176.4 | 114.1 | 0.73617 | 175.6638 | 113.3638 | 2.7597 | 1.3255 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | cast number on tissue capsule says 23
EN706 | 20230812 | MVCO | 24 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 182.3 | 119.1 | 0.73617 | 181.5638 | 118.3638 | 1.3998 | 0.73292 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | cast number on tissue capsule says 23
EN706 | 20230812 | MVCO | 24 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 190.9 | 124.7 | 0.73617 | 190.1638 | 123.9638 | 1.4663 | 0.76737 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | cast number on tissue capsule says 23
EN706 | 20230812 | MVCO | 24 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 117.5 | 74.63 | 0.73617 | 116.7638 | 73.8938 | 0.94952 | 0.38193 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | cast number on tissue capsule says 23
EN706 | 20230812 | MVCO | 24 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 124 | 80.26 | 0.73617 | 123.2638 | 79.5238 | 0.96879 | 0.4641 | 20240307 | 13-70 | 20240404 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | cast number on tissue capsule says 23
EN706 | 20230812 | MVCO | 24 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 113.8 | 74.29 | 0.5755 | 113.2245 | 73.7145 | 0.86614 | 0.44848 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR149
EN706 | 20230812 | MVCO | 24 | 7 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 125.1 | 79.06 | 0.5755 | 124.5245 | 78.4845 | 1.0093 | 0.39039 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR150
EN706 | 20230812 | MVCO | 24 | 7 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 159.7 | 100.8 | 0.5755 | 159.1245 | 100.2245 | 0.68537 | 0.26337 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR151
EN706 | 20230812 | MVCO | 24 | 7 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 144 | 94.79 | 0.5755 | 143.4245 | 94.2145 | 0.5747 | 0.3204 | 20240307 | 13-56 | 20240104 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR152
AR77 | 20231011 | L1 | 1 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 206.6 | 127.8 | 0.67983 | 205.9202 | 127.1202 | 1.7453 | 0.54517 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231011 | L1 | 1 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 226.6 | 137.4 | 0.48483 | 226.1152 | 136.9152 | 1.9757 | 0.49131 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231011 | L1 | 1 | 2 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 52.01 | 35.39 | 0.40133 | 51.6087 | 34.9887 | 0.36811 | 0.26233 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231011 | L1 | 1 | 2 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 51.23 | 35.43 | 0.48483 | 50.7452 | 34.9452 | 0.34995 | 0.2797 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231011 | L1 | 1 | 2 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 179.4 | 106.7 | 0.61383 | 178.7862 | 106.0862 | 1.5937 | 0.2982 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR5
AR77 | 20231011 | L1 | 1 | 2 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 174.1 | 103.7 | 0.61383 | 173.4862 | 103.0862 | 1.5433 | 0.29511 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR6
AR77 | 20231011 | L1 | 1 | 2 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 278.5 | 164.5 | 0.61383 | 277.8862 | 163.8862 | 1.3313 | 0.22567 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR11; logsheet says small drip about 1mL
AR77 | 20231011 | L1 | 1 | 2 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 295.6 | 172 | 0.61383 | 294.9862 | 171.3862 | 1.4408 | 0.18448 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR12
AR77 | 20231011 | L1 | 1 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 265.3 | 158.6 | 0.83067 | 264.4693 | 157.7693 | 2.3633 | 0.47947 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231011 | L1 | 1 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 298.3 | 177 | 0.40133 | 297.8987 | 176.5987 | 2.6867 | 0.49537 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231011 | L1 | 1 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 204.8 | 118.6 | 0.61383 | 204.1862 | 117.9862 | 1.8897 | 0.21447 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR3
AR77 | 20231011 | L1 | 1 | 7 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 173.2 | 101.8 | 0.61383 | 172.5862 | 101.1862 | 1.5599 | 0.23849 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR4
AR77 | 20231011 | L1 | 1 | 7 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 285.1 | 167.2 | 0.61383 | 284.4862 | 166.5862 | 1.3769 | 0.20578 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR9
AR77 | 20231011 | L1 | 1 | 7 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 359.8 | 208.7 | 0.61383 | 359.1862 | 208.0862 | 1.7646 | 0.21232 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR10
AR77 | 20231011 | L1 | 1 | 14 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 140.2 | 83.26 | 0.83067 | 139.3693 | 82.4293 | 2.5223 | 0.44818 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231011 | L1 | 1 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 255.8 | 150.2 | 0.69617 | 255.1038 | 149.5038 | 2.3389 | 0.3549 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231011 | L1 | 1 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 78.83 | 48.99 | 0.83067 | 77.9993 | 48.1593 | 0.66092 | 0.20683 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231011 | L1 | 1 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 80.19 | 49.23 | 0.40133 | 79.7887 | 48.8287 | 0.68573 | 0.19409 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231011 | L1 | 1 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 169.3 | 99.55 | 0.61383 | 168.6862 | 98.9362 | 1.5291 | 0.23535 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR1
AR77 | 20231011 | L1 | 1 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 184.5 | 107.5 | 0.61383 | 183.8862 | 106.8862 | 1.688 | 0.2182 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR2
AR77 | 20231011 | L1 | 1 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 276.4 | 159 | 0.61383 | 275.7862 | 158.3862 | 1.3711 | 0.13371 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR7
AR77 | 20231011 | L1 | 1 | 14 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 233.5 | 136.4 | 0.61383 | 232.8862 | 135.7862 | 1.1299 | 0.1555 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR8
AR77 | 20231012 | L2 | 2 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 94.71 | 69.28 | 0.83067 | 93.8793 | 68.4493 | 0.56325 | 0.6701 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl; depth =42
AR77 | 20231012 | L2 | 2 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 94.84 | 68.1 | 0.67983 | 94.1602 | 67.4202 | 0.59226 | 0.62254 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl; depth =42
AR77 | 20231012 | L2 | 2 | 1 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 48.23 | 37.12 | 0.48483 | 47.7452 | 36.6352 | 0.24607 | 0.41403 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl; depth =42
AR77 | 20231012 | L2 | 2 | 1 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 49.2 | 37.56 | 0.48483 | 48.7152 | 37.0752 | 0.25781 | 0.41022 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl; depth =42
AR77 | 20231012 | L2 | 2 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 158.7 | 103.9 | 0.48483 | 158.2152 | 103.4152 | 1.2138 | 0.64962 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L2 | 2 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 188.2 | 123.2 | 0.48483 | 187.7152 | 122.7152 | 1.4397 | 0.77145 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L2 | 2 | 3 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 104.5 | 64.67 | 0.61383 | 103.8862 | 64.0562 | 0.87316 | 0.26922 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR17; logsheet says manifold leaking a little here
AR77 | 20231012 | L2 | 2 | 3 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 105.6 | 65.55 | 0.61383 | 104.9862 | 64.9362 | 0.87798 | 0.28009 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR18
AR77 | 20231012 | L2 | 2 | 3 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 156.2 | 91.81 | 0.61383 | 155.5862 | 91.1962 | 0.75061 | 0.11423 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR23; logsheet says manifold leak small in same place
AR77 | 20231012 | L2 | 2 | 3 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 137.6 | 80.87 | 0.61383 | 136.9862 | 80.2562 | 0.66132 | 0.09978 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR24
AR77 | 20231012 | L2 | 2 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 255.9 | 161.9 | 0.67983 | 255.2202 | 161.2202 | 2.082 | 0.82294 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L2 | 2 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 249.9 | 160.2 | 0.48483 | 249.4152 | 159.7152 | 1.9868 | 0.89106 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L2 | 2 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 108.8 | 73.85 | 0.48483 | 108.3152 | 73.3652 | 0.7741 | 0.54782 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L2 | 2 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 123.9 | 83.39 | 0.40133 | 123.4987 | 82.9887 | 0.89725 | 0.59807 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L2 | 2 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 157.8 | 95.05 | 0.61383 | 157.1862 | 94.4362 | 1.3756 | 0.30855 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR15
AR77 | 20231012 | L2 | 2 | 8 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 156.8 | 96.46 | 0.61383 | 156.1862 | 95.8462 | 1.3183 | 0.38522 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR16
AR77 | 20231012 | L2 | 2 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 241.4 | 140.3 | 0.61383 | 240.7862 | 139.6862 | 1.1807 | 0.14641 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR21
AR77 | 20231012 | L2 | 2 | 8 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 211 | 122.6 | 0.61383 | 210.3862 | 121.9862 | 1.0324 | 0.12656 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR22
AR77 | 20231012 | L2 | 2 | 14 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 209.4 | 130.9 | 0.40133 | 208.9987 | 130.4987 | 3.4774 | 1.2254 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
AR77 | 20231012 | L2 | 2 | 14 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 177.7 | 112.6 | 0.67983 | 177.0202 | 111.9202 | 2.8838 | 1.1495 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
AR77 | 20231012 | L2 | 2 | 14 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 182.4 | 109.2 | 0.61383 | 181.7862 | 108.5862 | 1.5992 | 0.33069 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR25
AR77 | 20231012 | L2 | 2 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 186.3 | 111 | 0.61383 | 185.6862 | 110.3862 | 1.6451 | 0.3168 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR26
AR77 | 20231012 | L2 | 2 | 14 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 292.7 | 170.8 | 0.61383 | 292.0862 | 170.1862 | 1.421 | 0.19291 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR27
AR77 | 20231012 | L2 | 2 | 14 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 270.1 | 157.8 | 0.61383 | 269.4862 | 157.1862 | 1.3067 | 0.18121 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR28
AR77 | 20231012 | L2 | 2 | 18 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 120.9 | 77.35 | 0.67983 | 120.2202 | 76.6702 | 1.9292 | 0.83378 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L2 | 2 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 227.6 | 146.4 | 0.48483 | 227.1152 | 145.9152 | 1.7985 | 0.83067 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L2 | 2 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 127.6 | 84.8 | 0.40133 | 127.1987 | 84.3987 | 0.94797 | 0.57276 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L2 | 2 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 129.9 | 86.7 | 0.83067 | 129.0693 | 85.8693 | 0.95683 | 0.5904 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L2 | 2 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 152.7 | 91.87 | 0.61383 | 152.0862 | 91.2562 | 1.3335 | 0.29393 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR13
AR77 | 20231012 | L2 | 2 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 162.9 | 98.21 | 0.61383 | 162.2862 | 97.5962 | 1.4181 | 0.32238 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR14
AR77 | 20231012 | L2 | 2 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 252.5 | 146.6 | 0.61383 | 251.8862 | 145.9862 | 1.2368 | 0.15021 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR19 | poor replication
AR77 | 20231012 | L2 | 2 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 355.6 | 208.2 | 0.61383 | 354.9862 | 207.5862 | 1.7152 | 0.24988 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR20 | poor replication
AR77 | 20231012 | L5 | 5 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.52 | 21.26 | 0.83067 | 24.6893 | 20.4293 | 0.094354 | 0.27375 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 73m
AR77 | 20231012 | L5 | 5 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 28.2 | 22.42 | 0.40133 | 27.7987 | 22.0187 | 0.12802 | 0.26872 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 73m
AR77 | 20231012 | L5 | 5 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 31.56 | 24.81 | 0.40133 | 31.1587 | 24.4087 | 0.1495 | 0.2903 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 60m
AR77 | 20231012 | L5 | 5 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 28.27 | 23.24 | 0.40133 | 27.8687 | 22.8387 | 0.11141 | 0.30011 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 60m
AR77 | 20231012 | L5 | 5 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 39 | 25.41 | 0.67983 | 38.3202 | 24.7302 | 0.301 | 0.14459 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L5 | 5 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 44.7 | 30.5 | 0.48483 | 44.2152 | 30.0152 | 0.31451 | 0.22631 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L5 | 5 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 128.8 | 84.67 | 0.83067 | 127.9693 | 83.8393 | 0.97743 | 0.53322 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L5 | 5 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 137.2 | 87.56 | 0.67983 | 136.5202 | 86.8802 | 1.0995 | 0.46597 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L5 | 5 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 64.25 | 41.14 | 0.61383 | 63.6362 | 40.5262 | 0.50662 | 0.21612 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR33
AR77 | 20231012 | L5 | 5 | 8 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 61.74 | 38.58 | 0.61383 | 61.1262 | 37.9662 | 0.50772 | 0.16937 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR34
AR77 | 20231012 | L5 | 5 | 8 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 90.04 | 54.02 | 0.61383 | 89.4262 | 53.4062 | 0.4199 | 0.086574 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR39
AR77 | 20231012 | L5 | 5 | 8 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 108.5 | 64.78 | 0.61383 | 107.8862 | 64.1662 | 0.50966 | 0.098854 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR40
AR77 | 20231012 | L5 | 5 | 10 |  |  | a | 290 | 0 |  |  |  |  | Cal_030724 | 1.8135 | 0.57627 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR77 | 20231012 | L5 | 5 | 10 |  |  | b | 290 | 0 |  |  |  |  | Cal_030724 | 1.8135 | 0.57627 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR77 | 20231012 | L5 | 5 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 63.45 | 39.12 | 0.61383 | 62.8362 | 38.5062 | 0.53336 | 0.15335 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR31
AR77 | 20231012 | L5 | 5 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 56.84 | 35.56 | 0.61383 | 56.2262 | 34.9462 | 0.46492 | 0.15619 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR32
AR77 | 20231012 | L5 | 5 | 10 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 90.18 | 53.79 | 0.61383 | 89.5662 | 53.1762 | 0.42498 | 0.080226 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR37
AR77 | 20231012 | L5 | 5 | 10 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 102.8 | 61.42 | 0.61383 | 102.1862 | 60.8062 | 0.48326 | 0.09444 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR38
AR77 | 20231012 | L5 | 5 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 137.7 | 86.44 | 0.40133 | 137.2987 | 86.0387 | 1.1354 | 0.41493 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L5 | 5 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 131.8 | 83.5 | 0.40133 | 131.3987 | 83.0987 | 1.0698 | 0.42751 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L5 | 5 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 68.89 | 45.9 | 0.83067 | 68.0593 | 45.0693 | 0.5092 | 0.30287 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L5 | 5 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 77.05 | 50.09 | 0.48483 | 76.5652 | 49.6052 | 0.59713 | 0.29667 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231012 | L5 | 5 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 49.97 | 30.96 | 0.61383 | 49.3562 | 30.3462 | 0.41674 | 0.12445 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR29
AR77 | 20231012 | L5 | 5 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 57.62 | 35.5 | 0.61383 | 57.0062 | 34.8862 | 0.48492 | 0.13724 | 20240307 | 13-57 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR30
AR77 | 20231012 | L5 | 5 | 12 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 89.79 | 53.89 | 0.61383 | 89.1762 | 53.2762 | 0.41774 | 0.086583 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR35
AR77 | 20231012 | L5 | 5 | 12 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 99.44 | 59.51 | 0.61383 | 98.8262 | 58.8962 | 0.46632 | 0.093227 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR36
AR77 | 20231013 | L10 | 6 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13.78 | 10.49 | 0.48483 | 13.2952 | 10.0052 | 0.07287 | 0.10741 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth =61m
AR77 | 20231013 | L10 | 6 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.45 | 9.622 | 0.40133 | 12.0487 | 9.2207 | 0.062637 | 0.1035 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth =61m
AR77 | 20231013 | L10 | 6 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.286 | 6.947 | 0.40133 | 8.8847 | 6.5457 | 0.051806 | 0.066136 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth =61m
AR77 | 20231013 | L10 | 6 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.823 | 7.426 | 0.48483 | 9.3382 | 6.9412 | 0.053091 | 0.071978 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth =61m
AR77 | 20231013 | L10 | 6 | 5 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.938 | 6.6 | 0.61383 | 8.3242 | 5.9862 | 0.051254 | 0.055503 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR45 | Phaeo>Chl; depth = 61m
AR77 | 20231013 | L10 | 6 | 5 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.097 | 4.07 | 0.61383 | 4.4832 | 3.4562 | 0.022514 | 0.039123 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR46 | Phaeo>Chl; depth = 61m
AR77 | 20231013 | L10 | 6 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.919 | 2.195 | 0.61383 | 2.3052 | 1.5812 | 0.0084399 | 0.0065549 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR51
AR77 | 20231013 | L10 | 6 | 5 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.112 | 2.201 | 0.61383 | 2.4982 | 1.5872 | 0.01062 | 0.0044319 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR52
AR77 | 20231013 | L10 | 6 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 37.88 | 26.77 | 0.40133 | 37.4787 | 26.3687 | 0.24607 | 0.22905 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L10 | 6 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 34.54 | 24.97 | 0.83067 | 33.7093 | 24.1393 | 0.21196 | 0.22299 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl; depth =40m
AR77 | 20231013 | L10 | 6 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 71.28 | 49.19 | 0.48483 | 70.7952 | 48.7052 | 0.48927 | 0.38832 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L10 | 6 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 69.29 | 47.19 | 0.48483 | 68.8052 | 46.7052 | 0.48949 | 0.35206 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L10 | 6 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 60.61 | 41.2 | 0.40133 | 60.2087 | 40.7987 | 0.42991 | 0.30522 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L10 | 6 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 67.12 | 45.03 | 0.67983 | 66.4402 | 44.3502 | 0.48927 | 0.30985 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L10 | 6 | 9 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.15 | 11.32 | 0.61383 | 15.5362 | 10.7062 | 0.10552 | 0.08476 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR43
AR77 | 20231013 | L10 | 6 | 9 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.58 | 10.65 | 0.61383 | 14.9662 | 10.0362 | 0.10808 | 0.070908 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR44
AR77 | 20231013 | L10 | 6 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.975 | 4.307 | 0.61383 | 5.3612 | 3.6932 | 0.01948 | 0.015608 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR49
AR77 | 20231013 | L10 | 6 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.387 | 3.32 | 0.61383 | 3.7732 | 2.7062 | 0.012461 | 0.013249 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR50
AR77 | 20231013 | L10 | 6 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 55.99 | 37.96 | 0.48483 | 55.5052 | 37.4752 | 0.39934 | 0.2759 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L10 | 6 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 59.94 | 40.7 | 0.67983 | 59.2602 | 40.0202 | 0.42614 | 0.29496 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L10 | 6 | 15 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.81 | 8.203 | 0.61383 | 11.1962 | 7.5892 | 0.078804 | 0.05608 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR53
AR77 | 20231013 | L10 | 6 | 15 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.694 | 7.048 | 0.61383 | 9.0802 | 6.4342 | 0.057809 | 0.056548 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR54; logsheet says manifold leak little
AR77 | 20231013 | L10 | 6 | 15 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.454 | 2.743 | 0.61383 | 2.8402 | 2.1292 | 0.0082733 | 0.011882 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR55
AR77 | 20231013 | L10 | 6 | 15 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.721 | 2.698 | 0.61383 | 3.1072 | 2.0842 | 0.011925 | 0.0078395 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR56
AR77 | 20231013 | L10 | 6 | 18 |  |  | a | 290 | 0 |  |  |  |  | Cal_030724 | 1.8135 | 0.57627 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR77 | 20231013 | L10 | 6 | 18 |  |  | b | 290 | 0 |  |  |  |  | Cal_030724 | 1.8135 | 0.57627 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR77 | 20231013 | L10 | 6 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 34.75 | 22.81 | 0.48483 | 34.2652 | 22.3252 | 0.26446 | 0.13781 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L10 | 6 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 30.43 | 19.85 | 0.67983 | 29.7502 | 19.1702 | 0.23434 | 0.11108 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L10 | 6 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.311 | 6.574 | 0.61383 | 8.6972 | 5.9602 | 0.060001 | 0.046292 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR41
AR77 | 20231013 | L10 | 6 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.531 | 6.76 | 0.61383 | 8.9172 | 6.1462 | 0.060746 | 0.048864 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR42; logsheet says little manifold leak
AR77 | 20231013 | L10 | 6 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.039 | 2.883 | 0.61383 | 3.4252 | 2.2692 | 0.0135 | 0.0080581 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR47
AR77 | 20231013 | L10 | 6 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.686 | 2.481 | 0.61383 | 3.0722 | 1.8672 | 0.014022 | 0.0036534 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR48; logsheet says little manifold leak
AR77 | 20231013 | L11 | 9 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.88 | 7.17 | 0.48483 | 8.3952 | 6.6852 | 0.037875 | 0.082581 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth =63m
AR77 | 20231013 | L11 | 9 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.253 | 7.453 | 0.40133 | 8.8517 | 7.0517 | 0.039868 | 0.087192 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth =63m
AR77 | 20231013 | L11 | 9 | 5 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.082 | 3.779 | 0.61383 | 3.4682 | 3.1652 | 0.0066424 | 0.049805 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR61 | Phaeo>Chl; depth = 63m
AR77 | 20231013 | L11 | 9 | 5 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.079 | 3.495 | 0.61383 | 3.4652 | 2.8812 | 0.012803 | 0.03858 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR62 | Phaeo>Chl; depth = 63m
AR77 | 20231013 | L11 | 9 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 1.683 | 1.384 | 0.61383 | 1.0692 | 0.77017 | 0.0034855 | 0.0038182 | 20240307 | 13-59 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR67 | Phaeo>Chl; depth = 63m
AR77 | 20231013 | L11 | 9 | 5 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.252 | 1.633 | 0.61383 | 1.6382 | 1.0192 | 0.0072159 | 0.0024493 | 20240307 | 13-59 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR68
AR77 | 20231013 | L11 | 9 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 46.17 | 32.48 | 0.40133 | 45.7687 | 32.0787 | 0.30322 | 0.27479 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L11 | 9 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 45.94 | 32.7 | 0.48483 | 45.4552 | 32.2152 | 0.29325 | 0.28721 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L11 | 9 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 40.74 | 28.32 | 0.48483 | 40.2552 | 27.8352 | 0.27509 | 0.22646 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L11 | 9 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 39.93 | 27.9 | 0.48483 | 39.4452 | 27.4152 | 0.26645 | 0.22753 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L11 | 9 | 7 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 10.78 | 8.021 | 0.61383 | 10.1662 | 7.4072 | 0.060277 | 0.071372 | 20240307 | 13-59 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR69 | Phaeo>Chl; depth = 43m
AR77 | 20231013 | L11 | 9 | 7 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.65 | 8.491 | 0.61383 | 11.0362 | 7.8772 | 0.069016 | 0.070987 | 20240307 | 13-59 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR70 | Phaeo>Chl; depth = 43m
AR77 | 20231013 | L11 | 9 | 7 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.755 | 2.749 | 0.61383 | 3.1412 | 2.1352 | 0.011727 | 0.0085213 | 20240307 | 13-59 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR71
AR77 | 20231013 | L11 | 9 | 7 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.568 | 2.759 | 0.61383 | 2.9542 | 2.1452 | 0.0094137 | 0.010893 | 20240307 | 13-59 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR72 | Phaeo>Chl; depth = 43m
AR77 | 20231013 | L11 | 9 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 64.1 | 44.68 | 0.83067 | 63.2693 | 43.8493 | 0.43013 | 0.35996 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L11 | 9 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 61.54 | 43.16 | 0.69617 | 60.8438 | 42.4638 | 0.4071 | 0.35803 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L11 | 9 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 48.96 | 34.68 | 0.83067 | 48.1293 | 33.8493 | 0.31629 | 0.29362 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L11 | 9 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 54.7 | 38.01 | 0.48483 | 54.2152 | 37.5252 | 0.36966 | 0.30648 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L11 | 9 | 11 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.11 | 8.912 | 0.61383 | 11.4962 | 8.2982 | 0.069868 | 0.077617 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR59; logsheet says manifold little leak | Phaeo>Chl; depth = 33m
AR77 | 20231013 | L11 | 9 | 11 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.31 | 9.117 | 0.61383 | 11.6962 | 8.5032 | 0.069997 | 0.081647 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR60 | Phaeo>Chl; depth = 33m
AR77 | 20231013 | L11 | 9 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.91 | 3.164 | 0.61383 | 3.2962 | 2.5502 | 0.0087122 | 0.015516 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR65 | Phaeo>Chl; depth = 33m
AR77 | 20231013 | L11 | 9 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.458 | 3.395 | 0.61383 | 3.8442 | 2.7812 | 0.012414 | 0.014009 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR66 | Phaeo>Chl; depth = 33m
AR77 | 20231013 | L11 | 9 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 42.15 | 29.34 | 0.48483 | 41.6652 | 28.8552 | 0.28373 | 0.2362 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L11 | 9 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 44.88 | 30.73 | 0.67983 | 44.2002 | 30.0502 | 0.31341 | 0.22805 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L11 | 9 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 36.87 | 24.19 | 0.67983 | 36.1902 | 23.5102 | 0.28085 | 0.14277 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L11 | 9 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 35.38 | 23.23 | 0.40133 | 34.9787 | 22.8287 | 0.26911 | 0.14223 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L11 | 9 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 28.62 | 18.58 | 0.40133 | 28.2187 | 18.1787 | 0.22237 | 0.10518 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L11 | 9 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 29.6 | 19.4 | 0.48483 | 29.1152 | 18.9152 | 0.22592 | 0.1149 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L11 | 9 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.972 | 6.38 | 0.61383 | 8.3582 | 5.7662 | 0.056822 | 0.046011 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR57
AR77 | 20231013 | L11 | 9 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.293 | 5.777 | 0.61383 | 7.6792 | 5.1632 | 0.055156 | 0.036923 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR58
AR77 | 20231013 | L11 | 9 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.716 | 2.182 | 0.61383 | 2.1022 | 1.5682 | 0.0062363 | 0.0086622 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR63
AR77 | 20231013 | L11 | 9 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.813 | 2.116 | 0.61383 | 2.1992 | 1.5022 | 0.0081104 | 0.0061094 | 20240307 | 13-58 | 20240119 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR64
AR77 | 20231013 | L6 | 10 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 82.44 | 53.6 | 0.67983 | 81.7602 | 52.9202 | 0.63877 | 0.31476 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L6 | 10 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 90.14 | 58.64 | 0.40133 | 89.7387 | 58.2387 | 0.69769 | 0.35168 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L6 | 10 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 106.2 | 67.4 | 0.40133 | 105.7987 | 66.9987 | 0.85938 | 0.34783 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L6 | 10 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 111.2 | 69.9 | 0.67983 | 110.5202 | 69.2202 | 0.91475 | 0.33249 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L6 | 10 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 57.85 | 37.3 | 0.83067 | 57.0193 | 36.4693 | 0.45516 | 0.20196 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L6 | 10 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 46.71 | 30.91 | 0.40133 | 46.3087 | 30.5087 | 0.34995 | 0.19977 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L6 | 10 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 67.56 | 43.01 | 0.6925 | 66.8675 | 42.3175 | 0.53819 | 0.2165 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR77
AR77 | 20231013 | L6 | 10 | 6 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 62.68 | 39.41 | 0.6925 | 61.9875 | 38.7175 | 0.51013 | 0.18036 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR78
AR77 | 20231013 | L6 | 10 | 6 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 63.55 | 39.3 | 0.69667 | 62.8533 | 38.6033 | 0.28269 | 0.0834 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR83; logsheet says leaky manifold | poor replicate
AR77 | 20231013 | L6 | 10 | 6 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 101.6 | 62.19 | 0.69667 | 100.9033 | 61.4933 | 0.45941 | 0.12375 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR84 | poor replicate
AR77 | 20231013 | L6 | 10 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 199 | 124.5 | 0.40133 | 198.5987 | 124.0987 | 1.6501 | 0.58597 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L6 | 10 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 187.7 | 119.6 | 0.48483 | 187.2152 | 119.1152 | 1.5083 | 0.63793 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L6 | 10 | 9 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 133.5 | 81.21 | 0.6925 | 132.8075 | 80.5175 | 1.1424 | 0.28865 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR85
AR77 | 20231013 | L6 | 10 | 9 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 167.6 | 100.9 | 0.6925 | 166.9075 | 100.2075 | 1.4572 | 0.32378 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR86
AR77 | 20231013 | L6 | 10 | 9 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 289.9 | 171.2 | 0.6925 | 289.2075 | 170.5075 | 1.3837 | 0.23326 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR87
AR77 | 20231013 | L6 | 10 | 9 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 270.2 | 160.2 | 0.6925 | 269.5075 | 159.5075 | 1.28 | 0.22995 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR88
AR77 | 20231013 | L6 | 10 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 202 | 126.7 | 0.83067 | 201.1693 | 125.8693 | 1.6678 | 0.60015 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L6 | 10 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 207.3 | 130.5 | 0.48483 | 206.8152 | 130.0152 | 1.701 | 0.64163 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L6 | 10 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 123.7 | 80.09 | 0.40133 | 123.2987 | 79.6887 | 0.96591 | 0.46995 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L6 | 10 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 117.6 | 75.34 | 0.69617 | 116.9038 | 74.6438 | 0.93601 | 0.40895 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L6 | 10 | 13 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 116.8 | 70.26 | 0.69667 | 116.1033 | 69.5633 | 1.0203 | 0.22033 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR75
AR77 | 20231013 | L6 | 10 | 13 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 123.9 | 77.23 | 0.69667 | 123.2033 | 76.5333 | 1.0196 | 0.34062 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR76
AR77 | 20231013 | L6 | 10 | 13 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 195.1 | 116.1 | 0.69667 | 194.4033 | 115.4033 | 0.9226 | 0.1738 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR81
AR77 | 20231013 | L6 | 10 | 13 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 201.8 | 118.8 | 0.69667 | 201.1033 | 118.1033 | 0.96932 | 0.15274 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR82; logsheet says slight spill
AR77 | 20231013 | L6 | 10 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 129 | 81.34 | 0.48483 | 128.5152 | 80.8552 | 1.0556 | 0.40127 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
AR77 | 20231013 | L6 | 10 | 19 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 150.6 | 94.63 | 0.83067 | 149.7693 | 93.7993 | 2.4793 | 0.90088 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
AR77 | 20231013 | L6 | 10 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 137.2 | 87.75 | 0.48483 | 136.7152 | 87.2652 | 1.0953 | 0.47712 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L6 | 10 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 133.2 | 85.01 | 0.40133 | 132.7987 | 84.6087 | 1.0674 | 0.45716 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231013 | L6 | 10 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 92.51 | 55.98 | 0.6925 | 91.8175 | 55.2875 | 0.80081 | 0.18518 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR73
AR77 | 20231013 | L6 | 10 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 100.1 | 61.52 | 0.6925 | 99.4075 | 60.8275 | 0.84575 | 0.23904 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR74
AR77 | 20231013 | L6 | 10 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 139.5 | 83.95 | 0.6925 | 138.8075 | 83.2575 | 0.64874 | 0.14226 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR79; logsheet says slight spill
AR77 | 20231013 | L6 | 10 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 126.9 | 76.36 | 0.6925 | 126.2075 | 75.6675 | 0.58809 | 0.12819 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR80
AR77 | 20231014 | L9 | 13 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 43.57 | 29.01 | 0.48483 | 43.0852 | 28.5252 | 0.32249 | 0.19149 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L9 | 13 | 9 |  |  | b | 283 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 49.94 | 36.91 | 0.83067 | 49.1093 | 36.0793 | 0.29574 | 0.37043 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl; depth =40m
AR77 | 20231014 | L9 | 13 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 109.4 | 70.81 | 0.67983 | 108.7202 | 70.1302 | 0.85473 | 0.40891 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L9 | 13 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 115.4 | 73.64 | 0.67983 | 114.7202 | 72.9602 | 0.92494 | 0.38969 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L9 | 13 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 156.4 | 100.2 | 0.48483 | 155.9152 | 99.7152 | 1.2448 | 0.55194 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L9 | 13 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 156.2 | 99.34 | 0.67983 | 155.5202 | 98.6602 | 1.2594 | 0.51831 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L9 | 13 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 151.3 | 95.97 | 0.67983 | 150.6202 | 95.2902 | 1.2255 | 0.49148 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L9 | 13 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 135.8 | 86.42 | 0.40133 | 135.3987 | 86.0187 | 1.0937 | 0.45621 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L9 | 13 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 24.58 | 16.16 | 0.69667 | 23.8833 | 15.4633 | 0.18458 | 0.091188 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR91
AR77 | 20231014 | L9 | 13 | 16 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.82 | 16.31 | 0.6925 | 25.1275 | 15.6175 | 0.20777 | 0.069804 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR92
AR77 | 20231014 | L9 | 13 | 16 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13.55 | 8.527 | 0.69667 | 12.8533 | 7.8303 | 0.058661 | 0.015732 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR97
AR77 | 20231014 | L9 | 13 | 16 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 18.45 | 11.33 | 0.69667 | 17.7533 | 10.6333 | 0.083151 | 0.017872 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR98
AR77 | 20231014 | L9 | 13 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 128 | 82.36 | 0.69617 | 127.3038 | 81.6638 | 1.0109 | 0.46058 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L9 | 13 | 18 |  |  | b | 279 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 132.1 | 85.15 | 0.83067 | 131.2693 | 84.3193 | 1.0809 | 0.49831 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L9 | 13 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 17.72 | 11.7 | 0.69667 | 17.0233 | 11.0033 | 0.13197 | 0.064261 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR93
AR77 | 20231014 | L9 | 13 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 18.65 | 11.7 | 0.69667 | 17.9533 | 11.0033 | 0.15236 | 0.043874 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR94
AR77 | 20231014 | L9 | 13 | 18 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 10.15 | 6.713 | 0.69667 | 9.4533 | 6.0163 | 0.040066 | 0.016989 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR99
AR77 | 20231014 | L9 | 13 | 18 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.22 | 7.657 | 0.6925 | 11.5275 | 6.9645 | 0.053192 | 0.012855 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR100
AR77 | 20231014 | L9 | 13 | 21 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 130.1 | 83.04 | 0.40133 | 129.6987 | 82.6387 | 1.0423 | 0.44669 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L9 | 13 | 21 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 134 | 85.88 | 0.67983 | 133.3202 | 85.2002 | 1.0658 | 0.46937 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L9 | 13 | 21 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 113.5 | 73.05 | 0.67983 | 112.8202 | 72.3702 | 0.89592 | 0.40807 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L9 | 13 | 21 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 121.6 | 77.59 | 0.67983 | 120.9202 | 76.9102 | 0.97477 | 0.41103 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L9 | 13 | 21 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.61 | 9.966 | 0.69667 | 14.9133 | 9.2693 | 0.12373 | 0.04158 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR89
AR77 | 20231014 | L9 | 13 | 21 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 18.18 | 11.79 | 0.69667 | 17.4833 | 11.0933 | 0.14008 | 0.057755 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR90
AR77 | 20231014 | L9 | 13 | 21 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 10.45 | 6.789 | 0.6925 | 9.7575 | 6.0965 | 0.042755 | 0.015166 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR95
AR77 | 20231014 | L9 | 13 | 21 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.44 | 7.674 | 0.69667 | 10.7433 | 6.9773 | 0.043822 | 0.022227 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR96
AR77 | 20231014 | L8 | 16 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 41.9 | 27.68 | 0.83067 | 41.0693 | 26.8493 | 0.31496 | 0.16883 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L8 | 16 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 51.57 | 32.94 | 0.67983 | 50.8902 | 32.2602 | 0.41263 | 0.16864 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L8 | 16 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 20.32 | 13.9 | 0.40133 | 19.9187 | 13.4987 | 0.1422 | 0.10103 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L8 | 16 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 32.46 | 21.45 | 0.83067 | 31.6293 | 20.6193 | 0.24386 | 0.12767 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L8 | 16 | 5 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 36 | 21.56 | 0.6925 | 35.3075 | 20.8675 | 0.31656 | 0.055594 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR105
AR77 | 20231014 | L8 | 16 | 5 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 39.5 | 24.11 | 0.6925 | 38.8075 | 23.4175 | 0.33738 | 0.080245 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR106
AR77 | 20231014 | L8 | 16 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 66.77 | 39.11 | 0.6925 | 66.0775 | 38.4175 | 0.32244 | 0.041886 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR111
AR77 | 20231014 | L8 | 16 | 5 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 72.38 | 42.26 | 0.6925 | 71.6875 | 41.5675 | 0.35112 | 0.043082 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR112
AR77 | 20231014 | L8 | 16 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 166.2 | 106.2 | 0.83067 | 165.3693 | 105.3693 | 1.3289 | 0.56965 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L8 | 16 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 176.6 | 111.5 | 0.67983 | 175.9202 | 110.8202 | 1.4419 | 0.55491 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L8 | 16 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 128.3 | 82.59 | 0.40133 | 127.8987 | 82.1887 | 1.0124 | 0.46848 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L8 | 16 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 131.9 | 84.04 | 0.83067 | 131.0693 | 83.2093 | 1.06 | 0.43925 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L8 | 16 | 9 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 44.84 | 28.05 | 0.6925 | 44.1475 | 27.3575 | 0.36682 | 0.11941 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR103
AR77 | 20231014 | L8 | 16 | 9 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 51.41 | 32.56 | 0.6925 | 50.7175 | 31.8675 | 0.41323 | 0.15509 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR104
AR77 | 20231014 | L8 | 16 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 58.99 | 36.89 | 0.6925 | 58.2975 | 36.1975 | 0.25809 | 0.085804 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR109
AR77 | 20231014 | L8 | 16 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 55.88 | 35.5 | 0.6925 | 55.1875 | 34.8075 | 0.23801 | 0.092685 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR110
AR77 | 20231014 | L8 | 16 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 145.8 | 92.64 | 0.40133 | 145.3987 | 92.2387 | 1.1774 | 0.48456 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L8 | 16 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 141.8 | 90.93 | 0.48483 | 141.3152 | 90.4452 | 1.1267 | 0.50296 | 20240307 | 15-15 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L8 | 16 | 14 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 38.14 | 24.26 | 0.6925 | 37.4475 | 23.5675 | 0.30324 | 0.11563 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR113
AR77 | 20231014 | L8 | 16 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 43.2 | 28.09 | 0.6925 | 42.5075 | 27.3975 | 0.33012 | 0.15683 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR114; logsheet says small spill
AR77 | 20231014 | L8 | 16 | 14 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 70.63 | 44.1 | 0.6925 | 69.9375 | 43.4075 | 0.30927 | 0.10238 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR115
AR77 | 20231014 | L8 | 16 | 14 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 72.08 | 45.13 | 0.6925 | 71.3875 | 44.4375 | 0.3136 | 0.10706 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR116
AR77 | 20231014 | L8 | 16 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 123 | 78 | 0.48483 | 122.5152 | 77.5152 | 0.9967 | 0.4 | 20240307 | 15-16 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L8 | 16 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 124.4 | 80.36 | 0.48483 | 123.9152 | 79.8752 | 0.97544 | 0.46379 | 20240307 | 15-16 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L8 | 16 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 100.7 | 63.81 | 0.40133 | 100.2987 | 63.4087 | 0.81707 | 0.32545 | 20240307 | 15-13 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L8 | 16 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 97.74 | 62.67 | 0.40133 | 97.3387 | 62.2687 | 0.77676 | 0.34522 | 20240307 | 15-14 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L8 | 16 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 30.54 | 19.54 | 0.6925 | 29.8475 | 18.8475 | 0.24114 | 0.094982 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR101
AR77 | 20231014 | L8 | 16 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 37.5 | 24.08 | 0.6925 | 36.8075 | 23.3875 | 0.29419 | 0.1229 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR102
AR77 | 20231014 | L8 | 16 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 47.91 | 30.19 | 0.6925 | 47.2175 | 29.4975 | 0.20694 | 0.073301 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR107
AR77 | 20231014 | L8 | 16 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 40.75 | 25.36 | 0.6925 | 40.0575 | 24.6675 | 0.17908 | 0.054426 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR108
AR77 | 20231014 | L7 | 17 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 49.59 | 33.02 | 0.67983 | 48.9102 | 32.3402 | 0.36701 | 0.21571 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L7 | 17 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 38.17 | 25.7 | 0.67983 | 37.4902 | 25.0202 | 0.2762 | 0.17463 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L7 | 17 | 10 |  |  | a | 290 | 0 |  |  |  |  | Cal_030724 | 1.8135 | 0.57627 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR77 | 20231014 | L7 | 17 | 10 |  |  | b | 290 | 0 |  |  |  |  | Cal_030724 | 1.8135 | 0.57627 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR77 | 20231014 | L7 | 17 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 36.03 | 22.85 | 0.6925 | 35.3375 | 22.1575 | 0.28893 | 0.10622 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR121
AR77 | 20231014 | L7 | 17 | 10 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 38.26 | 24.75 | 0.6925 | 37.5675 | 24.0575 | 0.29617 | 0.13287 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR122
AR77 | 20231014 | L7 | 17 | 10 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 43.08 | 27.68 | 0.6925 | 42.3875 | 26.9875 | 0.17952 | 0.07641 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR127
AR77 | 20231014 | L7 | 17 | 10 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 39.88 | 25.85 | 0.6925 | 39.1875 | 25.1575 | 0.16355 | 0.075026 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR128
AR77 | 20231014 | L7 | 17 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 189.9 | 119.1 | 0.40133 | 189.4987 | 118.6987 | 1.5681 | 0.57062 | 20240307 | 15-14 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L7 | 17 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 172.1 | 110.9 | 0.40133 | 171.6987 | 110.4987 | 1.3555 | 0.6355 | 20240307 | 15-14 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L7 | 17 | 11 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 76.39 | 47.5 | 0.6925 | 75.6975 | 46.8075 | 0.63333 | 0.20143 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR119
AR77 | 20231014 | L7 | 17 | 11 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 73.56 | 45.52 | 0.6925 | 72.8675 | 44.8275 | 0.6126 | 0.18413 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR120
AR77 | 20231014 | L7 | 17 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 127.5 | 76.32 | 0.6925 | 126.8075 | 75.6275 | 0.59771 | 0.1208 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR125
AR77 | 20231014 | L7 | 17 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 137.6 | 82.56 | 0.6925 | 136.9075 | 81.8675 | 0.64278 | 0.13501 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR126
AR77 | 20231014 | L7 | 17 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 192.3 | 122.7 | 0.83067 | 191.4693 | 121.8693 | 1.5416 | 0.65433 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L7 | 17 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 189.3 | 119.6 | 0.69617 | 188.6038 | 118.9038 | 1.5438 | 0.59868 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L7 | 17 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 122.4 | 78.08 | 0.40133 | 121.9987 | 77.6787 | 0.98164 | 0.41801 | 20240307 | 15-14 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L7 | 17 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 114.2 | 73.27 | 0.67983 | 113.5202 | 72.5902 | 0.90655 | 0.40141 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L7 | 17 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 79.51 | 49.03 | 0.6925 | 78.8175 | 48.3375 | 0.66819 | 0.19386 | 20240307 | 13-60 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR117
AR77 | 20231014 | L7 | 17 | 15 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 90.39 | 56.73 | 0.6925 | 89.6975 | 56.0375 | 0.7379 | 0.26147 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR118
AR77 | 20231014 | L7 | 17 | 15 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 113.2 | 68.86 | 0.6925 | 112.5075 | 68.1675 | 0.51595 | 0.12934 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR123
AR77 | 20231014 | L7 | 17 | 15 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 97.81 | 58.83 | 0.6925 | 97.1175 | 58.1375 | 0.45523 | 0.097114 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR124
AR77 | 20231014 | L3 | 18 | 1 |  |  | a | 290 | 0 |  |  |  |  | Cal_030724 | 1.8135 | 0.57627 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR77 | 20231014 | L3 | 18 | 1 |  |  | b | 290 | 0 |  |  |  |  | Cal_030724 | 1.8135 | 0.57627 |  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AR77 | 20231014 | L3 | 18 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 53.71 | 34.34 | 0.67983 | 53.0302 | 33.6602 | 0.42902 | 0.17748 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L3 | 18 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 54.94 | 35.46 | 0.83067 | 54.1093 | 34.6293 | 0.43146 | 0.19251 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L3 | 18 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 151.7 | 96.44 | 0.40133 | 151.2987 | 96.0387 | 1.2239 | 0.50652 | 20240307 | 15-14 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L3 | 18 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 137.3 | 88.09 | 0.40133 | 136.8987 | 87.6887 | 1.0899 | 0.49006 | 20240307 | 15-14 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L3 | 18 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 60.47 | 39.03 | 0.69667 | 59.7733 | 38.3333 | 0.47001 | 0.21362 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR131
AR77 | 20231014 | L3 | 18 | 6 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 69.49 | 43.47 | 0.69667 | 68.7933 | 42.7733 | 0.56847 | 0.19175 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR132
AR77 | 20231014 | L3 | 18 | 6 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 99.15 | 60.7 | 0.69667 | 98.4533 | 60.0033 | 0.44904 | 0.12103 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR137
AR77 | 20231014 | L3 | 18 | 6 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 89.8 | 56.54 | 0.69667 | 89.1033 | 55.8433 | 0.38843 | 0.14212 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR138
AR77 | 20231014 | L3 | 18 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 140.3 | 91.01 | 0.48483 | 139.8152 | 90.5252 | 1.0917 | 0.5394 | 20240307 | 15-16 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L3 | 18 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 132.8 | 85.64 | 0.40133 | 132.3987 | 85.2387 | 1.0445 | 0.49132 | 20240307 | 15-14 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L3 | 18 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 64.48 | 40.83 | 0.69667 | 63.7833 | 40.1333 | 0.51846 | 0.19728 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR133
AR77 | 20231014 | L3 | 18 | 10 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 64.46 | 40.6 | 0.69667 | 63.7633 | 39.9033 | 0.52306 | 0.18857 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR134
AR77 | 20231014 | L3 | 18 | 10 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 85.93 | 52.91 | 0.69667 | 85.2333 | 52.2133 | 0.38492 | 0.11023 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR139
AR77 | 20231014 | L3 | 18 | 10 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 102 | 63.93 | 0.69667 | 101.3033 | 63.2333 | 0.44379 | 0.15587 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR140
AR77 | 20231014 | L3 | 18 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 162.8 | 105.3 | 0.67983 | 162.1202 | 104.6202 | 1.2736 | 0.61153 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L3 | 18 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 179 | 113 | 0.40133 | 178.5987 | 112.5987 | 1.4618 | 0.56702 | 20240307 | 15-14 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L3 | 18 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 101.2 | 64.6 | 0.67983 | 100.5202 | 63.9202 | 0.81065 | 0.34109 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L3 | 18 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 100.8 | 65.48 | 0.40133 | 100.3987 | 65.0787 | 0.7823 | 0.39032 | 20240307 | 15-14 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L3 | 18 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 57.3 | 35.81 | 0.69667 | 56.6033 | 35.1133 | 0.47111 | 0.1551 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR129
AR77 | 20231014 | L3 | 18 | 16 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 55.89 | 35.84 | 0.69667 | 55.1933 | 35.1433 | 0.43954 | 0.18721 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR130; logsheet says filtered slower than the rest
AR77 | 20231014 | L3 | 18 | 16 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 76.15 | 45.86 | 0.69667 | 75.4533 | 45.1633 | 0.35374 | 0.075338 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR135
AR77 | 20231014 | L3 | 18 | 16 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 78.64 | 47.82 | 0.69667 | 77.9433 | 47.1233 | 0.35863 | 0.087451 | 20240307 | 13-68 | 20240328 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR136
AR77 | 20231014 | L4 | 19 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 67.47 | 44.12 | 0.83067 | 66.6393 | 43.2893 | 0.51718 | 0.26283 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L4 | 19 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 69.27 | 44.05 | 0.69617 | 68.5738 | 43.3538 | 0.55859 | 0.22257 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L4 | 19 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 20.92 | 14.58 | 0.48483 | 20.4352 | 14.0952 | 0.14042 | 0.11355 | 20240307 | 15-16 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L4 | 19 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 40.33 | 26.25 | 0.40133 | 39.9287 | 25.8487 | 0.31186 | 0.1539 | 20240307 | 15-14 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L4 | 19 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 59.63 | 36.49 | 0.6925 | 58.9375 | 35.7975 | 0.50728 | 0.13113 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR145
AR77 | 20231014 | L4 | 19 | 6 |  |  | b | 210 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 43.59 | 26.2 | 0.6925 | 42.8975 | 25.5075 | 0.5319 | 0.10279 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR146
AR77 | 20231014 | L4 | 19 | 6 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 84.33 | 49.84 | 0.6925 | 83.6375 | 49.1475 | 0.40133 | 0.063908 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR151
AR77 | 20231014 | L4 | 19 | 6 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 103 | 60.6 | 0.6925 | 102.3075 | 59.9075 | 0.49517 | 0.07399 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR152
AR77 | 20231014 | L4 | 19 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 142.9 | 91.89 | 0.40133 | 142.4987 | 91.4887 | 1.1298 | 0.51867 | 20240307 | 15-14 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L4 | 19 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 150.1 | 97.4 | 0.67983 | 149.4202 | 96.7202 | 1.1672 | 0.5755 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L4 | 19 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 90.9 | 60.26 | 0.67983 | 90.2202 | 59.5802 | 0.67864 | 0.3949 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L4 | 19 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 99.47 | 66.05 | 0.67983 | 98.7902 | 65.3702 | 0.74022 | 0.43765 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L4 | 19 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 48.18 | 29.74 | 0.6925 | 47.4875 | 29.0475 | 0.40424 | 0.11379 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR143
AR77 | 20231014 | L4 | 19 | 9 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 50.5 | 32.3 | 0.6925 | 49.8075 | 31.6075 | 0.39763 | 0.16414 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR144
AR77 | 20231014 | L4 | 19 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 60.33 | 35.57 | 0.6925 | 59.6375 | 34.8775 | 0.28916 | 0.042198 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR149
AR77 | 20231014 | L4 | 19 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 64.6 | 39.4 | 0.6925 | 63.9075 | 38.7075 | 0.2943 | 0.073447 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR150
AR77 | 20231014 | L4 | 19 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 141.8 | 92.7 | 0.83067 | 140.9693 | 91.8693 | 1.0875 | 0.56783 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L4 | 19 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 147.8 | 96 | 0.83067 | 146.9693 | 95.1693 | 1.1473 | 0.56749 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L4 | 19 | 15 |  |  | a | 146 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 27.42 | 17.49 | 0.6925 | 26.7275 | 16.7975 | 0.43686 | 0.16432 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR153
AR77 | 20231014 | L4 | 19 | 15 |  |  | b | 146 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 26.99 | 17.29 | 0.6925 | 26.2975 | 16.5975 | 0.42675 | 0.16728 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR154
AR77 | 20231014 | L4 | 19 | 15 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 57.72 | 34.6 | 0.6925 | 57.0275 | 33.9075 | 0.26952 | 0.052041 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR155
AR77 | 20231014 | L4 | 19 | 15 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 74.6 | 44.17 | 0.6925 | 73.9075 | 43.4775 | 0.35409 | 0.057477 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR156
AR77 | 20231014 | L4 | 19 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 138.4 | 91.12 | 0.67983 | 137.7202 | 90.4402 | 1.0472 | 0.58239 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L4 | 19 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 143.3 | 92.92 | 0.40133 | 142.8987 | 92.5187 | 1.1159 | 0.55118 | 20240307 | 15-14 | 20250311 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L4 | 19 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 100.9 | 65.47 | 0.69617 | 100.2038 | 64.7738 | 0.78473 | 0.38239 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L4 | 19 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 88.57 | 58.56 | 0.83067 | 87.7393 | 57.7293 | 0.66469 | 0.3755 | 20240307 | 13-98 | 20241101 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231014 | L4 | 19 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 61.75 | 38.91 | 0.6925 | 61.0575 | 38.2175 | 0.5007 | 0.18087 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR141
AR77 | 20231014 | L4 | 19 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 64.77 | 39.67 | 0.6925 | 64.0775 | 38.9775 | 0.55024 | 0.14488 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR142
AR77 | 20231014 | L4 | 19 | 19 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 72.38 | 43.34 | 0.6925 | 71.6875 | 42.6475 | 0.33853 | 0.065914 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR147
AR77 | 20231014 | L4 | 19 | 19 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 86.67 | 52.07 | 0.6925 | 85.9775 | 51.3775 | 0.40334 | 0.083889 | 20240307 | 13-61 | 20240201 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR148
AR77 | 20231016 | MVCO | 35 | 1 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 190.1 | 126.6 | 0.48483 | 189.6152 | 126.1152 | 2.8129 | 1.7319 | 20240307 | 15-16 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231016 | MVCO | 35 | 1 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 187.8 | 125 | 0.48483 | 187.3152 | 124.5152 | 2.7819 | 1.7052 | 20240307 | 15-16 | 20250313 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
AR77 | 20231016 | MVCO | 35 | 4 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 171.9 | 112.7 | 0.69617 | 171.2038 | 112.0038 | 2.6224 | 1.4138 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
AR77 | 20231016 | MVCO | 35 | 4 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 78.97 | 51.8 | 0.69617 | 78.2738 | 51.1038 | 1.2036 | 0.63805 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
AR77 | 20231016 | MVCO | 35 | 4 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 93.56 | 65.64 | 0.69617 | 92.8638 | 64.9438 | 0.6184 | 0.55179 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
AR77 | 20231016 | MVCO | 35 | 4 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 165.2 | 113.7 | 0.67983 | 164.5202 | 113.0202 | 1.1407 | 0.89578 | 20240307 | 15-09 | 20250304 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN712 | 20240209 | L1 | 1 | 4 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 380.8 | 230.6 | 0.41983 | 380.3802 | 230.1802 | 6.6535 | 1.6414 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN712 | 20240209 | L1 | 1 | 4 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 351.2 | 214.6 | 0.41983 | 350.7802 | 214.1802 | 6.0511 | 1.6673 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN712 | 20240209 | L1 | 1 | 4 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 73.73 | 54.11 | 0.41983 | 73.3102 | 53.6902 | 0.43456 | 0.53285 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl; depth = 17
EN712 | 20240209 | L1 | 1 | 4 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 75.25 | 53.15 | 0.778 | 74.472 | 52.372 | 0.48949 | 0.45417 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240209 | L1 | 1 | 4 |  |  | a | 293 | 5 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 299.2 | 176.2 | 0.73967 | 298.4603 | 175.4603 | 5.3928 | 0.86545 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR5 | poor replicate
EN712 | 20240209 | L1 | 1 | 4 |  |  | b | 294 | 5 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 273.8 | 165.2 | 0.73967 | 273.0603 | 164.4603 | 4.7453 | 1.1007 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR6 | poor replicate
EN712 | 20240209 | L1 | 1 | 4 |  |  | a | 550 | 20 | 10 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 166.9 | 99.44 | 0.73967 | 166.1603 | 98.7003 | 4.727 | 0.89929 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR11 | poor replicate
EN712 | 20240209 | L1 | 1 | 4 |  |  | b | 550 | 20 | 10 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 212.1 | 124.5 | 0.73967 | 211.3603 | 123.7603 | 6.1382 | 0.91658 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR12 | poor replicate
EN712 | 20240209 | L1 | 1 | 9 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 471.4 | 279.4 | 0.778 | 470.622 | 278.622 | 8.5052 | 1.5355 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240209 | L1 | 1 | 9 |  |  | b | 290 | 0 | 15 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 322.6 | 190.4 | 0.41983 | 322.1802 | 189.9802 | 8.7842 | 1.4852 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240209 | L1 | 1 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 81.5 | 56.7 | 0.778 | 80.722 | 55.922 | 0.54929 | 0.45833 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240209 | L1 | 1 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 89.59 | 62.04 | 0.41983 | 89.1702 | 61.6202 | 0.6102 | 0.5001 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240209 | L1 | 1 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 400.1 | 238.7 | 0.73967 | 399.3603 | 237.9603 | 3.5382 | 0.70554 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR3 | poor replicate
EN712 | 20240209 | L1 | 1 | 9 |  |  | b | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 243 | 147.6 | 0.73967 | 242.2603 | 146.8603 | 6.2741 | 1.5832 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR4 | poor replicate
EN712 | 20240209 | L1 | 1 | 9 |  |  | a | 551 | 20 | 10 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 189.5 | 113.7 | 0.73967 | 188.7603 | 112.9603 | 5.3017 | 1.1257 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR9 | poor replicate
EN712 | 20240209 | L1 | 1 | 9 |  |  | b | 551 | 20 | 10 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 250.4 | 145 | 0.73967 | 249.6603 | 144.2603 | 7.3721 | 0.83637 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR10 | poor replicate
EN712 | 20240209 | L1 | 1 | 17 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 395.5 | 235.3 | 0.41983 | 395.0802 | 234.8802 | 7.0965 | 1.3678 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240209 | L1 | 1 | 17 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 414.8 | 245.1 | 0.36183 | 414.4382 | 244.7382 | 7.5173 | 1.3023 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240209 | L1 | 1 | 17 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 86.26 | 59.53 | 0.41983 | 85.8402 | 59.1102 | 0.59204 | 0.47303 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240209 | L1 | 1 | 17 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 82.06 | 56.44 | 0.778 | 81.282 | 55.662 | 0.56745 | 0.43549 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240209 | L1 | 1 | 17 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 321.2 | 189.8 | 0.73967 | 320.4603 | 189.0603 | 2.8806 | 0.49112 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR1 | poor replicate
EN712 | 20240209 | L1 | 1 | 17 |  |  | b | 243 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 197.6 | 118.9 | 0.73967 | 196.8603 | 118.1603 | 2.0803 | 0.46059 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR2; small leak 50mL lost | poor replicate
EN712 | 20240209 | L1 | 1 | 17 |  |  | a | 550 | 20 | 10 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 223.2 | 131 | 0.73967 | 222.4603 | 130.2603 | 6.4605 | 0.96477 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR7 | poor replicate
EN712 | 20240209 | L1 | 1 | 17 |  |  | b | 552 | 20 | 10 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 172.2 | 104.5 | 0.73967 | 171.4603 | 103.7603 | 4.7266 | 1.1667 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR8 | poor replicate
EN712 | 20240210 | L5 | 2 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 87.88 | 57.96 | 0.41983 | 87.4602 | 57.5402 | 0.66269 | 0.37409 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L5 | 2 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 102.4 | 65.9 | 0.778 | 101.622 | 65.122 | 0.80843 | 0.36496 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L5 | 2 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 99.68 | 64.25 | 0.41983 | 99.2602 | 63.8302 | 0.78473 | 0.36538 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L5 | 2 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 95.23 | 62.6 | 0.41983 | 94.8102 | 62.1802 | 0.72272 | 0.39767 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L5 | 2 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 93.73 | 60.78 | 0.41983 | 93.3102 | 60.3602 | 0.72981 | 0.35779 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L5 | 2 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 88.66 | 58.24 | 0.41983 | 88.2402 | 57.8202 | 0.67377 | 0.36806 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L5 | 2 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 123.2 | 77.69 | 0.36183 | 122.8382 | 77.3282 | 1.008 | 0.38533 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L5 | 2 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 112.3 | 71 | 0.41983 | 111.8802 | 70.5802 | 0.91475 | 0.35699 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L5 | 2 | 11 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 79.07 | 49.32 | 0.73967 | 78.3303 | 48.5803 | 0.65218 | 0.2142 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR17
EN712 | 20240210 | L5 | 2 | 11 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 77.73 | 48.48 | 0.73967 | 76.9903 | 47.7403 | 0.63904 | 0.20946 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR18
EN712 | 20240210 | L5 | 2 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 99.68 | 59.45 | 0.73967 | 98.9403 | 58.7103 | 0.46983 | 0.087958 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR19
EN712 | 20240210 | L5 | 2 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 94.24 | 56.01 | 0.73967 | 93.5003 | 55.2703 | 0.44647 | 0.078633 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR20; small leak
EN712 | 20240210 | L5 | 2 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 100.3 | 63.53 | 0.41983 | 99.8802 | 63.1102 | 0.81441 | 0.32273 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L5 | 2 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 102.6 | 65.6 | 0.36183 | 102.2382 | 65.2382 | 0.81951 | 0.35598 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L5 | 2 | 17 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 54.79 | 36.06 | 0.36183 | 54.4282 | 35.6982 | 0.41485 | 0.22838 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L5 | 2 | 17 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 51.47 | 33.55 | 0.41983 | 51.0502 | 33.1302 | 0.39691 | 0.20004 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L5 | 2 | 17 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 82.43 | 50.68 | 0.73967 | 81.6903 | 49.9403 | 0.69603 | 0.19461 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR13
EN712 | 20240210 | L5 | 2 | 17 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 78.65 | 48.98 | 0.73967 | 77.9103 | 48.2403 | 0.65043 | 0.20989 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR14
EN712 | 20240210 | L5 | 2 | 17 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 84.74 | 50.29 | 0.73967 | 84.0003 | 49.5503 | 0.40232 | 0.068434 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR15
EN712 | 20240210 | L5 | 2 | 17 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 71.92 | 42.68 | 0.73967 | 71.1803 | 41.9403 | 0.34024 | 0.056773 | 20240307 | 13-95 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR16
EN712 | 20240210 | L6 | 5 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 94.79 | 61.94 | 0.41983 | 94.3702 | 61.5202 | 0.72759 | 0.3809 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | volume filtered not marked on logshee; assume 290
EN712 | 20240210 | L6 | 5 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 95.36 | 62.33 | 0.41983 | 94.9402 | 61.9102 | 0.73158 | 0.38394 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | volume filtered not marked on logshee; assume 290
EN712 | 20240210 | L6 | 5 | 7 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 52.51 | 34.45 | 0.73967 | 51.7703 | 33.7103 | 0.39457 | 0.20458 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR27
EN712 | 20240210 | L6 | 5 | 7 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 65.15 | 41.75 | 0.73967 | 64.4103 | 41.0103 | 0.51123 | 0.21765 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR28
EN712 | 20240210 | L6 | 5 | 7 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 71.59 | 43.55 | 0.73967 | 70.8503 | 42.8103 | 0.32687 | 0.079115 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR35
EN712 | 20240210 | L6 | 5 | 7 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 69.67 | 42.52 | 0.73967 | 68.9303 | 41.7803 | 0.31592 | 0.079578 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR36
EN712 | 20240210 | L6 | 5 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 102.9 | 65.38 | 0.41983 | 102.4802 | 64.9602 | 0.83103 | 0.33945 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L6 | 5 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 91.37 | 59.47 | 0.41983 | 90.9502 | 59.0502 | 0.70655 | 0.35744 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L6 | 5 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 42.59 | 29.31 | 0.41983 | 42.1702 | 28.8902 | 0.29414 | 0.22642 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L6 | 5 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 46.55 | 31.44 | 0.41983 | 46.1302 | 31.0202 | 0.33467 | 0.22426 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L6 | 5 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 48.54 | 30.97 | 0.73967 | 47.8003 | 30.2303 | 0.38517 | 0.15395 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR25
EN712 | 20240210 | L6 | 5 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 46.41 | 30.55 | 0.73967 | 45.6703 | 29.8103 | 0.3465 | 0.18332 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR26
EN712 | 20240210 | L6 | 5 | 10 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 64.09 | 39.07 | 0.73967 | 63.3503 | 38.3303 | 0.2922 | 0.071966 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR33
EN712 | 20240210 | L6 | 5 | 10 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 64.27 | 38.49 | 0.73967 | 63.5303 | 37.7503 | 0.30107 | 0.05758 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR34
EN712 | 20240210 | L6 | 5 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 96.25 | 63.11 | 0.41983 | 95.8302 | 62.6902 | 0.73401 | 0.39556 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L6 | 5 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 96.88 | 63.13 | 0.41983 | 96.4602 | 62.7102 | 0.74752 | 0.38241 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L6 | 5 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 48.86 | 32.08 | 0.41983 | 48.4402 | 31.6602 | 0.37166 | 0.19881 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L6 | 5 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 44.91 | 30.01 | 0.36183 | 44.5482 | 29.6482 | 0.33002 | 0.20419 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L6 | 5 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 70.09 | 44.18 | 0.73967 | 69.3503 | 43.4403 | 0.568 | 0.20671 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR23
EN712 | 20240210 | L6 | 5 | 15 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 65.64 | 42.49 | 0.73967 | 64.9003 | 41.7503 | 0.5075 | 0.23708 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR24
EN712 | 20240210 | L6 | 5 | 15 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 83.46 | 50.97 | 0.73967 | 82.7203 | 50.2303 | 0.37875 | 0.097607 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR31
EN712 | 20240210 | L6 | 5 | 15 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 84.63 | 51.64 | 0.73967 | 83.8903 | 50.9003 | 0.38457 | 0.098132 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR32
EN712 | 20240210 | L6 | 5 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 100.6 | 65.84 | 0.36183 | 100.2382 | 65.4782 | 0.76989 | 0.40992 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L6 | 5 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 96.08 | 61.94 | 0.41983 | 95.6602 | 61.5202 | 0.75616 | 0.35233 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L6 | 5 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 46.93 | 31.76 | 0.41983 | 46.5102 | 31.3402 | 0.336 | 0.2287 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L6 | 5 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 46.05 | 31.83 | 0.41983 | 45.6302 | 31.4102 | 0.31496 | 0.251 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L6 | 5 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 72.97 | 45.64 | 0.73967 | 72.2303 | 44.9003 | 0.59913 | 0.20162 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR21
EN712 | 20240210 | L6 | 5 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 68.31 | 43.71 | 0.73967 | 67.5703 | 42.9703 | 0.53928 | 0.22705 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR22
EN712 | 20240210 | L6 | 5 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 83.83 | 50.54 | 0.73967 | 83.0903 | 49.8003 | 0.38878 | 0.084356 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR29
EN712 | 20240210 | L6 | 5 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 71.89 | 43.91 | 0.73967 | 71.1503 | 43.1703 | 0.32558 | 0.083078 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR30
EN712 | 20240210 | L7 | 6 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 30.91 | 21.38 | 0.41983 | 30.4902 | 20.9602 | 0.21108 | 0.16659 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 29.64 | 21.35 | 0.41983 | 29.2202 | 20.9302 | 0.18361 | 0.19351 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth =102
EN712 | 20240210 | L7 | 6 | 3 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 24.08 | 17.08 | 0.36183 | 23.7182 | 16.7182 | 0.15504 | 0.14619 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 3 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 24.15 | 17.05 | 0.41983 | 23.7302 | 16.6302 | 0.15726 | 0.14239 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 3 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.105 | 6.902 | 0.73967 | 8.3653 | 6.1623 | 0.048294 | 0.061604 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR41 | Phaeo>Chl
EN712 | 20240210 | L7 | 6 | 3 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.14 | 7.842 | 0.73967 | 10.4003 | 7.1023 | 0.072299 | 0.054363 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR42
EN712 | 20240210 | L7 | 6 | 3 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.622 | 4.369 | 0.73967 | 5.8823 | 3.6293 | 0.026264 | 0.0081544 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR47
EN712 | 20240210 | L7 | 6 | 3 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.186 | 4.943 | 0.73967 | 6.4463 | 4.2033 | 0.026147 | 0.013714 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR48
EN712 | 20240210 | L7 | 6 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 48 | 32.22 | 0.41983 | 47.5802 | 31.8002 | 0.34951 | 0.22348 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 46.19 | 30.91 | 0.41983 | 45.7702 | 30.4902 | 0.33843 | 0.21095 | 20240307 | 15-19 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 70.34 | 45.66 | 0.41983 | 69.9202 | 45.2402 | 0.54663 | 0.26852 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 65.91 | 43.38 | 0.41983 | 65.4902 | 42.9602 | 0.49901 | 0.27506 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 53.35 | 35.16 | 0.41983 | 52.9302 | 34.7402 | 0.40289 | 0.22307 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 51.83 | 34.07 | 0.41983 | 51.4102 | 33.6502 | 0.39336 | 0.21296 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 24.9 | 16.76 | 0.73967 | 24.1603 | 16.0203 | 0.17845 | 0.10726 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR39
EN712 | 20240210 | L7 | 6 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 23.61 | 15.5 | 0.73967 | 22.8703 | 14.7603 | 0.17718 | 0.085156 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR40
EN712 | 20240210 | L7 | 6 | 10 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 36.9 | 23.63 | 0.73967 | 36.1603 | 22.8903 | 0.15497 | 0.062499 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR45
EN712 | 20240210 | L7 | 6 | 10 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 1.3 | 1.046 | 0.73967 | 0.56033 | 0.30633 | 0.0029663 | -5.5984e-05 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 4 |  | TAR46 | there was definitely a filter in vial
EN712 | 20240210 | L7 | 6 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 90.7 | 59.92 | 0.41983 | 90.2802 | 59.5002 | 0.68174 | 0.39036 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 92.12 | 60.84 | 0.41983 | 91.7002 | 60.4202 | 0.69282 | 0.39586 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 97.94 | 62.73 | 0.41983 | 97.5202 | 62.3102 | 0.77986 | 0.34287 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 108.2 | 68.4 | 0.41983 | 107.7802 | 67.9802 | 0.88153 | 0.34337 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 96.47 | 60.84 | 0.41983 | 96.0502 | 60.4202 | 0.78916 | 0.29951 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 105.5 | 67.53 | 0.41983 | 105.0802 | 67.1102 | 0.84099 | 0.36823 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 16 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 50.25 | 33.33 | 0.41983 | 49.8302 | 32.9102 | 0.37476 | 0.21823 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 16 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 50.72 | 33.71 | 0.41983 | 50.3002 | 33.2902 | 0.37675 | 0.22308 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240210 | L7 | 6 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 57.96 | 35.71 | 0.73967 | 57.2203 | 34.9703 | 0.48777 | 0.13589 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR37
EN712 | 20240210 | L7 | 6 | 16 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 57.27 | 35.93 | 0.73967 | 56.5303 | 35.1903 | 0.46782 | 0.15976 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR38
EN712 | 20240210 | L7 | 6 | 16 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 50.6 | 31.18 | 0.73967 | 49.8603 | 30.4403 | 0.2268 | 0.062405 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR43
EN712 | 20240210 | L7 | 6 | 16 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 47.8 | 29.15 | 0.73967 | 47.0603 | 28.4103 | 0.21701 | 0.051923 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR44
EN712 | 20240211 | L10 | 7 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 36.43 | 24.52 | 0.41983 | 36.0102 | 24.1002 | 0.26379 | 0.17045 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | there are three tissue capsules marked C7N5a; two of them are entered as C7N4a and C7N4b because those tissue capsules were missing
EN712 | 20240211 | L10 | 7 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 34.63 | 24.11 | 0.41983 | 34.2102 | 23.6902 | 0.23301 | 0.19385 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | there are three tissue capsules marked C7N5a; two of them are entered as C7N4a and C7N4b because those tissue capsules were missing
EN712 | 20240211 | L10 | 7 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 44.68 | 30.73 | 0.41983 | 44.2602 | 30.3102 | 0.30898 | 0.23716 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | there are three tissue capsules marked C7N5a; two of them are entered as C7N4a and C7N4b because those tissue capsules were missing
EN712 | 20240211 | L10 | 7 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 45.72 | 31.3 | 0.36183 | 45.3582 | 30.9382 | 0.31939 | 0.23807 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L10 | 7 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 39.43 | 27.16 | 0.36183 | 39.0682 | 26.7982 | 0.27177 | 0.21109 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L10 | 7 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 41.31 | 28.52 | 0.41983 | 40.8902 | 28.1002 | 0.28328 | 0.22304 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L10 | 7 | 5 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 10.02 | 7.213 | 0.73967 | 9.2803 | 6.4733 | 0.061535 | 0.05391 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR53
EN712 | 20240211 | L10 | 7 | 5 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.727 | 6.404 | 0.73967 | 7.9873 | 5.6643 | 0.050925 | 0.050092 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR54
EN712 | 20240211 | L10 | 7 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.113 | 3.691 | 0.73967 | 4.3733 | 2.9513 | 0.016577 | 0.011412 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR59
EN712 | 20240211 | L10 | 7 | 5 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.96 | 3.492 | 0.73967 | 4.2203 | 2.7523 | 0.017113 | 0.0089885 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR60
EN712 | 20240211 | L10 | 7 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 87.3 | 55.59 | 0.41983 | 86.8802 | 55.1702 | 0.70234 | 0.29174 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | there are three tissue capsules marked C7N11b; two of them were entered as C7N8a and C7N8b because those tissue capsules were missing
EN712 | 20240211 | L10 | 7 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 93.23 | 58.72 | 0.36183 | 92.8682 | 58.3582 | 0.76436 | 0.28716 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | there are three tissue capsules marked C7N11b; two of them were entered as C7N8a and C7N8b because those tissue capsules were missing
EN712 | 20240211 | L10 | 7 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 95.2 | 61.59 | 0.41983 | 94.7802 | 61.1702 | 0.74442 | 0.35776 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L10 | 7 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 96.5 | 63.3 | 0.41983 | 96.0802 | 62.8802 | 0.73534 | 0.39766 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L10 | 7 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 117.4 | 74.13 | 0.36183 | 117.0382 | 73.7682 | 0.95838 | 0.3708 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L10 | 7 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 111.4 | 69.58 | 0.778 | 110.622 | 68.802 | 0.92627 | 0.31344 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | there are three tissue capsules marked C7N11b; two of them were entered as C7N8a and C7N8b because those tissue capsules were missing
EN712 | 20240211 | L10 | 7 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 95.96 | 60.27 | 0.778 | 95.182 | 59.492 | 0.79049 | 0.28146 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L10 | 7 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 92.59 | 58.4 | 0.36183 | 92.2282 | 58.0382 | 0.75727 | 0.28848 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L10 | 7 | 11 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 24.49 | 15.84 | 0.73967 | 23.7503 | 15.1003 | 0.18963 | 0.079672 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR51
EN712 | 20240211 | L10 | 7 | 11 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 26.19 | 17.12 | 0.73967 | 25.4503 | 16.3803 | 0.19816 | 0.092975 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR52
EN712 | 20240211 | L10 | 7 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.96 | 8.247 | 0.73967 | 12.2203 | 7.5073 | 0.055041 | 0.016284 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR57
EN712 | 20240211 | L10 | 7 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.61 | 7.282 | 0.73967 | 10.8703 | 6.5423 | 0.050545 | 0.011612 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR58
EN712 | 20240211 | L10 | 7 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 84.42 | 54.08 | 0.36183 | 84.0582 | 53.7182 | 0.672 | 0.29592 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L10 | 7 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 85.68 | 54.32 | 0.41983 | 85.2602 | 53.9002 | 0.69459 | 0.27661 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L10 | 7 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 60.58 | 38.56 | 0.41983 | 60.1602 | 38.1402 | 0.48772 | 0.19951 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L10 | 7 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 64.5 | 41.24 | 0.36183 | 64.1382 | 40.8782 | 0.51518 | 0.22138 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L10 | 7 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 22.17 | 14.13 | 0.73967 | 21.4303 | 13.3903 | 0.17625 | 0.062548 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR49
EN712 | 20240211 | L10 | 7 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 24 | 15.53 | 0.73967 | 23.2603 | 14.7903 | 0.18568 | 0.078089 | 20240307 | 13-96 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR50
EN712 | 20240211 | L10 | 7 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.98 | 8.047 | 0.73967 | 12.2403 | 7.3073 | 0.05761 | 0.011814 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR55
EN712 | 20240211 | L10 | 7 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 10.34 | 6.612 | 0.73967 | 9.6003 | 5.8723 | 0.04338 | 0.012209 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR56
EN712 | 20240211 | L11 | 10 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 33.75 | 23.02 | 0.36183 | 33.3882 | 22.6582 | 0.23766 | 0.17061 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 32.54 | 22.21 | 0.36183 | 32.1782 | 21.8482 | 0.2288 | 0.16487 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 33.76 | 23.72 | 0.41983 | 33.3402 | 23.3002 | 0.22237 | 0.19746 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 35.39 | 24.3 | 0.36183 | 35.0282 | 23.9382 | 0.24563 | 0.1857 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 29.37 | 20.76 | 0.36183 | 29.0082 | 20.3982 | 0.1907 | 0.17684 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 30.43 | 21.14 | 0.36183 | 30.0682 | 20.7782 | 0.20576 | 0.16863 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 6 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.582 | 6.917 | 0.73967 | 8.8423 | 6.1773 | 0.058224 | 0.051568 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR67
EN712 | 20240211 | L11 | 10 | 6 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.934 | 6.507 | 0.73967 | 8.1943 | 5.7673 | 0.053024 | 0.04948 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR68
EN712 | 20240211 | L11 | 10 | 6 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.023 | 3.062 | 0.73967 | 3.2833 | 2.3223 | 0.011203 | 0.010821 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR75
EN712 | 20240211 | L11 | 10 | 6 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.062 | 2.956 | 0.73967 | 3.3223 | 2.2163 | 0.01287 | 0.0081106 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR76
EN712 | 20240211 | L11 | 10 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 69.24 | 44.77 | 0.41983 | 68.8202 | 44.3502 | 0.54198 | 0.25714 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 68.94 | 45.18 | 0.36183 | 68.5782 | 44.8182 | 0.52626 | 0.28129 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 55.94 | 36.01 | 0.36183 | 55.5782 | 35.6482 | 0.44143 | 0.2009 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 63.96 | 40.76 | 0.36183 | 63.5982 | 40.3982 | 0.51385 | 0.21406 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 20.59 | 13.46 | 0.73967 | 19.8503 | 12.7203 | 0.1563 | 0.070549 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR65
EN712 | 20240211 | L11 | 10 | 9 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 17.64 | 11.93 | 0.73967 | 16.9003 | 11.1903 | 0.12518 | 0.074392 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR66
EN712 | 20240211 | L11 | 10 | 9 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.82 | 7.332 | 0.73967 | 11.0803 | 6.5923 | 0.052318 | 0.0102 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR73
EN712 | 20240211 | L11 | 10 | 9 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 10.89 | 6.952 | 0.73967 | 10.1503 | 6.2123 | 0.045906 | 0.013007 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR74
EN712 | 20240211 | L11 | 10 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 88.53 | 56.28 | 0.41983 | 88.1102 | 55.8602 | 0.7143 | 0.29221 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 88.34 | 56.35 | 0.41983 | 87.9202 | 55.9302 | 0.70854 | 0.29923 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.01 | 15.97 | 0.73967 | 24.2703 | 15.2303 | 0.19818 | 0.073441 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR63
EN712 | 20240211 | L11 | 10 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 24.79 | 15.94 | 0.73967 | 24.0503 | 15.2003 | 0.19335 | 0.076809 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR64
EN712 | 20240211 | L11 | 10 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.09 | 8.767 | 0.73967 | 13.3503 | 8.0273 | 0.062165 | 0.0141 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR71
EN712 | 20240211 | L11 | 10 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14 | 8.858 | 0.73967 | 13.2603 | 8.1183 | 0.060051 | 0.017078 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR72
EN712 | 20240211 | L11 | 10 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 87.54 | 56.07 | 0.36183 | 87.1782 | 55.7082 | 0.69703 | 0.30675 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 84.46 | 53.34 | 0.778 | 83.682 | 52.562 | 0.68927 | 0.25781 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 65.41 | 41.96 | 0.36183 | 65.0482 | 41.5982 | 0.51939 | 0.23014 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 66.67 | 43.22 | 0.41983 | 66.2502 | 42.8002 | 0.51939 | 0.2518 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L11 | 10 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 24.32 | 15.58 | 0.73967 | 23.5803 | 14.8403 | 0.1916 | 0.073062 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR61
EN712 | 20240211 | L11 | 10 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 19.47 | 12.6 | 0.73967 | 18.7303 | 11.8603 | 0.1506 | 0.060911 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR62
EN712 | 20240211 | L11 | 10 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.02 | 8.739 | 0.73967 | 13.2803 | 7.9993 | 0.061674 | 0.014324 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR69
EN712 | 20240211 | L11 | 10 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.83 | 9.858 | 0.73967 | 15.0903 | 9.1183 | 0.069491 | 0.016825 | 20240307 | 13-97 | 20241030 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR70
EN712 | 20240211 | L9 | 11 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.36 | 9.926 | 0.41983 | 13.9402 | 9.5062 | 0.098208 | 0.073078 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13.99 | 10.54 | 0.36183 | 13.6282 | 10.1782 | 0.076414 | 0.10698 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 125
EN712 | 20240211 | L9 | 11 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 32.87 | 22.63 | 0.36183 | 32.5082 | 22.2682 | 0.2268 | 0.17443 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 32.87 | 22.6 | 0.778 | 32.092 | 21.822 | 0.22747 | 0.16573 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 34.57 | 23.27 | 0.778 | 33.792 | 22.492 | 0.25028 | 0.15499 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 35.27 | 24.04 | 0.36183 | 34.9082 | 23.6782 | 0.24873 | 0.17791 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 37.8 | 26.38 | 0.36183 | 37.4382 | 26.0182 | 0.25294 | 0.21587 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 40.7 | 27.83 | 0.41983 | 40.2802 | 27.4102 | 0.28506 | 0.20883 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 40.31 | 27.64 | 0.36183 | 39.9482 | 27.2782 | 0.28063 | 0.21088 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 40.09 | 27.13 | 0.36183 | 39.7282 | 26.7682 | 0.28705 | 0.19527 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.88 | 8.113 | 0.57567 | 11.3043 | 7.5373 | 0.082581 | 0.05184 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR81
EN712 | 20240211 | L9 | 11 | 16 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 10.76 | 7.332 | 0.57567 | 10.1843 | 6.7563 | 0.075149 | 0.045343 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR82
EN712 | 20240211 | L9 | 11 | 16 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.056 | 4.717 | 0.57567 | 6.4803 | 4.1413 | 0.027266 | 0.012007 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR87
EN712 | 20240211 | L9 | 11 | 16 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.214 | 4.889 | 0.57567 | 6.6383 | 4.3133 | 0.027103 | 0.013802 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR88
EN712 | 20240211 | L9 | 11 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 97.79 | 63.49 | 0.778 | 97.012 | 62.712 | 0.75971 | 0.37026 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 93.14 | 60.66 | 0.36183 | 92.7782 | 60.2982 | 0.7194 | 0.36708 | 20240307 | 15-17 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 17.76 | 11.85 | 0.57567 | 17.1843 | 11.2743 | 0.12956 | 0.071506 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR79
EN712 | 20240211 | L9 | 11 | 19 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 20.03 | 13.36 | 0.57567 | 19.4543 | 12.7843 | 0.14572 | 0.081496 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR80
EN712 | 20240211 | L9 | 11 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.55 | 5.434 | 0.57567 | 7.9743 | 4.8583 | 0.03639 | 0.0097669 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR85
EN712 | 20240211 | L9 | 11 | 19 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.572 | 5.175 | 0.57567 | 6.9963 | 4.5993 | 0.027993 | 0.015703 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR86
EN712 | 20240211 | L9 | 11 | 22 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 92.45 | 58.35 | 0.36183 | 92.0882 | 57.9882 | 0.75528 | 0.28958 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 22 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 83.89 | 53.34 | 0.41983 | 83.4702 | 52.9202 | 0.67665 | 0.27689 | 20240307 | 15-20 | 20250320 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 22 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 68.31 | 43.51 | 0.36183 | 67.9482 | 43.1482 | 0.54929 | 0.22817 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 22 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 66.05 | 42.26 | 0.36183 | 65.6882 | 41.8982 | 0.52692 | 0.22802 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240211 | L9 | 11 | 22 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 21.95 | 14.23 | 0.57567 | 21.3743 | 13.6543 | 0.16924 | 0.074272 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR77
EN712 | 20240211 | L9 | 11 | 22 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 23.98 | 15.12 | 0.57567 | 23.4043 | 14.5443 | 0.19423 | 0.065153 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR78
EN712 | 20240211 | L9 | 11 | 22 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.457 | 5.889 | 0.57567 | 8.8813 | 5.3133 | 0.041669 | 0.008811 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR83
EN712 | 20240211 | L9 | 11 | 22 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.8 | 7.08 | 0.57567 | 11.2243 | 6.5043 | 0.054923 | 0.0066485 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR84
EN712 | 20240212 | L8 | 12 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 32.59 | 22.4 | 0.36183 | 32.2282 | 22.0382 | 0.2257 | 0.1714 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 30.4 | 21.36 | 0.36183 | 30.0382 | 20.9982 | 0.20023 | 0.17813 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 62.46 | 40.54 | 0.778 | 61.682 | 39.762 | 0.4855 | 0.23094 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 58.81 | 38.85 | 0.778 | 58.032 | 38.072 | 0.44209 | 0.24391 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 46.25 | 31.16 | 0.778 | 45.472 | 30.382 | 0.33423 | 0.21321 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 47.56 | 32.4 | 0.778 | 46.782 | 31.622 | 0.33578 | 0.234 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 5 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 18.79 | 12.66 | 0.57567 | 18.2143 | 12.0843 | 0.13393 | 0.080853 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR95
EN712 | 20240212 | L8 | 12 | 5 |  |  | b | 250 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.41 | 10.36 | 0.57567 | 14.8343 | 9.7843 | 0.12975 | 0.074758 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR96; ~50mL lost
EN712 | 20240212 | L8 | 12 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.02 | 7.945 | 0.57567 | 11.4443 | 7.3693 | 0.047504 | 0.022383 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR103
EN712 | 20240212 | L8 | 12 | 5 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.08 | 6.816 | 0.57567 | 10.5043 | 6.2403 | 0.049617 | 0.0094555 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR104
EN712 | 20240212 | L8 | 12 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 68.74 | 45 | 0.778 | 67.962 | 44.222 | 0.52581 | 0.271 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 68.39 | 45.49 | 0.36183 | 68.0282 | 45.1282 | 0.50721 | 0.30593 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 98.65 | 62.98 | 0.778 | 97.872 | 62.202 | 0.79005 | 0.33073 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 98.21 | 62.06 | 0.36183 | 97.8482 | 61.6982 | 0.80068 | 0.31102 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 64.23 | 41.23 | 0.778 | 63.452 | 40.452 | 0.50942 | 0.21946 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 67.42 | 43.37 | 0.778 | 66.642 | 42.592 | 0.53268 | 0.23476 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 34.5 | 22.36 | 0.57567 | 33.9243 | 21.7843 | 0.26613 | 0.12237 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR93
EN712 | 20240212 | L8 | 12 | 10 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 31.09 | 20.07 | 0.57567 | 30.5143 | 19.4943 | 0.24158 | 0.10608 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR94
EN712 | 20240212 | L8 | 12 | 10 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 23.47 | 15.15 | 0.57567 | 22.8943 | 14.5743 | 0.096989 | 0.041225 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR101
EN712 | 20240212 | L8 | 12 | 10 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 28.47 | 17.61 | 0.57567 | 27.8943 | 17.0343 | 0.1266 | 0.034945 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR102
EN712 | 20240212 | L8 | 12 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 91.57 | 58.64 | 0.36183 | 91.2082 | 58.2782 | 0.72936 | 0.32072 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 100.7 | 63.95 | 0.778 | 99.922 | 63.172 | 0.81397 | 0.32429 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 36.43 | 23.12 | 0.57567 | 35.8543 | 22.5443 | 0.29178 | 0.11027 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR91
EN712 | 20240212 | L8 | 12 | 16 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 33.79 | 21.88 | 0.57567 | 33.2143 | 21.3043 | 0.2602 | 0.11844 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR92
EN712 | 20240212 | L8 | 12 | 16 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 38.03 | 22.94 | 0.57567 | 37.4543 | 22.3643 | 0.17623 | 0.036246 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR99
EN712 | 20240212 | L8 | 12 | 16 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 26.95 | 16.77 | 0.57567 | 26.3743 | 16.1943 | 0.11889 | 0.034969 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR100
EN712 | 20240212 | L8 | 12 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 84.7 | 54.12 | 0.36183 | 84.3382 | 53.7582 | 0.67731 | 0.29132 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 77.77 | 50.09 | 0.36183 | 77.4082 | 49.7282 | 0.61308 | 0.28294 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 55.75 | 36.01 | 0.36183 | 55.3882 | 35.6482 | 0.43722 | 0.2051 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 64.6 | 41.11 | 0.778 | 63.822 | 40.332 | 0.52028 | 0.20644 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L8 | 12 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 34.67 | 22.03 | 0.57567 | 34.0943 | 21.4543 | 0.2771 | 0.10552 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR89
EN712 | 20240212 | L8 | 12 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 30.12 | 19.43 | 0.57567 | 29.5443 | 18.8543 | 0.23435 | 0.1019 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR90
EN712 | 20240212 | L8 | 12 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 23.86 | 15.1 | 0.57567 | 23.2843 | 14.5243 | 0.1023 | 0.035686 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR97
EN712 | 20240212 | L8 | 12 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 21.45 | 13.11 | 0.57567 | 20.8743 | 12.5343 | 0.097046 | 0.021607 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR98
EN712 | 20240212 | L4 | 15 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 124.6 | 83.07 | 0.778 | 123.822 | 82.292 | 0.91984 | 0.56293 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 123.8 | 84.55 | 0.36183 | 123.4382 | 84.1882 | 0.86934 | 0.64759 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 155.7 | 99.28 | 0.778 | 154.922 | 98.502 | 1.2496 | 0.52521 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 148.7 | 94.27 | 0.778 | 147.922 | 93.492 | 1.2056 | 0.47901 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 4 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 72.22 | 48.3 | 0.778 | 71.442 | 47.522 | 0.5298 | 0.32647 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 4 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 78.33 | 51.94 | 0.36183 | 77.9682 | 51.5782 | 0.58451 | 0.34485 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 163.5 | 102.8 | 0.778 | 162.722 | 102.022 | 1.3444 | 0.49384 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 177.5 | 112 | 0.778 | 176.722 | 111.222 | 1.4508 | 0.55329 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 49.31 | 33.4 | 0.778 | 48.532 | 32.622 | 0.35239 | 0.23541 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 55.11 | 37.85 | 0.778 | 54.332 | 37.072 | 0.38229 | 0.28569 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 6 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 119.1 | 74.17 | 0.57567 | 118.5243 | 73.5943 | 0.98161 | 0.3264 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR111
EN712 | 20240212 | L4 | 15 | 6 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 97.53 | 59.89 | 0.57567 | 96.9543 | 59.3143 | 0.82234 | 0.23187 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR112
EN712 | 20240212 | L4 | 15 | 6 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 206.8 | 122.1 | 0.57567 | 206.2243 | 121.5243 | 0.98737 | 0.16509 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR119
EN712 | 20240212 | L4 | 15 | 6 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 149.9 | 89.95 | 0.57567 | 149.3243 | 89.3743 | 0.69759 | 0.14845 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR120
EN712 | 20240212 | L4 | 15 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 161.4 | 101.8 | 0.778 | 160.622 | 101.022 | 1.3201 | 0.50018 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 169.9 | 107.5 | 0.778 | 169.122 | 106.722 | 1.3821 | 0.54087 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 11 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 137.8 | 83.86 | 0.57567 | 137.2243 | 83.2843 | 1.1825 | 0.30281 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR109
EN712 | 20240212 | L4 | 15 | 11 |  |  | b | 250 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 123.7 | 75.32 | 0.57567 | 123.1243 | 74.7443 | 1.243 | 0.31924 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR110; filter replaced
EN712 | 20240212 | L4 | 15 | 11 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 206.2 | 120.3 | 0.57567 | 205.6243 | 119.7243 | 1.0014 | 0.13403 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR117
EN712 | 20240212 | L4 | 15 | 11 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 214.1 | 126 | 0.57567 | 213.5243 | 125.4243 | 1.027 | 0.16244 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR118
EN712 | 20240212 | L4 | 15 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 163.3 | 101.8 | 0.778 | 162.522 | 101.022 | 1.3622 | 0.4581 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 174.4 | 109.9 | 0.778 | 173.622 | 109.122 | 1.4286 | 0.5376 | 20240307 | 15-10 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 128.5 | 79.82 | 0.57567 | 127.9243 | 79.2443 | 1.0672 | 0.34607 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR107
EN712 | 20240212 | L4 | 15 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 130.5 | 81.71 | 0.57567 | 129.9243 | 81.1343 | 1.0659 | 0.37608 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR108
EN712 | 20240212 | L4 | 15 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 235.1 | 138.4 | 0.57567 | 234.5243 | 137.8243 | 1.1293 | 0.1801 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR115
EN712 | 20240212 | L4 | 15 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 237.1 | 139 | 0.57567 | 236.5243 | 138.4243 | 1.1457 | 0.16946 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR116
EN712 | 20240212 | L4 | 15 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 168 | 105.2 | 0.778 | 167.222 | 104.422 | 1.3909 | 0.49057 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 183.3 | 113.7 | 0.778 | 182.522 | 112.922 | 1.5416 | 0.49311 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 46.28 | 32.19 | 0.778 | 45.502 | 31.412 | 0.31208 | 0.25392 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 48.65 | 34.28 | 0.36183 | 48.2882 | 33.9182 | 0.31828 | 0.29287 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240212 | L4 | 15 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 156.9 | 93.85 | 0.57567 | 156.3243 | 93.2743 | 1.3822 | 0.28126 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR105
EN712 | 20240212 | L4 | 15 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 135.8 | 83.61 | 0.57567 | 135.2243 | 83.0343 | 1.1441 | 0.33671 | 20240307 | 13-87 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR106
EN712 | 20240212 | L4 | 15 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 214.8 | 126.5 | 0.57567 | 214.2243 | 125.9243 | 1.0312 | 0.16515 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR113
EN712 | 20240212 | L4 | 15 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 216.6 | 127.3 | 0.57567 | 216.0243 | 126.7243 | 1.0391 | 0.16049 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR114
EN712 | 20240213 | L3 | 16 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 357.8 | 224.1 | 0.778 | 357.022 | 223.322 | 2.9613 | 1.0626 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L3 | 16 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 380.2 | 234.2 | 0.778 | 379.422 | 233.422 | 3.2337 | 0.97216 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L3 | 16 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 359.7 | 220.5 | 0.778 | 358.922 | 219.722 | 3.0831 | 0.87592 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L3 | 16 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 334.6 | 206.4 | 0.778 | 333.822 | 205.622 | 2.8395 | 0.86549 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L3 | 16 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 193.5 | 116.7 | 0.778 | 192.722 | 115.922 | 5.1031 | 1.1631 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN712 | 20240213 | L3 | 16 | 7 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 252.4 | 153.6 | 0.36183 | 252.0382 | 153.2382 | 4.3766 | 1.1456 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN712 | 20240213 | L3 | 16 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 170.4 | 100.6 | 0.57567 | 169.8243 | 100.0243 | 4.5905 | 0.76099 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR125 | poor replicate or bad dilution
EN712 | 20240213 | L3 | 16 | 7 |  |  | b | 294 | 5 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 187.7 | 110.5 | 0.57567 | 187.1243 | 109.9243 | 5.0599 | 0.80125 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR126 | poor replicate or bad dilution
EN712 | 20240213 | L3 | 16 | 7 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 297.1 | 172 | 0.57567 | 296.5243 | 171.4243 | 4.3829 | 0.50297 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR131
EN712 | 20240213 | L3 | 16 | 7 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 331.7 | 193.7 | 0.57567 | 331.1243 | 193.1243 | 4.8349 | 0.6695 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR132
EN712 | 20240213 | L3 | 16 | 12 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 252.1 | 152.1 | 0.778 | 251.322 | 151.322 | 4.4298 | 1.0234 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L3 | 16 | 12 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 227.1 | 137.9 | 0.778 | 226.322 | 137.122 | 3.9514 | 0.99008 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L3 | 16 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 408.5 | 244.9 | 0.57567 | 407.9243 | 244.3243 | 3.5865 | 0.7708 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR123 | poor replicate
EN712 | 20240213 | L3 | 16 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 167.3 | 97.04 | 0.57567 | 166.7243 | 96.4643 | 4.6207 | 0.54027 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR124 | poor replicate or bad dilution
EN712 | 20240213 | L3 | 16 | 12 |  |  | a | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 260.9 | 151.5 | 0.57567 | 260.3243 | 150.9243 | 3.8259 | 0.46789 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR129
EN712 | 20240213 | L3 | 16 | 12 |  |  | b | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 255.5 | 148.1 | 0.57567 | 254.9243 | 147.5243 | 3.756 | 0.4411 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR130
EN712 | 20240213 | L3 | 16 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 439.4 | 264.7 | 0.778 | 438.622 | 263.922 | 3.8694 | 0.88604 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L3 | 16 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 429.7 | 251.3 | 0.778 | 428.922 | 250.522 | 3.9514 | 0.56265 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L3 | 16 | 17 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 354.7 | 211.8 | 0.778 | 353.922 | 211.022 | 3.1651 | 0.63721 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | very poor replicate
EN712 | 20240213 | L3 | 16 | 17 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 38.12 | 27.32 | 0.36183 | 37.7582 | 26.9582 | 0.23921 | 0.24654 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | very poor replicate
EN712 | 20240213 | L3 | 16 | 17 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 435.2 | 253.9 | 0.57567 | 434.6243 | 253.3243 | 3.9745 | 0.54329 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR121 | poor replicate
EN712 | 20240213 | L3 | 16 | 17 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 362.7 | 215.9 | 0.57567 | 362.1243 | 215.3243 | 3.2182 | 0.62191 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR122 | poor replicate
EN712 | 20240213 | L3 | 16 | 17 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 236.6 | 138.3 | 0.57567 | 236.0243 | 137.7243 | 3.444 | 0.48141 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR127
EN712 | 20240213 | L3 | 16 | 17 |  |  | b | 552 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 239 | 138.1 | 0.57567 | 238.4243 | 137.5243 | 3.5223 | 0.38322 | 20240307 | 13-88 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR128
EN712 | 20240213 | L2 | 17 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 355.7 | 221.8 | 0.778 | 354.922 | 221.022 | 2.9657 | 1.0167 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN712 | 20240213 | L2 | 17 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 419.4 | 254.1 | 0.778 | 418.622 | 253.322 | 3.6612 | 0.90325 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN712 | 20240213 | L2 | 17 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 318.7 | 195.5 | 0.778 | 317.922 | 194.722 | 2.7287 | 0.77984 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L2 | 17 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 330 | 197.1 | 0.778 | 329.222 | 196.322 | 2.9436 | 0.59382 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L2 | 17 | 4 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 34.76 | 25.15 | 0.778 | 33.982 | 24.372 | 0.21285 | 0.22629 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L2 | 17 | 4 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 39.4 | 27.93 | 0.778 | 38.622 | 27.152 | 0.25405 | 0.23519 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L2 | 17 | 4 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 214.2 | 130.2 | 0.57567 | 213.6243 | 129.6243 | 1.8415 | 0.47025 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR137 | poor replicate
EN712 | 20240213 | L2 | 17 | 4 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 364.4 | 216.6 | 0.57567 | 363.8243 | 216.0243 | 3.2401 | 0.61247 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR138 | poor replicate
EN712 | 20240213 | L2 | 17 | 4 |  |  | a | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 213.9 | 126.3 | 0.57567 | 213.3243 | 125.7243 | 3.0635 | 0.51333 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR143
EN712 | 20240213 | L2 | 17 | 4 |  |  | b | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 223.3 | 130.3 | 0.57567 | 222.7243 | 129.7243 | 3.2524 | 0.43828 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR144
EN712 | 20240213 | L2 | 17 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 255.1 | 156.5 | 0.778 | 254.322 | 155.722 | 2.1839 | 0.62198 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L2 | 17 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 254 | 155.4 | 0.36183 | 253.6382 | 155.0382 | 2.1839 | 0.60966 | 20240307 | 15-18 | 20250319 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L2 | 17 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 317.7 | 189.1 | 0.57567 | 317.1243 | 188.5243 | 2.8192 | 0.54295 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR135
EN712 | 20240213 | L2 | 17 | 8 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 362.8 | 214.9 | 0.57567 | 362.2243 | 214.3243 | 3.2312 | 0.57799 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR136
EN712 | 20240213 | L2 | 17 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 174.1 | 101.7 | 0.57567 | 173.5243 | 101.1243 | 2.5366 | 0.34566 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR141
EN712 | 20240213 | L2 | 17 | 8 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 167.7 | 96.43 | 0.57567 | 167.1243 | 95.8543 | 2.497 | 0.23504 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR142
EN712 | 20240213 | L2 | 17 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 195.3 | 117.3 | 0.778 | 194.522 | 116.522 | 1.7276 | 0.37193 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L2 | 17 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 180.8 | 109.2 | 0.778 | 180.022 | 108.422 | 1.5859 | 0.36773 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L2 | 17 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 38.78 | 25.54 | 0.778 | 38.002 | 24.762 | 0.29325 | 0.15292 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L2 | 17 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 38.23 | 26.79 | 0.778 | 37.452 | 26.012 | 0.25338 | 0.21531 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | L2 | 17 | 13 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 201.2 | 119.4 | 0.57567 | 200.6243 | 118.8243 | 1.7932 | 0.32588 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR133
EN712 | 20240213 | L2 | 17 | 13 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 207.7 | 122.8 | 0.57567 | 207.1243 | 122.2243 | 1.8612 | 0.31855 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR134
EN712 | 20240213 | L2 | 17 | 13 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 278.4 | 162.5 | 0.57567 | 277.8243 | 161.9243 | 1.3535 | 0.18484 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR139
EN712 | 20240213 | L2 | 17 | 13 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 289.3 | 167.1 | 0.57567 | 288.7243 | 166.5243 | 1.4219 | 0.15441 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR140
EN712 | 20240213 | MVCO | 20 | 1 |  |  | a | 290 | 0 | 15 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 218 | 135.3 | 0.778 | 217.222 | 134.522 | 16.4854 | 5.3294 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN712 | 20240213 | MVCO | 20 | 1 |  |  | b | 290 | 0 | 15 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 206 | 126.8 | 0.778 | 205.222 | 126.022 | 15.7877 | 4.6487 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN712 | 20240213 | MVCO | 20 | 1 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 142 | 101.2 | 0.778 | 141.222 | 100.422 | 0.90367 | 0.90577 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl; depth = 15
EN712 | 20240213 | MVCO | 20 | 1 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 113.7 | 82.74 | 0.778 | 112.922 | 81.962 | 0.68573 | 0.7911 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl; depth = 15
EN712 | 20240213 | MVCO | 20 | 4 |  |  | a | 290 | 0 | 15 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 203.7 | 125.4 | 0.778 | 202.922 | 124.622 | 15.6083 | 4.6011 | 20240307 | 15-11 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN712 | 20240213 | MVCO | 20 | 4 |  |  | b | 290 | 0 | 15 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 216.7 | 133.3 | 0.778 | 215.922 | 132.522 | 16.6249 | 4.8656 | 20240307 | 15-12 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN712 | 20240213 | MVCO | 20 | 4 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 65.38 | 45.6 | 0.778 | 64.602 | 44.822 | 0.4381 | 0.36952 | 20240307 | 15-12 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN712 | 20240213 | MVCO | 20 | 4 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 96.88 | 68.82 | 0.778 | 96.102 | 68.042 | 0.6215 | 0.60451 | 20240307 | 15-12 | 20250305 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240503 | L1 | 1 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 153.5 | 91.97 | 0.75067 | 152.7493 | 91.2193 | 1.3662 | 0.261 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240503 | L1 | 1 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 161.8 | 98.22 | 0.75067 | 161.0493 | 97.4693 | 1.4118 | 0.32698 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240503 | L1 | 1 | 2 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 38.43 | 24.07 | 0.75067 | 37.6793 | 23.3193 | 0.31886 | 0.097133 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240503 | L1 | 1 | 2 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 40.98 | 26.48 | 0.75067 | 40.2293 | 25.7293 | 0.32197 | 0.13702 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240503 | L1 | 1 | 2 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 88.44 | 53.07 | 0.57567 | 87.8643 | 52.4943 | 0.77538 | 0.16079 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR5
EN715 | 20240503 | L1 | 1 | 2 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 81.72 | 49.6 | 0.57567 | 81.1443 | 49.0243 | 0.70414 | 0.17016 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR6
EN715 | 20240503 | L1 | 1 | 2 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 78.49 | 47.1 | 0.57567 | 77.9143 | 46.5243 | 0.36592 | 0.075285 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR11
EN715 | 20240503 | L1 | 1 | 2 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 71.67 | 43.29 | 0.57567 | 71.0943 | 42.7143 | 0.33083 | 0.074242 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR12
EN715 | 20240503 | L1 | 1 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 166 | 99.53 | 0.75067 | 165.2493 | 98.7793 | 1.4759 | 0.28617 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240503 | L1 | 1 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 184.2 | 111.2 | 0.75067 | 183.4493 | 110.4493 | 1.6209 | 0.34936 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240503 | L1 | 1 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 86.89 | 52.88 | 0.75067 | 86.1393 | 52.1293 | 0.75518 | 0.17475 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240503 | L1 | 1 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 71.86 | 44.3 | 0.75067 | 71.1093 | 43.5493 | 0.61196 | 0.16491 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240503 | L1 | 1 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 81.34 | 50.06 | 0.57567 | 80.7643 | 49.4843 | 0.68572 | 0.19678 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR3
EN715 | 20240503 | L1 | 1 | 7 |  |  | b | 264 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 79.15 | 48.84 | 0.57567 | 78.5743 | 48.2643 | 0.73745 | 0.21784 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR4; spilled ~30mL
EN715 | 20240503 | L1 | 1 | 7 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 84.61 | 51.06 | 0.57567 | 84.0343 | 50.4843 | 0.39181 | 0.087818 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR9
EN715 | 20240503 | L1 | 1 | 7 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 70.47 | 41.74 | 0.57567 | 69.8943 | 41.1643 | 0.33552 | 0.055563 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR10
EN715 | 20240503 | L1 | 1 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 201.4 | 121.7 | 0.75067 | 200.6493 | 120.9493 | 1.7697 | 0.3879 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240503 | L1 | 1 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 199.5 | 120.5 | 0.75067 | 198.7493 | 119.7493 | 1.7542 | 0.38203 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240503 | L1 | 1 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 87.2 | 53.35 | 0.75067 | 86.4493 | 52.5993 | 0.75162 | 0.18669 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240503 | L1 | 1 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 94.02 | 58.06 | 0.75067 | 93.2693 | 57.3093 | 0.79848 | 0.22386 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240503 | L1 | 1 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 72.95 | 44.72 | 0.57567 | 72.3743 | 44.1443 | 0.61886 | 0.16841 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR1
EN715 | 20240503 | L1 | 1 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 74.38 | 46.21 | 0.57567 | 73.8043 | 45.6343 | 0.61755 | 0.19629 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR2
EN715 | 20240503 | L1 | 1 | 14 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 88.65 | 52.58 | 0.57567 | 88.0743 | 52.0043 | 0.41972 | 0.072566 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR7
EN715 | 20240503 | L1 | 1 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 88.61 | 50.73 | 0.57567 | 88.0343 | 50.1543 | 0.44238 | 0.034115 | 20240307 | 13-89 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR8
EN715 | 20240504 | L2 | 2 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 173.9 | 110.6 | 0.75067 | 173.1493 | 109.8493 | 1.4055 | 0.55404 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L2 | 2 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 173.5 | 111.8 | 0.75067 | 172.7493 | 111.0493 | 1.37 | 0.61097 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L2 | 2 | 1 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 40.27 | 30.19 | 0.75067 | 39.5193 | 29.4393 | 0.22382 | 0.30134 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN715 | 20240504 | L2 | 2 | 1 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 32.26 | 24.89 | 0.75067 | 31.5093 | 24.1393 | 0.16365 | 0.26697 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN715 | 20240504 | L2 | 2 | 4 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240504 | L2 | 2 | 4 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240504 | L2 | 2 | 4 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 109.9 | 71.14 | 0.57567 | 109.3243 | 70.5643 | 0.8497 | 0.40874 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR17
EN715 | 20240504 | L2 | 2 | 4 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 92.06 | 59.84 | 0.57567 | 91.4843 | 59.2643 | 0.70633 | 0.35059 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR18
EN715 | 20240504 | L2 | 2 | 4 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 151.8 | 94.6 | 0.57567 | 151.2243 | 94.0243 | 0.6668 | 0.22487 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR23
EN715 | 20240504 | L2 | 2 | 4 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 123.9 | 77.76 | 0.57567 | 123.3243 | 77.1843 | 0.53787 | 0.1941 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR24
EN715 | 20240504 | L2 | 2 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 110.2 | 68.63 | 0.75067 | 109.4493 | 67.8793 | 0.92304 | 0.28785 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L2 | 2 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 115.2 | 69.64 | 0.75067 | 114.4493 | 68.8893 | 1.0116 | 0.21727 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L2 | 2 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 43.84 | 26.9 | 0.75067 | 43.0893 | 26.1493 | 0.37614 | 0.090329 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L2 | 2 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 46.67 | 28.95 | 0.75067 | 45.9193 | 28.1993 | 0.39346 | 0.10958 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L2 | 2 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 57.69 | 36.3 | 0.57567 | 57.1143 | 35.7243 | 0.46891 | 0.16819 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR15
EN715 | 20240504 | L2 | 2 | 8 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 61.73 | 37.78 | 0.57567 | 61.1543 | 37.2043 | 0.52325 | 0.13799 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR16
EN715 | 20240504 | L2 | 2 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 75.44 | 46.99 | 0.57567 | 74.8643 | 46.4143 | 0.33225 | 0.10871 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR21
EN715 | 20240504 | L2 | 2 | 8 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 74.36 | 45.32 | 0.57567 | 73.7843 | 44.7443 | 0.33914 | 0.085955 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR22
EN715 | 20240504 | L2 | 2 | 15 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240504 | L2 | 2 | 15 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240504 | L2 | 2 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 115.6 | 68.1 | 0.75067 | 114.8493 | 67.3493 | 1.0547 | 0.14672 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L2 | 2 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 119.6 | 70.22 | 0.75067 | 118.8493 | 69.4693 | 1.0965 | 0.14279 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L2 | 2 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 29.47 | 18.21 | 0.75067 | 28.7193 | 17.4593 | 0.25002 | 0.061431 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L2 | 2 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 35.97 | 22.06 | 0.75067 | 35.2193 | 21.3093 | 0.30887 | 0.071269 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L2 | 2 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 71.07 | 42.71 | 0.57567 | 70.4943 | 42.1343 | 0.62171 | 0.12971 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR13
EN715 | 20240504 | L2 | 2 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 72.07 | 43.01 | 0.57567 | 71.4943 | 42.4343 | 0.63706 | 0.11971 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR14
EN715 | 20240504 | L2 | 2 | 18 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 98.43 | 59.41 | 0.57567 | 97.8543 | 58.8343 | 0.45404 | 0.10289 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR19
EN715 | 20240504 | L2 | 2 | 18 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 104.7 | 61.94 | 0.57567 | 104.1243 | 61.3643 | 0.49937 | 0.083626 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR20
EN715 | 20240504 | L4 | 5 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 51.78 | 40.08 | 0.75067 | 51.0293 | 39.3293 | 0.25979 | 0.4418 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN715 | 20240504 | L4 | 5 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 48.94 | 38.07 | 0.75067 | 48.1893 | 37.3193 | 0.24136 | 0.42437 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN715 | 20240504 | L4 | 5 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 76.57 | 57.77 | 0.75067 | 75.8193 | 57.0193 | 0.41745 | 0.59971 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN715 | 20240504 | L4 | 5 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 77.59 | 58.34 | 0.75067 | 76.8393 | 57.5893 | 0.42744 | 0.59989 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN715 | 20240504 | L4 | 5 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 96.32 | 70.94 | 0.75067 | 95.5693 | 70.1893 | 0.56355 | 0.68855 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN715 | 20240504 | L4 | 5 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 93.97 | 70 | 0.75067 | 93.2193 | 69.2493 | 0.53224 | 0.70309 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN715 | 20240504 | L4 | 5 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 70.03 | 52.62 | 0.75067 | 69.2793 | 51.8693 | 0.38658 | 0.53871 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN715 | 20240504 | L4 | 5 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 77.77 | 57.59 | 0.75067 | 77.0193 | 56.8393 | 0.44809 | 0.56586 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN715 | 20240504 | L4 | 5 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 41.37 | 29.71 | 0.68567 | 40.6843 | 29.0243 | 0.25561 | 0.26201 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR29 | Phaeo>Chl
EN715 | 20240504 | L4 | 5 | 7 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 39.84 | 29.49 | 0.68567 | 39.1543 | 28.8043 | 0.22689 | 0.2868 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR30 | Phaeo>Chl
EN715 | 20240504 | L4 | 5 | 7 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.42 | 9.926 | 0.68567 | 13.7343 | 9.2403 | 0.052388 | 0.035242 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR35
EN715 | 20240504 | L4 | 5 | 7 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.6 | 8.424 | 0.68567 | 10.9143 | 7.7383 | 0.037024 | 0.036362 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR36
EN715 | 20240504 | L4 | 5 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 173.4 | 121 | 0.75067 | 172.6493 | 120.2493 | 1.1635 | 0.98159 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L4 | 5 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 179.2 | 126.5 | 0.75067 | 178.4493 | 125.7493 | 1.1702 | 1.073 | 20250529 | 15-29 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L4 | 5 | 13 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 215.1 | 131.7 | 0.75067 | 214.3493 | 130.9493 | 3.7037 | 0.96826 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L4 | 5 | 13 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 224.2 | 137.5 | 0.75067 | 223.4493 | 136.7493 | 3.8503 | 1.0286 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L4 | 5 | 13 |  |  | a | 290 | 10 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 211.3 | 127.8 | 0.75067 | 210.5493 | 127.0493 | 3.7082 | 0.82467 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN715 | 20240504 | L4 | 5 | 13 |  |  | b | 290 | 10 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 185.3 | 113.7 | 0.75067 | 184.5493 | 112.9493 | 3.1797 | 0.85008 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN715 | 20240504 | L4 | 5 | 13 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 183.1 | 114.3 | 0.68567 | 182.4143 | 113.6143 | 1.5082 | 0.51795 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR27 | poor replicate
EN715 | 20240504 | L4 | 5 | 13 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 242.1 | 146.4 | 0.68567 | 241.4143 | 145.7143 | 2.0908 | 0.49901 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR28 | poor replicate
EN715 | 20240504 | L4 | 5 | 13 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 20.62 | 13.61 | 0.68567 | 19.9343 | 12.9243 | 0.081866 | 0.040923 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR33
EN715 | 20240504 | L4 | 5 | 13 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 28.2 | 18.8 | 0.68567 | 27.5143 | 18.1143 | 0.10978 | 0.062319 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR34
EN715 | 20240504 | L4 | 5 | 18 |  |  | a | 290 | 0 | 15 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 119.7 | 72.58 | 0.75067 | 118.9493 | 71.8293 | 3.1388 | 0.70523 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L4 | 5 | 18 |  |  | b | 290 | 0 | 15 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 139.9 | 85.23 | 0.75067 | 139.1493 | 84.4793 | 3.6418 | 0.87928 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L4 | 5 | 18 |  |  | a | 290 | 10 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 195.6 | 118.9 | 0.75067 | 194.8493 | 118.1493 | 3.4062 | 0.80912 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L4 | 5 | 18 |  |  | b | 290 | 10 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 200.1 | 122.2 | 0.75067 | 199.3493 | 121.4493 | 3.4595 | 0.87357 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L4 | 5 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 72.23 | 45.51 | 0.68567 | 71.6543 | 44.9343 | 0.58576 | 0.2156 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR25 | poor replicate or mixed up filter
EN715 | 20240504 | L4 | 5 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 202.4 | 123.7 | 0.68567 | 201.8243 | 123.1243 | 1.7253 | 0.47052 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR26 | poor replicate or mixed up filter
EN715 | 20240504 | L4 | 5 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 22.3 | 14.66 | 0.68567 | 21.7243 | 14.0843 | 0.089224 | 0.044586 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR31
EN715 | 20240504 | L4 | 5 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 27.02 | 18.13 | 0.68567 | 26.4443 | 17.5543 | 0.10345 | 0.062727 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR32
EN715 | 20240504 | L5 | 6 | 2 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240504 | L5 | 6 | 2 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240504 | L5 | 6 | 6 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240504 | L5 | 6 | 6 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240504 | L5 | 6 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 82.43 | 55.22 | 0.75067 | 81.6793 | 54.4693 | 0.60419 | 0.36749 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L5 | 6 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 85.08 | 57.17 | 0.75067 | 84.3293 | 56.4193 | 0.61973 | 0.38673 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L5 | 6 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 159.4 | 98.05 | 0.75067 | 158.6493 | 97.2993 | 1.3622 | 0.37346 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L5 | 6 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 167 | 103.1 | 0.75067 | 166.2493 | 102.3493 | 1.4189 | 0.40692 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L5 | 6 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 130.5 | 80.69 | 0.75067 | 129.7493 | 79.9393 | 1.106 | 0.32002 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L5 | 6 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 144.9 | 89.21 | 0.75067 | 144.1493 | 88.4593 | 1.2366 | 0.34144 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L5 | 6 | 12 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 57.88 | 35.39 | 0.57567 | 57.3043 | 34.8143 | 0.49135 | 0.12741 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR39
EN715 | 20240504 | L5 | 6 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 62.94 | 38.5 | 0.57567 | 62.3643 | 37.9243 | 0.53578 | 0.14056 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR40
EN715 | 20240504 | L5 | 6 | 12 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 34.62 | 20.59 | 0.57567 | 34.0443 | 20.0143 | 0.16385 | 0.026299 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR43
EN715 | 20240504 | L5 | 6 | 12 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 32.51 | 19.64 | 0.57567 | 31.9343 | 19.0643 | 0.1503 | 0.030821 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR44
EN715 | 20240504 | L5 | 6 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 204.3 | 122.4 | 0.75067 | 203.5493 | 121.6493 | 1.8186 | 0.35153 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L5 | 6 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 173.7 | 104.7 | 0.75067 | 172.9493 | 103.9493 | 1.5321 | 0.32222 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L5 | 6 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 120 | 71.95 | 0.75067 | 119.2493 | 71.1993 | 1.0669 | 0.20319 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L5 | 6 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 117.7 | 71.32 | 0.75067 | 116.9493 | 70.5693 | 1.0298 | 0.22903 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L5 | 6 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 99.6 | 60.69 | 0.75067 | 98.8493 | 59.9393 | 0.86398 | 0.20527 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L5 | 6 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 92.27 | 56.03 | 0.75067 | 91.5193 | 55.2793 | 0.80469 | 0.18143 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240504 | L5 | 6 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 39.21 | 23.14 | 0.57567 | 38.6343 | 22.5643 | 0.35229 | 0.050122 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR37
EN715 | 20240504 | L5 | 6 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 35.41 | 21.61 | 0.57567 | 34.8343 | 21.0343 | 0.30252 | 0.0726 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR38; leaked
EN715 | 20240504 | L5 | 6 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 26.49 | 16.04 | 0.57567 | 25.9143 | 15.4643 | 0.12204 | 0.02488 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR41; leaked
EN715 | 20240504 | L5 | 6 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 26.1 | 15.78 | 0.57567 | 25.5243 | 15.2043 | 0.12009 | 0.023842 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR42
EN715 | 20240505 | L10 | 7 | 2 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240505 | L10 | 7 | 2 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240505 | L10 | 7 | 4 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240505 | L10 | 7 | 4 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240505 | L10 | 7 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 52.49 | 37.52 | 0.75067 | 51.7393 | 36.7693 | 0.3324 | 0.32352 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L10 | 7 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 54.52 | 39.37 | 0.75067 | 53.7693 | 38.6193 | 0.3364 | 0.35253 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN715 | 20240505 | L10 | 7 | 7 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 43.13 | 30.74 | 0.75067 | 42.3793 | 29.9893 | 0.27511 | 0.25986 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L10 | 7 | 7 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 37.99 | 27.54 | 0.75067 | 37.2393 | 26.7893 | 0.23204 | 0.24585 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN715 | 20240505 | L10 | 7 | 7 |  |  | a | 193 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 17.41 | 12.92 | 0.57567 | 16.8343 | 12.3443 | 0.14943 | 0.18478 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR49; ~100mL leaked
EN715 | 20240505 | L10 | 7 | 7 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 21.11 | 15.05 | 0.57567 | 20.5343 | 14.4743 | 0.13285 | 0.12529 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR50
EN715 | 20240505 | L10 | 7 | 7 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.54 | 11 | 0.57567 | 14.9643 | 10.4243 | 0.052924 | 0.045934 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR55
EN715 | 20240505 | L10 | 7 | 7 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.35 | 10.54 | 0.57567 | 14.7743 | 9.9643 | 0.056072 | 0.038424 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR56
EN715 | 20240505 | L10 | 7 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 184.2 | 121.1 | 0.75067 | 183.4493 | 120.3493 | 1.4011 | 0.74579 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L10 | 7 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 181.4 | 122.6 | 0.75067 | 180.6493 | 121.8493 | 1.3056 | 0.86803 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L10 | 7 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 85.32 | 58.79 | 0.75067 | 84.5693 | 58.0393 | 0.58909 | 0.44627 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L10 | 7 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 111.9 | 77.53 | 0.75067 | 111.1493 | 76.7793 | 0.76317 | 0.60649 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L10 | 7 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 115.6 | 73.52 | 0.57567 | 115.0243 | 72.9443 | 0.92248 | 0.3784 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR47
EN715 | 20240505 | L10 | 7 | 12 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 120.7 | 78.37 | 0.57567 | 120.1243 | 77.7943 | 0.92481 | 0.45785 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR48
EN715 | 20240505 | L10 | 7 | 12 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 104 | 64.37 | 0.57567 | 103.4243 | 63.7943 | 0.46282 | 0.14327 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR53
EN715 | 20240505 | L10 | 7 | 12 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 124.6 | 76.45 | 0.57567 | 124.0243 | 75.8743 | 0.56232 | 0.15853 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR54
EN715 | 20240505 | L10 | 7 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 102.3 | 68.51 | 0.75067 | 101.5493 | 67.7593 | 0.75029 | 0.45846 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L10 | 7 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 111.1 | 73.99 | 0.75067 | 110.3493 | 73.2393 | 0.82401 | 0.4825 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L10 | 7 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 76.87 | 50.77 | 0.75067 | 76.1193 | 50.0193 | 0.57954 | 0.31275 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L10 | 7 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 83.54 | 55.75 | 0.75067 | 82.7893 | 54.9993 | 0.61706 | 0.36406 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L10 | 7 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 46.18 | 31.28 | 0.75067 | 45.4293 | 30.5293 | 0.33085 | 0.21376 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L10 | 7 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 44.98 | 29.98 | 0.75067 | 44.2293 | 29.2293 | 0.33307 | 0.18835 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L10 | 7 | 20 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 41.9 | 28.28 | 0.57567 | 41.3243 | 27.7043 | 0.29858 | 0.1955 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR45
EN715 | 20240505 | L10 | 7 | 20 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 42.49 | 29.38 | 0.57567 | 41.9143 | 28.8043 | 0.2874 | 0.2263 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR46
EN715 | 20240505 | L10 | 7 | 20 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 40.04 | 25.61 | 0.57567 | 39.4643 | 25.0343 | 0.16852 | 0.069321 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR51
EN715 | 20240505 | L10 | 7 | 20 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 46.76 | 30.07 | 0.57567 | 46.1843 | 29.4943 | 0.19421 | 0.084991 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR52
EN715 | 20240505 | L11 | 10 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 16.62 | 11.73 | 0.75067 | 15.8693 | 10.9793 | 0.10858 | 0.087279 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L11 | 10 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 14.47 | 10.6 | 0.5535 | 13.9165 | 10.0465 | 0.085932 | 0.093286 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phae>Chl
EN715 | 20240505 | L11 | 10 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 4.531 | 3.943 | 0.5535 | 3.9775 | 3.3895 | 0.013056 | 0.047409 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phae>Chl;poor replicate
EN715 | 20240505 | L11 | 10 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 13.52 | 9.583 | 0.75067 | 12.7693 | 8.8323 | 0.087419 | 0.070139 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN715 | 20240505 | L11 | 10 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.75 | 9.981 | 0.57567 | 14.1743 | 9.4053 | 0.10455 | 0.063187 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR61
EN715 | 20240505 | L11 | 10 | 6 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.3 | 8.015 | 0.57567 | 10.7243 | 7.4393 | 0.072014 | 0.060658 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR62
EN715 | 20240505 | L11 | 10 | 6 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13.57 | 9.239 | 0.57567 | 12.9943 | 8.6633 | 0.050488 | 0.03167 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR67
EN715 | 20240505 | L11 | 10 | 6 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.45 | 10.56 | 0.57567 | 15.8743 | 9.9843 | 0.068662 | 0.026024 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR68
EN715 | 20240505 | L11 | 10 | 10 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 171.3 | 105.3 | 0.75067 | 170.5493 | 104.5493 | 2.931 | 0.79908 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L11 | 10 | 10 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 173.1 | 106.1 | 0.75067 | 172.3493 | 105.3493 | 2.9754 | 0.78321 | 20250529 | 15-30 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L11 | 10 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 46.05 | 31.43 | 0.5535 | 45.4965 | 30.8765 | 0.32463 | 0.22617 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L11 | 10 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 48.72 | 32.72 | 0.75067 | 47.9693 | 31.9693 | 0.35527 | 0.21502 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L11 | 10 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 375 | 223.4 | 0.57567 | 374.4243 | 222.8243 | 3.3234 | 0.65044 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR59
EN715 | 20240505 | L11 | 10 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 315.9 | 190.8 | 0.57567 | 315.3243 | 190.2243 | 2.7331 | 0.64778 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR60
EN715 | 20240505 | L11 | 10 | 10 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 170.5 | 102 | 0.57567 | 169.9243 | 101.4243 | 2.3999 | 0.49085 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR65 | poor replicate or bad dilution
EN715 | 20240505 | L11 | 10 | 10 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 104.9 | 60.81 | 0.57567 | 104.3243 | 60.2343 | 1.5447 | 0.17207 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR66 | poor replicate or bad dilution
EN715 | 20240505 | L11 | 10 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 145.5 | 93.73 | 0.75067 | 144.7493 | 92.9793 | 1.1495 | 0.50912 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L11 | 10 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 140.9 | 90.9 | 0.75067 | 140.1493 | 90.1493 | 1.1102 | 0.49793 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L11 | 10 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 117.7 | 72.82 | 0.5535 | 117.1465 | 72.2665 | 0.99654 | 0.29261 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L11 | 10 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 118.2 | 73.05 | 0.75067 | 117.4493 | 72.2993 | 1.0025 | 0.2872 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L11 | 10 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 102.9 | 62.44 | 0.5535 | 102.3465 | 61.8865 | 0.8984 | 0.20559 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L11 | 10 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 91.21 | 56.04 | 0.75067 | 90.4593 | 55.2893 | 0.78093 | 0.20536 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L11 | 10 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 57.65 | 36.05 | 0.75067 | 56.8993 | 35.2993 | 0.47962 | 0.15008 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L11 | 10 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 50.51 | 32.4 | 0.75067 | 49.7593 | 31.6493 | 0.40212 | 0.16246 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240505 | L11 | 10 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 48.81 | 29.63 | 0.57567 | 48.2343 | 29.0543 | 0.42047 | 0.097687 | 20240307 | 13-90 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR57
EN715 | 20240505 | L11 | 10 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 45.11 | 26.33 | 0.57567 | 44.5343 | 25.7543 | 0.4117 | 0.047604 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR58
EN715 | 20240505 | L11 | 10 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 36.76 | 21.85 | 0.57567 | 36.1843 | 21.2743 | 0.17413 | 0.027993 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR63
EN715 | 20240505 | L11 | 10 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 34.6 | 20.73 | 0.57567 | 34.0243 | 20.1543 | 0.16139 | 0.029391 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR64
EN715 | 20240506 | L6 | 13 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 120.5 | 89.55 | 0.75067 | 119.7493 | 88.7993 | 0.68723 | 0.89685 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl
EN715 | 20240506 | L6 | 13 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 134.8 | 99.43 | 0.75067 | 134.0493 | 98.6793 | 0.78537 | 0.97495 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl
EN715 | 20240506 | L6 | 13 | 5 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 259 | 172.4 | 0.5535 | 258.4465 | 171.8465 | 3.8458 | 2.2853 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L6 | 13 | 5 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 278 | 184.6 | 0.75067 | 277.2493 | 183.8493 | 4.1478 | 2.4115 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L6 | 13 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 77.52 | 60.24 | 0.75067 | 76.7693 | 59.4893 | 0.38369 | 0.67753 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl
EN715 | 20240506 | L6 | 13 | 5 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 45.53 | 33.25 | 0.75067 | 44.7793 | 32.4993 | 0.27267 | 0.30708 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl
EN715 | 20240506 | L6 | 13 | 5 |  |  | a | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 180.8 | 120.3 | 0.57567 | 180.2243 | 119.7243 | 3.9789 | 2.4266 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR73 | poor replicate or bad dilution
EN715 | 20240506 | L6 | 13 | 5 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 411.3 | 281.2 | 0.57567 | 410.7243 | 280.6243 | 2.8424 | 2.1452 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR74 | poor replicate
EN715 | 20240506 | L6 | 13 | 5 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 313.3 | 199 | 0.57567 | 312.7243 | 198.4243 | 4.0046 | 1.6509 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR79
EN715 | 20240506 | L6 | 13 | 5 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 319.7 | 202.6 | 0.57567 | 319.1243 | 202.0243 | 4.1027 | 1.6554 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR80
EN715 | 20240506 | L6 | 13 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 97.61 | 60.92 | 0.75067 | 96.8593 | 60.1693 | 0.81468 | 0.25867 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L6 | 13 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 96.09 | 61.65 | 0.75067 | 95.3393 | 60.8993 | 0.76472 | 0.32165 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L6 | 13 | 12 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 62.46 | 39.49 | 0.75067 | 61.7093 | 38.7393 | 0.51004 | 0.18103 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L6 | 13 | 12 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 52.43 | 33.69 | 0.75067 | 51.6793 | 32.9393 | 0.41611 | 0.17149 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L6 | 13 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 54.71 | 35.26 | 0.57567 | 54.1343 | 34.6843 | 0.42638 | 0.19217 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR71
EN715 | 20240506 | L6 | 13 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 57.05 | 36.68 | 0.57567 | 56.4743 | 36.1043 | 0.44655 | 0.19733 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR72
EN715 | 20240506 | L6 | 13 | 12 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 69.01 | 40.86 | 0.57567 | 68.4343 | 40.2843 | 0.32815 | 0.053878 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR77
EN715 | 20240506 | L6 | 13 | 12 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 65.47 | 39.54 | 0.57567 | 64.8943 | 38.9643 | 0.30227 | 0.067239 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR78
EN715 | 20240506 | L6 | 13 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 71.28 | 44.4 | 0.75067 | 70.5293 | 43.6493 | 0.59686 | 0.1818 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L6 | 13 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 72.32 | 45.16 | 0.75067 | 71.5693 | 44.4093 | 0.60308 | 0.18914 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L6 | 13 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 67.95 | 41.27 | 0.75067 | 67.1993 | 40.5193 | 0.59242 | 0.1304 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L6 | 13 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 74.98 | 45.86 | 0.75067 | 74.2293 | 45.1093 | 0.6466 | 0.1581 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L6 | 13 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 82.43 | 51.09 | 0.75067 | 81.6793 | 50.3393 | 0.69589 | 0.20211 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L6 | 13 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 68.35 | 42.06 | 0.75067 | 67.5993 | 41.3093 | 0.58376 | 0.15315 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L6 | 13 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 42.07 | 26.69 | 0.75067 | 41.3193 | 25.9393 | 0.34151 | 0.12122 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L6 | 13 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 30.38 | 19.21 | 0.5535 | 29.8265 | 18.6565 | 0.24802 | 0.084786 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L6 | 13 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 32.24 | 19.36 | 0.57567 | 31.6643 | 18.7843 | 0.28236 | 0.052642 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR69
EN715 | 20240506 | L6 | 13 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 31.43 | 19.18 | 0.57567 | 30.8543 | 18.6043 | 0.26855 | 0.063242 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR70
EN715 | 20240506 | L6 | 13 | 19 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 39.24 | 23.23 | 0.57567 | 38.6643 | 22.6543 | 0.18697 | 0.028257 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR75
EN715 | 20240506 | L6 | 13 | 19 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 36.69 | 21.81 | 0.57567 | 36.1143 | 21.2343 | 0.17315 | 0.027862 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR76
EN715 | 20240506 | L8 | 14 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 202.6 | 157.2 | 0.5535 | 202.0465 | 156.6465 | 1.0081 | 1.7863 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 189.3 | 148.8 | 0.5535 | 188.7465 | 148.2465 | 0.89928 | 1.7453 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phae>Chl
EN715 | 20240506 | L8 | 14 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 307.1 | 213.6 | 0.75067 | 306.3493 | 212.8493 | 2.0761 | 1.7209 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 4 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 289.8 | 198 | 0.75067 | 289.0493 | 197.2493 | 2.0384 | 1.4803 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 158.8 | 105.7 | 0.75067 | 158.0493 | 104.9493 | 1.1791 | 0.69311 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 161.5 | 109.3 | 0.75067 | 160.7493 | 108.5493 | 1.1591 | 0.77732 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 60.82 | 44.68 | 0.5535 | 60.2665 | 44.1265 | 0.35838 | 0.42878 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phae>Chl
EN715 | 20240506 | L8 | 14 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 46.06 | 34.27 | 0.5535 | 45.5065 | 33.7165 | 0.26179 | 0.33967 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phae>Chl
EN715 | 20240506 | L8 | 14 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 129 | 87.42 | 0.57567 | 128.4243 | 86.8443 | 0.91152 | 0.63725 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR85
EN715 | 20240506 | L8 | 14 | 6 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 127.2 | 87.39 | 0.57567 | 126.6243 | 86.8143 | 0.86975 | 0.67322 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR86
EN715 | 20240506 | L8 | 14 | 6 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 111.5 | 71.55 | 0.57567 | 110.9243 | 70.9743 | 0.46656 | 0.20774 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR91
EN715 | 20240506 | L8 | 14 | 6 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 93.17 | 60.01 | 0.57567 | 92.5943 | 59.4343 | 0.38726 | 0.1774 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR92
EN715 | 20240506 | L8 | 14 | 12 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 113.3 | 72.03 | 0.75067 | 112.5493 | 71.2793 | 0.91638 | 0.35516 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 12 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 109.2 | 69.54 | 0.5535 | 108.6465 | 68.9865 | 0.88063 | 0.35001 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 71.92 | 45.89 | 0.75067 | 71.1693 | 45.1393 | 0.57798 | 0.22725 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 75.33 | 46.75 | 0.75067 | 74.5793 | 45.9993 | 0.63461 | 0.18597 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 51.09 | 33.03 | 0.75067 | 50.3393 | 32.2793 | 0.40101 | 0.17481 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 56.32 | 36.35 | 0.5535 | 55.7665 | 35.7965 | 0.44342 | 0.19514 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 32.49 | 20.73 | 0.57567 | 31.9143 | 20.1543 | 0.2578 | 0.10163 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR83
EN715 | 20240506 | L8 | 14 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 33 | 20.96 | 0.57567 | 32.4243 | 20.3843 | 0.26394 | 0.09959 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR84
EN715 | 20240506 | L8 | 14 | 14 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 33.84 | 20.71 | 0.57567 | 33.2643 | 20.1343 | 0.15306 | 0.037881 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR89
EN715 | 20240506 | L8 | 14 | 14 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 37.08 | 22.75 | 0.57567 | 36.5043 | 22.1743 | 0.16705 | 0.043238 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR90
EN715 | 20240506 | L8 | 14 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 56.48 | 33.85 | 0.75067 | 55.7293 | 33.0993 | 0.50249 | 0.087965 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 49.95 | 30.5 | 0.75067 | 49.1993 | 29.7493 | 0.43188 | 0.098816 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 56.13 | 34.31 | 0.75067 | 55.3793 | 33.5593 | 0.4845 | 0.11416 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 53.68 | 32.42 | 0.5535 | 53.1265 | 31.8665 | 0.47207 | 0.096393 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 32.1 | 19.98 | 0.5535 | 31.5465 | 19.4265 | 0.26912 | 0.077427 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 34.76 | 21.95 | 0.5535 | 34.2065 | 21.3965 | 0.28444 | 0.097249 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L8 | 14 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 33.03 | 19.27 | 0.57567 | 32.4543 | 18.6943 | 0.30165 | 0.031745 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR81
EN715 | 20240506 | L8 | 14 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 31.88 | 19.44 | 0.57567 | 31.3043 | 18.8643 | 0.27271 | 0.063714 | 20240307 | 13-91 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR82
EN715 | 20240506 | L8 | 14 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 38.54 | 22.31 | 0.57567 | 37.9643 | 21.7343 | 0.18954 | 0.016947 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR87
EN715 | 20240506 | L8 | 14 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 36.69 | 21.14 | 0.57567 | 36.1143 | 20.5643 | 0.18094 | 0.013723 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR88
EN715 | 20240506 | L3 | 15 | 4 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 57.92 | 45.62 | 0.75067 | 57.1693 | 44.8693 | 0.27312 | 0.5273 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phae>Chl
EN715 | 20240506 | L3 | 15 | 4 |  |  | b | 279 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 62.66 | 49.37 | 0.5535 | 62.1065 | 48.8165 | 0.30673 | 0.59843 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phae>Chl
EN715 | 20240506 | L3 | 15 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 112.6 | 78.78 | 0.5535 | 112.0465 | 78.2265 | 0.75096 | 0.64451 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L3 | 15 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 117.2 | 81.66 | 0.5535 | 116.6465 | 81.1065 | 0.78915 | 0.6577 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L3 | 15 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 241.9 | 152.6 | 0.5535 | 241.3465 | 152.0465 | 1.9829 | 0.72947 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L3 | 15 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 238.3 | 151.5 | 0.5535 | 237.7465 | 150.9465 | 1.9274 | 0.76536 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L3 | 15 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 69.41 | 47.83 | 0.57567 | 68.8343 | 47.2543 | 0.47308 | 0.36965 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR97
EN715 | 20240506 | L3 | 15 | 9 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 70.01 | 46.66 | 0.57567 | 69.4343 | 46.0843 | 0.51188 | 0.30998 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR98
EN715 | 20240506 | L3 | 15 | 9 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 26.09 | 20.45 | 0.57567 | 25.5143 | 19.8743 | 0.065747 | 0.12273 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR103 | Phaeo>Chl
EN715 | 20240506 | L3 | 15 | 9 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 19.31 | 14.39 | 0.57567 | 18.7343 | 13.8143 | 0.057354 | 0.073653 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR104 | Phaeo>Chl
EN715 | 20240506 | L3 | 15 | 12 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 202.2 | 121.2 | 0.75067 | 201.4493 | 120.4493 | 3.5971 | 0.70022 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L3 | 15 | 12 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 202.2 | 121.6 | 0.5535 | 201.6465 | 121.0465 | 3.5794 | 0.73929 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L3 | 15 | 12 |  |  | a | 290 | 10 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 165.9 | 99.34 | 0.5535 | 165.3465 | 98.7865 | 2.9559 | 0.56861 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN715 | 20240506 | L3 | 15 | 12 |  |  | b | 290 | 10 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 201.8 | 119.7 | 0.5535 | 201.2465 | 119.1465 | 3.646 | 0.60489 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate
EN715 | 20240506 | L3 | 15 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 90.25 | 56.43 | 0.57567 | 89.6743 | 55.8543 | 0.74141 | 0.2547 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR95
EN715 | 20240506 | L3 | 15 | 12 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 91.58 | 57.68 | 0.57567 | 91.0043 | 57.1043 | 0.74063 | 0.2743 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR96
EN715 | 20240506 | L3 | 15 | 12 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.97 | 11.35 | 0.57567 | 14.3943 | 10.7743 | 0.042276 | 0.060087 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR101 | Phaeo>Chl
EN715 | 20240506 | L3 | 15 | 12 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.67 | 11.63 | 0.57567 | 16.0943 | 11.0543 | 0.05886 | 0.046163 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR102
EN715 | 20240506 | L3 | 15 | 17 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 201.9 | 121.5 | 0.5535 | 201.3465 | 120.9465 | 3.5705 | 0.7446 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L3 | 15 | 17 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 214.9 | 129 | 0.5535 | 214.3465 | 128.4465 | 3.8147 | 0.76794 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L3 | 15 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 167.7 | 103.6 | 0.75067 | 166.9493 | 102.8493 | 1.4233 | 0.4114 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L3 | 15 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 180.6 | 108.5 | 0.5535 | 180.0465 | 107.9465 | 1.6009 | 0.32469 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L3 | 15 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 165 | 101.2 | 0.5535 | 164.4465 | 100.6465 | 1.4166 | 0.37877 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L3 | 15 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 146.1 | 89.13 | 0.5535 | 145.5465 | 88.5765 | 1.265 | 0.31511 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240506 | L3 | 15 | 20 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 55.34 | 34.28 | 0.57567 | 54.7643 | 33.7043 | 0.46168 | 0.1394 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR93
EN715 | 20240506 | L3 | 15 | 20 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 55.01 | 33.93 | 0.57567 | 54.4343 | 33.3543 | 0.46212 | 0.13272 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR94
EN715 | 20240506 | L3 | 15 | 20 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 17.65 | 11.84 | 0.57567 | 17.0743 | 11.2643 | 0.067852 | 0.039166 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR99
EN715 | 20240506 | L3 | 15 | 20 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 19.28 | 12.13 | 0.57567 | 18.7043 | 11.5543 | 0.083199 | 0.026177 | 20240307 | 13-92 | 20240808 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR100
EN715 | 20240507 | L9 | 18 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 59.75 | 42.07 | 0.5535 | 59.1965 | 41.5165 | 0.39258 | 0.34803 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L9 | 18 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 61.6 | 42.62 | 0.5535 | 61.0465 | 42.0665 | 0.42144 | 0.32898 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L9 | 18 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 149.6 | 98.45 | 0.5535 | 149.0465 | 97.8965 | 1.1358 | 0.6106 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L9 | 18 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 153.9 | 100.5 | 0.5535 | 153.3465 | 99.9465 | 1.1857 | 0.59721 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L9 | 18 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 111.4 | 72.61 | 0.75067 | 110.6493 | 71.8593 | 0.86131 | 0.42057 | 20250529 | 15-31 | 20250709 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L9 | 18 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 110.4 | 71.53 | 0.5535 | 109.8465 | 70.9765 | 0.86309 | 0.40305 | 20250529 | 15-27 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L9 | 18 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 81.91 | 52.75 | 0.68567 | 81.2243 | 52.0643 | 0.64586 | 0.29225 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR109
EN715 | 20240507 | L9 | 18 | 9 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 79.25 | 51.1 | 0.68567 | 78.5643 | 50.4143 | 0.62349 | 0.28489 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR110
EN715 | 20240507 | L9 | 18 | 9 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 37.61 | 25.56 | 0.68567 | 36.9243 | 24.8743 | 0.26416 | 0.17945 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR115
EN715 | 20240507 | L9 | 18 | 9 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 44.09 | 32.38 | 0.68567 | 43.4043 | 31.6943 | 0.25671 | 0.30853 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR116 | Phaeo>Chl
EN715 | 20240507 | L9 | 18 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 147.5 | 96.89 | 0.5965 | 146.9035 | 96.2935 | 0.59146 | 0.31263 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L9 | 18 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 148.8 | 98.67 | 0.5535 | 148.2465 | 98.1165 | 0.58585 | 0.33535 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L9 | 18 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 91.15 | 61.91 | 0.68567 | 90.4643 | 61.2243 | 0.641 | 0.45087 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR107
EN715 | 20240507 | L9 | 18 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 94.67 | 63.29 | 0.68567 | 93.9843 | 62.6043 | 0.68558 | 0.42711 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR108
EN715 | 20240507 | L9 | 18 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 62.23 | 42.55 | 0.68567 | 61.5443 | 41.8643 | 0.22983 | 0.1679 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR113
EN715 | 20240507 | L9 | 18 | 14 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 59.23 | 42.65 | 0.68567 | 58.5443 | 41.9643 | 0.19363 | 0.20506 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR114 | Phaeo>Chl
EN715 | 20240507 | L9 | 18 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 58.71 | 37.4 | 0.5535 | 58.1565 | 36.8465 | 0.47318 | 0.18412 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L9 | 18 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 59.85 | 37.33 | 0.5535 | 59.2965 | 36.7765 | 0.50005 | 0.156 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L9 | 18 | 22 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 53.05 | 32.95 | 0.5535 | 52.4965 | 32.3965 | 0.44631 | 0.13161 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L9 | 18 | 22 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 58.41 | 36.21 | 0.5535 | 57.8565 | 35.6565 | 0.49294 | 0.14313 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L9 | 18 | 22 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 30.75 | 19.26 | 0.5535 | 30.1965 | 18.7065 | 0.25513 | 0.078572 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L9 | 18 | 22 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 40.54 | 25.36 | 0.5535 | 39.9865 | 24.8065 | 0.33706 | 0.10545 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L9 | 18 | 22 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 27.93 | 17.04 | 0.68567 | 27.2443 | 16.3543 | 0.23873 | 0.05293 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR105
EN715 | 20240507 | L9 | 18 | 22 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 26.21 | 15.81 | 0.68567 | 25.5243 | 15.1243 | 0.22799 | 0.041736 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR106
EN715 | 20240507 | L9 | 18 | 22 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 29.87 | 18.21 | 0.68567 | 29.1843 | 17.5243 | 0.13617 | 0.030321 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR111
EN715 | 20240507 | L9 | 18 | 22 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 31.62 | 18.95 | 0.68567 | 30.9343 | 18.2643 | 0.14743 | 0.025463 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR112
EN715 | 20240507 | L7 | 19 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 160.3 | 128.8 | 0.5965 | 159.7035 | 128.2035 | 0.69944 | 1.5876 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate; Phaeo>Chl
EN715 | 20240507 | L7 | 19 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 97.14 | 79.15 | 0.5965 | 96.5435 | 78.5535 | 0.39946 | 1.0018 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | poor replicate; Phaeo>Chl
EN715 | 20240507 | L7 | 19 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 97.71 | 77.79 | 0.5535 | 97.1565 | 77.2365 | 0.44231 | 0.9355 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN715 | 20240507 | L7 | 19 | 5 |  |  | b | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240507 | L7 | 19 | 9 |  |  | a | 290 | 0 | 15 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 308.2 | 197.2 | 0.5535 | 307.6465 | 196.6465 | 7.3941 | 3.1297 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L7 | 19 | 9 |  |  | b | 290 | 0 | 15 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 310 | 194.9 | 0.5535 | 309.4465 | 194.3465 | 7.6672 | 2.7335 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L7 | 19 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 51.59 | 41.93 | 0.5965 | 50.9935 | 41.3335 | 0.2145 | 0.52285 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN715 | 20240507 | L7 | 19 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 |  |  | 0.5535 | -0.5535 | -0.5535 | 0 | -0.0098738 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446
EN715 | 20240507 | L7 | 19 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 263.3 | 176.7 | 0.68567 | 262.6143 | 176.0143 | 5.6954 | 3.7217 | 20240307 | 13-94 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR121 | poor replicate
EN715 | 20240507 | L7 | 19 | 9 |  |  | b | 294 | 5 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 283.6 | 182.8 | 0.68567 | 282.9143 | 182.1143 | 6.6067 | 3.1036 | 20240307 | 13-94 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR122 | poor replicate
EN715 | 20240507 | L7 | 19 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 444.7 | 297.3 | 0.68567 | 444.0143 | 296.6143 | 5.1642 | 3.2898 | 20240307 | 13-94 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR127
EN715 | 20240507 | L7 | 19 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 452.8 | 297.1 | 0.68567 | 452.1143 | 296.4143 | 5.455 | 2.9933 | 20240307 | 13-94 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR128
EN715 | 20240507 | L7 | 19 | 12 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 357.4 | 237.2 | 0.5965 | 356.8035 | 236.6035 | 5.338 | 3.1035 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L7 | 19 | 12 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240507 | L7 | 19 | 12 |  |  | a | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240507 | L7 | 19 | 12 |  |  | b | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN715 | 20240507 | L7 | 19 | 12 |  |  | a | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 185.7 | 134.5 | 0.68567 | 185.0143 | 133.8143 | 3.3672 | 3.7921 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR119 | Phaeo>Chl
EN715 | 20240507 | L7 | 19 | 12 |  |  | b | 293 | 5 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 202.7 | 146.4 | 0.68567 | 202.0143 | 145.7143 | 3.7026 | 4.0933 | 20240307 | 13-94 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR120 | Phaeo>Chl
EN715 | 20240507 | L7 | 19 | 12 |  |  | a | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 342.6 | 239.1 | 0.68567 | 341.9143 | 238.4143 | 3.6196 | 3.1633 | 20240307 | 13-94 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR125
EN715 | 20240507 | L7 | 19 | 12 |  |  | b | 551 | 20 | 5 | 1 | 0 | 0.33333 | Cal_030724 | 1.8135 | 0.57627 | 342.7 | 238.3 | 0.68567 | 342.0143 | 237.6143 | 3.6511 | 3.1091 | 20240307 | 13-94 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR126
EN715 | 20240507 | L7 | 19 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 86.31 | 54.65 | 0.5535 | 85.7565 | 54.0965 | 0.703 | 0.26202 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L7 | 19 | 15 |  |  | b | 283 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 82.52 | 52.45 | 0.5535 | 81.9665 | 51.8965 | 0.68421 | 0.26447 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L7 | 19 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 69.88 | 43.02 | 0.5965 | 69.2835 | 42.4235 | 0.59641 | 0.16037 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L7 | 19 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 71.27 | 44.06 | 0.5535 | 70.7165 | 43.5065 | 0.60419 | 0.17192 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L7 | 19 | 21 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 72.58 | 45.15 | 0.5535 | 72.0265 | 44.5965 | 0.60907 | 0.18648 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L7 | 19 | 21 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 64.5 | 40.09 | 0.5965 | 63.9035 | 39.4935 | 0.54201 | 0.16251 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L7 | 19 | 21 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 50.93 | 32.01 | 0.5965 | 50.3335 | 31.4135 | 0.42011 | 0.14027 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L7 | 19 | 21 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 40.03 | 25.53 | 0.5965 | 39.4335 | 24.9335 | 0.32197 | 0.12282 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240507 | L7 | 19 | 21 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 23.01 | 13.83 | 0.68567 | 22.3243 | 13.1443 | 0.20124 | 0.03317 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR117
EN715 | 20240507 | L7 | 19 | 21 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 21.02 | 12.97 | 0.68567 | 20.3343 | 12.2843 | 0.17647 | 0.042605 | 20240307 | 13-93 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR118
EN715 | 20240507 | L7 | 19 | 21 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.07 | 15.57 | 0.68567 | 24.3843 | 14.8843 | 0.11095 | 0.030465 | 20240307 | 13-94 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR123
EN715 | 20240507 | L7 | 19 | 21 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.47 | 15.04 | 0.68567 | 24.7843 | 14.3543 | 0.12137 | 0.014515 | 20240307 | 13-94 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR124
EN715 | 20240508 | MVCO | 23 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 167.9 | 107.1 | 0.5535 | 167.3465 | 106.5465 | 1.35 | 0.55063 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240508 | MVCO | 23 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 167.7 | 106.7 | 0.5965 | 167.1035 | 106.1035 | 1.3545 | 0.53829 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240508 | MVCO | 23 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 136.3 | 87.38 | 0.5965 | 135.7035 | 86.7835 | 1.0862 | 0.46187 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240508 | MVCO | 23 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 133.1 | 83.81 | 0.5965 | 132.5035 | 83.2135 | 1.0945 | 0.38997 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240508 | MVCO | 23 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 122.1 | 78.09 | 0.5965 | 121.5035 | 77.4935 | 0.97722 | 0.40517 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240508 | MVCO | 23 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 125.9 | 79.66 | 0.5965 | 125.3035 | 79.0635 | 1.0267 | 0.38366 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240508 | MVCO | 23 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 57.23 | 37.92 | 0.5535 | 56.6765 | 37.3665 | 0.42877 | 0.23781 | 20250529 | 15-28 | 20250701 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240508 | MVCO | 23 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 54.04 | 35.78 | 0.5965 | 53.4435 | 35.1835 | 0.40545 | 0.22218 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN715 | 20240508 | MVCO | 23 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 68.21 | 42.42 | 0.68567 | 67.5243 | 41.7343 | 0.56537 | 0.17892 | 20240307 | 13-94 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR129
EN715 | 20240508 | MVCO | 23 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 64.13 | 39.93 | 0.68567 | 63.4443 | 39.2443 | 0.53051 | 0.16937 | 20240307 | 13-94 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR130
EN715 | 20240508 | MVCO | 23 | 14 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 57.25 | 34.63 | 0.68567 | 56.5643 | 33.9443 | 0.26417 | 0.058324 | 20240307 | 13-94 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR131
EN715 | 20240508 | MVCO | 23 | 14 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 62.59 | 38.88 | 0.68567 | 61.9043 | 38.1943 | 0.27589 | 0.085661 | 20240307 | 13-94 | 20241017 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR132
EN720 | 20240906 | L1 | 2 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 90.73 | 56.84 | 0.68367 | 90.0463 | 56.1563 | 0.75063 | 0.26122 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240906 | L1 | 2 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 94.43 | 60.37 | 0.68367 | 93.7463 | 59.6863 | 0.75439 | 0.32106 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240906 | L1 | 2 | 2 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 69.38 | 44.51 | 0.69667 | 68.6833 | 43.8133 | 0.55223 | 0.22935 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240906 | L1 | 2 | 2 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 63.2 | 41.23 | 0.68367 | 62.5163 | 40.5463 | 0.48661 | 0.24397 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240906 | L1 | 2 | 2 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 29.3 | 19.04 | 0.59 | 28.71 | 18.45 | 0.22492 | 0.10412 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR5
EN720 | 20240906 | L1 | 2 | 2 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 26.54 | 16.72 | 0.59 | 25.95 | 16.13 | 0.21527 | 0.072386 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR6
EN720 | 20240906 | L1 | 2 | 2 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 21.03 | 13.86 | 0.59 | 20.44 | 13.27 | 0.083583 | 0.042261 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR11
EN720 | 20240906 | L1 | 2 | 2 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.82 | 10.98 | 0.59 | 16.23 | 10.39 | 0.068079 | 0.030454 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR12
EN720 | 20240906 | L1 | 2 | 7 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN720 | 20240906 | L1 | 2 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 64.48 | 40.32 | 0.5965 | 63.8835 | 39.7235 | 0.53646 | 0.17216 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240906 | L1 | 2 | 7 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 24.64 | 15.59 | 0.59 | 24.05 | 15 | 0.19839 | 0.069114 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR3
EN720 | 20240906 | L1 | 2 | 7 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.1 | 15.54 | 0.59 | 24.51 | 14.95 | 0.20886 | 0.056848 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR4
EN720 | 20240906 | L1 | 2 | 7 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 19.95 | 12.2 | 0.59 | 19.36 | 11.61 | 0.090508 | 0.019794 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR9
EN720 | 20240906 | L1 | 2 | 7 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 19.05 | 11.67 | 0.59 | 18.46 | 11.08 | 0.086187 | 0.019079 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR10
EN720 | 20240906 | L1 | 2 | 13 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 92.37 | 58.68 | 0.5965 | 91.7735 | 58.0835 | 0.74807 | 0.28807 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240906 | L1 | 2 | 13 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 99.18 | 62.33 | 0.69667 | 98.4833 | 61.6333 | 0.81824 | 0.28123 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240906 | L1 | 2 | 13 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 65.01 | 41.18 | 0.5965 | 64.4135 | 40.5835 | 0.52913 | 0.19483 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240906 | L1 | 2 | 13 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 72.19 | 46.57 | 0.68367 | 71.5063 | 45.8863 | 0.56745 | 0.25934 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240906 | L1 | 2 | 13 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 38.96 | 24.09 | 0.59 | 38.37 | 23.5 | 0.32598 | 0.093115 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR1
EN720 | 20240906 | L1 | 2 | 13 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 35.73 | 22.04 | 0.59 | 35.14 | 21.45 | 0.30011 | 0.082424 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR2
EN720 | 20240906 | L1 | 2 | 13 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 40.56 | 24.71 | 0.59 | 39.97 | 24.12 | 0.1851 | 0.044051 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR7
EN720 | 20240906 | L1 | 2 | 13 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 38.8 | 23.99 | 0.59 | 38.21 | 23.4 | 0.17233 | 0.049177 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR8
EN720 | 20240907 | L2 | 3 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 87.78 | 62.13 | 0.69667 | 87.0833 | 61.4333 | 0.56955 | 0.52635 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 1 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 45.31 | 32.03 | 0.69667 | 44.6133 | 31.3333 | 0.58975 | 0.52815 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 141.9 | 94.45 | 0.69667 | 141.2033 | 93.7533 | 1.0536 | 0.61885 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 133.5 | 89.4 | 0.69617 | 132.8038 | 88.7038 | 0.97677 | 0.62154 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 6 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 137.8 | 88.06 | 0.69667 | 137.1033 | 87.3633 | 2.2089 | 0.90802 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 265.3 | 170.7 | 0.69667 | 264.6033 | 170.0033 | 2.1005 | 0.93212 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 6 |  |  | a | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN720 | 20240907 | L2 | 3 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 172.1 | 112 | 0.7975 | 171.3025 | 111.2025 | 1.3311 | 0.67254 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 139.5 | 88.99 | 0.59 | 138.91 | 88.4 | 1.1073 | 0.46923 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR17
EN720 | 20240907 | L2 | 3 | 6 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 149.3 | 96.46 | 0.59 | 148.71 | 95.87 | 1.1584 | 0.55137 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR18
EN720 | 20240907 | L2 | 3 | 6 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 133.3 | 81.65 | 0.59 | 132.71 | 81.06 | 0.6021 | 0.16662 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR23
EN720 | 20240907 | L2 | 3 | 6 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 165.3 | 101.8 | 0.59 | 164.71 | 101.21 | 0.74024 | 0.21957 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR24
EN720 | 20240907 | L2 | 3 | 10 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 167.8 | 109.9 | 0.5965 | 167.2035 | 109.3035 | 2.5713 | 1.3284 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 10 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 169.5 | 109.3 | 0.69667 | 168.8033 | 108.6033 | 2.6734 | 1.2013 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 10 |  |  | a | 290 | 10 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 132 | 85.25 | 0.69667 | 131.3033 | 84.5533 | 2.0761 | 0.94054 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 10 |  |  | b | 290 | 10 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 137.7 | 88.21 | 0.5965 | 137.1035 | 87.6135 | 2.1978 | 0.92804 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 111.3 | 73.62 | 0.59 | 110.71 | 73.03 | 0.82602 | 0.47639 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR15
EN720 | 20240907 | L2 | 3 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 106.2 | 69.73 | 0.59 | 105.61 | 69.14 | 0.79678 | 0.43206 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR16
EN720 | 20240907 | L2 | 3 | 10 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 54.47 | 35.11 | 0.59 | 53.88 | 34.52 | 0.2261 | 0.10187 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR21
EN720 | 20240907 | L2 | 3 | 10 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 54.75 | 35.52 | 0.59 | 54.16 | 34.93 | 0.22458 | 0.10728 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR22
EN720 | 20240907 | L2 | 3 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 38.39 | 24.63 | 0.69667 | 37.6933 | 23.9333 | 0.30553 | 0.12141 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 37.62 | 24.7 | 0.5965 | 37.0235 | 24.1035 | 0.28688 | 0.1431 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 23.6 | 15.87 | 0.5965 | 23.0035 | 15.2735 | 0.17164 | 0.10082 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 25.21 | 16.83 | 0.69667 | 24.5133 | 16.1333 | 0.18607 | 0.10173 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L2 | 3 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.75 | 10.43 | 0.59 | 16.16 | 9.84 | 0.13855 | 0.036938 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR13
EN720 | 20240907 | L2 | 3 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.96 | 10.14 | 0.59 | 15.37 | 9.55 | 0.12759 | 0.042727 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR14
EN720 | 20240907 | L2 | 3 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 22.88 | 14.32 | 0.59 | 22.29 | 13.73 | 0.099968 | 0.030475 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR19
EN720 | 20240907 | L2 | 3 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 19.45 | 12.04 | 0.59 | 18.86 | 11.45 | 0.086224 | 0.022164 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR20
EN720 | 20240907 | L4 | 6 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 62.84 | 44.71 | 0.68367 | 62.1563 | 44.0263 | 0.40156 | 0.39173 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | C6N3 chl collection delayed; tissue capsule was not marked with an A or a B
EN720 | 20240907 | L4 | 6 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 62.66 | 43.55 | 0.5965 | 62.0635 | 42.9535 | 0.42433 | 0.34191 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | C6N3 chl collection delayed; tissue capsule was not marked with an A or a B
EN720 | 20240907 | L4 | 6 | 5 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 187 | 123.7 | 0.68367 | 186.3163 | 123.0163 | 2.804 | 1.6291 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | tissue capsule was not marked with an A, B, <10 A, or <10 B
EN720 | 20240907 | L4 | 6 | 5 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 199.1 | 132.6 | 0.5965 | 198.5035 | 132.0035 | 2.9532 | 1.7564 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | tissue capsule was not marked with an A, B, <10 A, or <10 B; tissue capsule marked as C6N11, but numbers match better with C6N5
EN720 | 20240907 | L4 | 6 | 5 |  |  | a | 290 | 10 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 176 | 114.4 | 0.69667 | 175.3033 | 113.7033 | 2.7356 | 1.3211 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | tissue capsule was not marked with an A, B, <10 A, or <10 B
EN720 | 20240907 | L4 | 6 | 5 |  |  | b | 290 | 10 | 10 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 165.1 | 107.5 | 0.5965 | 164.5035 | 106.9035 | 2.558 | 1.2561 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | tissue capsule was not marked with an A, B, <10 A, or <10 B
EN720 | 20240907 | L4 | 6 | 5 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 121.4 | 79.87 | 0.59 | 120.81 | 79.28 | 0.91042 | 0.50345 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR29
EN720 | 20240907 | L4 | 6 | 5 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 116.1 | 77.55 | 0.59 | 115.51 | 76.96 | 0.8451 | 0.5274 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR30
EN720 | 20240907 | L4 | 6 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 34.64 | 24.57 | 0.59 | 34.05 | 23.98 | 0.11739 | 0.11002 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR35
EN720 | 20240907 | L4 | 6 | 5 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 29.02 | 20.35 | 0.59 | 28.43 | 19.76 | 0.10107 | 0.086323 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR36
EN720 | 20240907 | L4 | 6 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 41.89 | 26.84 | 0.69667 | 41.1933 | 26.1433 | 0.33418 | 0.13219 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | tissue capsule not marked with an A or a B
EN720 | 20240907 | L4 | 6 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 27.36 | 17.49 | 0.5965 | 26.7635 | 16.8935 | 0.21916 | 0.082202 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | tissue capsule not marked with an A or a B
EN720 | 20240907 | L4 | 6 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 36.62 | 22.81 | 0.68367 | 35.9363 | 22.1263 | 0.30588 | 0.092805 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3
EN720 | 20240907 | L4 | 6 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 35.62 | 22.21 | 0.7975 | 34.8225 | 21.4125 | 0.29702 | 0.088803 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3
EN720 | 20240907 | L4 | 6 | 11 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.645 | 6.048 | 0.59 | 8.055 | 5.458 | 0.056932 | 0.040406 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR27
EN720 | 20240907 | L4 | 6 | 11 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.024 | 5.427 | 0.59 | 7.434 | 4.837 | 0.056738 | 0.029231 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR28
EN720 | 20240907 | L4 | 6 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.819 | 3.742 | 0.59 | 4.229 | 3.152 | 0.012578 | 0.017368 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR33 | Phaeo>Chl; depth = 24
EN720 | 20240907 | L4 | 6 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.617 | 5.291 | 0.59 | 7.027 | 4.701 | 0.027164 | 0.017498 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR34
EN720 | 20240907 | L4 | 6 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 40.75 | 25.32 | 0.68367 | 40.0663 | 24.6363 | 0.34176 | 0.10215 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | C6N14 chl collection delayed; tissue capsule was not marked with an A or a B
EN720 | 20240907 | L4 | 6 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 31.82 | 19.91 | 0.69667 | 31.1233 | 19.2133 | 0.26446 | 0.078288 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | C6N14 chl collection delayed; tissue capsule was not marked with an A or a B
EN720 | 20240907 | L4 | 6 | 14 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.385 | 3.691 | 0.59 | 4.795 | 3.101 | 0.03701 | 0.018105 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR37
EN720 | 20240907 | L4 | 6 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.849 | 5.447 | 0.59 | 7.259 | 4.857 | 0.052478 | 0.033847 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR38
EN720 | 20240907 | L4 | 6 | 14 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.032 | 4.696 | 0.59 | 5.442 | 4.106 | 0.015574 | 0.023365 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR39 | Phaeo>Chl; depth = 15
EN720 | 20240907 | L4 | 6 | 14 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.226 | 3.724 | 0.59 | 4.636 | 3.134 | 0.017478 | 0.012189 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR40
EN720 | 20240907 | L4 | 6 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 27.76 | 17.83 | 0.68367 | 27.0763 | 17.1463 | 0.21994 | 0.089011 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L4 | 6 | 18 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN720 | 20240907 | L4 | 6 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 25.5 | 15.67 | 0.5965 | 24.9035 | 15.0735 | 0.21827 | 0.050623 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L4 | 6 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 25.23 | 15.85 | 0.69667 | 24.5333 | 15.1533 | 0.20828 | 0.06204 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L4 | 6 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.324 | 3.721 | 0.59 | 4.734 | 3.131 | 0.035141 | 0.020697 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR25
EN720 | 20240907 | L4 | 6 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.575 | 5.36 | 0.59 | 6.985 | 4.77 | 0.048557 | 0.03651 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR26
EN720 | 20240907 | L4 | 6 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.417 | 6.589 | 0.59 | 8.827 | 5.999 | 0.033027 | 0.023967 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR31
EN720 | 20240907 | L4 | 6 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.027 | 3.168 | 0.59 | 3.437 | 2.578 | 0.0099955 | 0.014408 | 20240307 | 15-21 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR32 | Phaeo>Chl; depth = 3
EN720 | 20240907 | L5 | 7 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 47.02 | 34.92 | 0.69667 | 46.3233 | 34.2233 | 0.26867 | 0.34183 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN720 | 20240907 | L5 | 7 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 47.73 | 35.18 | 0.5965 | 47.1335 | 34.5835 | 0.27867 | 0.33826 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Phaeo>Chl
EN720 | 20240907 | L5 | 7 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 122.3 | 85.49 | 0.5965 | 121.7035 | 84.8935 | 0.81735 | 0.69705 | 20250529 | 15-25 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L5 | 7 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 125 | 86.34 | 0.5965 | 124.4035 | 85.7435 | 0.85843 | 0.67114 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L5 | 7 | 6 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 18.62 | 13.38 | 0.59 | 18.03 | 12.79 | 0.11487 | 0.11322 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR43
EN720 | 20240907 | L5 | 7 | 6 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.73 | 11.99 | 0.59 | 16.14 | 11.4 | 0.10356 | 0.099058 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR44
EN720 | 20240907 | L5 | 7 | 6 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.785 | 3.567 | 0.59 | 4.195 | 2.977 | 0.014224 | 0.014059 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR49 | Phaeo=Chl; depth = 42
EN720 | 20240907 | L5 | 7 | 6 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.617 | 3.647 | 0.59 | 4.027 | 3.057 | 0.011328 | 0.017715 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR50 | Phaeo>Chl; depth = 42
EN720 | 20240907 | L5 | 7 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 60.36 | 38.67 | 0.5965 | 59.7635 | 38.0735 | 0.48162 | 0.19757 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L5 | 7 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 60.12 | 38.41 | 0.69667 | 59.4233 | 37.7133 | 0.48206 | 0.1907 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L5 | 7 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 10.02 | 6.712 | 0.59 | 9.43 | 6.122 | 0.072518 | 0.036661 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR45
EN720 | 20240907 | L5 | 7 | 9 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.1 | 7.679 | 0.59 | 10.51 | 7.089 | 0.074995 | 0.051429 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR46
EN720 | 20240907 | L5 | 7 | 9 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.331 | 3.787 | 0.59 | 4.741 | 3.197 | 0.017999 | 0.012319 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR51
EN720 | 20240907 | L5 | 7 | 9 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.006 | 2.831 | 0.59 | 3.416 | 2.241 | 0.013697 | 0.0075549 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR52
EN720 | 20240907 | L5 | 7 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 30.78 | 19.15 | 0.5965 | 30.1835 | 18.5535 | 0.25824 | 0.072734 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L5 | 7 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 36.77 | 22.78 | 0.5965 | 36.1735 | 22.1835 | 0.31064 | 0.085087 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L5 | 7 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 21.07 | 13.8 | 0.69667 | 20.3733 | 13.1033 | 0.16143 | 0.072321 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L5 | 7 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 19.33 | 12 | 0.5965 | 18.7335 | 11.4035 | 0.16276 | 0.040666 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L5 | 7 | 15 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 15.44 | 9.94 | 0.5965 | 14.8435 | 9.3435 | 0.12212 | 0.044552 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L5 | 7 | 15 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 18.82 | 12.18 | 0.68367 | 18.1363 | 11.4963 | 0.14707 | 0.060077 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240907 | L5 | 7 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.43 | 4.697 | 0.59 | 5.84 | 4.107 | 0.037991 | 0.035253 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR41
EN720 | 20240907 | L5 | 7 | 15 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.808 | 3.283 | 0.59 | 4.218 | 2.693 | 0.033431 | 0.014596 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR42
EN720 | 20240907 | L5 | 7 | 15 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.159 | 2.925 | 0.59 | 3.569 | 2.335 | 0.014411 | 0.0077727 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR47
EN720 | 20240907 | L5 | 7 | 15 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.462 | 2.352 | 0.59 | 2.872 | 1.762 | 0.012916 | 0.0037633 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR48
EN720 | 20240908 | L10 | 8 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 28.29 | 20.49 | 0.69667 | 27.5933 | 19.7933 | 0.1732 | 0.1799 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl
EN720 | 20240908 | L10 | 8 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 26.16 | 18.88 | 0.69667 | 25.4633 | 18.1833 | 0.16165 | 0.16272 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl
EN720 | 20240908 | L10 | 8 | 6 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 23.66 | 16.78 | 0.69667 | 22.9633 | 16.0833 | 0.15277 | 0.13414 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L10 | 8 | 6 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 22.7 | 16.2 | 0.69667 | 22.0033 | 15.5033 | 0.14433 | 0.13223 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L10 | 8 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 144.9 | 99.68 | 0.5965 | 144.3035 | 99.0835 | 1.0041 | 0.76345 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L10 | 8 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 147.5 | 103 | 0.5965 | 146.9035 | 102.4035 | 0.9881 | 0.83866 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L10 | 8 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 80.19 | 55.71 | 0.5965 | 79.5935 | 55.1135 | 0.54357 | 0.43959 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L10 | 8 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 83.43 | 58.79 | 0.68367 | 82.7463 | 58.1063 | 0.54575 | 0.50124 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L10 | 8 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 72.12 | 49.81 | 0.77 | 71.35 | 49.04 | 0.48908 | 0.38549 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR57
EN720 | 20240908 | L10 | 8 | 9 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 76.03 | 52.31 | 0.77 | 75.26 | 51.54 | 0.51822 | 0.39781 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR58
EN720 | 20240908 | L10 | 8 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 80.1 | 55.18 | 0.77 | 79.33 | 54.41 | 0.29103 | 0.2259 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR63
EN720 | 20240908 | L10 | 8 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 83.33 | 55.6 | 0.77 | 82.56 | 54.83 | 0.32384 | 0.19707 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR64
EN720 | 20240908 | L10 | 8 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 91.2 | 64.94 | 0.7975 | 90.4025 | 64.1425 | 0.58163 | 0.57412 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L10 | 8 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 90.88 | 63.51 | 0.7975 | 90.0825 | 62.7125 | 0.60621 | 0.52376 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L10 | 8 | 17 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN720 | 20240908 | L10 | 8 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 37.53 | 25.65 | 0.7975 | 36.7325 | 24.8525 | 0.26313 | 0.18467 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L10 | 8 | 17 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.088 | 3.597 | 0.77 | 4.318 | 2.827 | 0.032686 | 0.017731 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR55
EN720 | 20240908 | L10 | 8 | 17 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.885 | 3.705 | 0.77 | 4.115 | 2.935 | 0.025868 | 0.026474 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR56 | Phaeo = Chl
EN720 | 20240908 | L10 | 8 | 17 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.889 | 6.707 | 0.77 | 8.119 | 5.937 | 0.025436 | 0.030867 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR61 | poor replicate
EN720 | 20240908 | L10 | 8 | 17 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.535 | 2.371 | 0.77 | 1.765 | 1.601 | 0.0019118 | 0.013271 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR62 | poor replicate; Phaeo>Chl
EN720 | 20240908 | L10 | 8 | 21 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 16.1 | 10.71 | 0.5965 | 15.5035 | 10.1135 | 0.11968 | 0.060731 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L10 | 8 | 21 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 16.05 | 11.02 | 0.5965 | 15.4535 | 10.4235 | 0.11169 | 0.074254 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L10 | 8 | 21 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 14.2 | 9.606 | 0.5965 | 13.6035 | 9.0095 | 0.10201 | 0.058712 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L10 | 8 | 21 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 14.73 | 10 | 0.5965 | 14.1335 | 9.4035 | 0.10503 | 0.06272 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L10 | 8 | 21 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.159 | 2.396 | 0.77 | 2.389 | 1.626 | 0.016727 | 0.012271 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR53
EN720 | 20240908 | L10 | 8 | 21 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.514 | 2.363 | 0.77 | 2.744 | 1.593 | 0.025232 | 0.0031771 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR54
EN720 | 20240908 | L10 | 8 | 21 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.627 | 2.015 | 0.77 | 1.857 | 1.245 | 0.0071472 | 0.004681 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR59
EN720 | 20240908 | L10 | 8 | 21 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.035 | 1.612 | 0.77 | 1.265 | 0.842 | 0.0049221 | 0.0030484 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR60
EN720 | 20240908 | L11 | 11 | 4 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.645 | 3.374 | 0.77 | 2.875 | 2.604 | 0.0059207 | 0.040361 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR71 | Phaeo>Chl; depth = 100m
EN720 | 20240908 | L11 | 11 | 4 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.668 | 3.156 | 0.77 | 2.898 | 2.386 | 0.011186 | 0.031221 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR72 | Phaeo>Chl; depth = 100m
EN720 | 20240908 | L11 | 11 | 4 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.391 | 1.965 | 0.77 | 1.621 | 1.195 | 0.004966 | 0.0063666 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR79 | Phaeo>Chl; depth = 100m
EN720 | 20240908 | L11 | 11 | 4 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.342 | 1.873 | 0.59 | 1.752 | 1.283 | 0.0054574 | 0.0066878 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR80 | Phaeo>Chl; depth = 100m
EN720 | 20240908 | L11 | 11 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 27.92 | 20.82 | 0.7975 | 27.1225 | 20.0225 | 0.15726 | 0.20352 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 67
EN720 | 20240908 | L11 | 11 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 27.47 | 20.59 | 0.5965 | 26.8735 | 19.9935 | 0.15277 | 0.20389 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 108.6 | 76.27 | 0.69667 | 107.9033 | 75.5733 | 0.71787 | 0.63027 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 111.8 | 79.05 | 0.68367 | 111.1163 | 78.3663 | 0.72538 | 0.68666 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 88.16 | 62.62 | 0.7975 | 87.3625 | 61.8225 | 0.56568 | 0.54826 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 96.04 | 68.21 | 0.68367 | 95.3563 | 67.5263 | 0.6164 | 0.60031 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 20.48 | 14.52 | 0.77 | 19.71 | 13.75 | 0.13066 | 0.11456 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR69
EN720 | 20240908 | L11 | 11 | 9 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 20.59 | 14.77 | 0.77 | 19.82 | 14 | 0.12715 | 0.12167 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR70
EN720 | 20240908 | L11 | 11 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.224 | 6.107 | 0.77 | 7.454 | 5.337 | 0.024723 | 0.025981 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR77 | Phaeo>Chl; depth = 47m
EN720 | 20240908 | L11 | 11 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.91 | 5.488 | 0.77 | 7.14 | 4.718 | 0.028285 | 0.016539 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR78
EN720 | 20240908 | L11 | 11 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 23.43 | 16.04 | 0.7975 | 22.6325 | 15.2425 | 0.16368 | 0.11097 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 23.4 | 15.88 | 0.68367 | 22.7163 | 15.1963 | 0.16656 | 0.10725 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 14 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 23.41 | 15.85 | 0.69667 | 22.7133 | 15.1533 | 0.16787 | 0.10245 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 14 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 23.5 | 15.53 | 0.7975 | 22.7025 | 14.7325 | 0.17653 | 0.08893 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 14 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.872 | 3.51 | 0.77 | 4.102 | 2.74 | 0.029858 | 0.019007 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR67
EN720 | 20240908 | L11 | 11 | 14 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.119 | 3.08 | 0.77 | 3.349 | 2.31 | 0.022777 | 0.018419 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR68
EN720 | 20240908 | L11 | 11 | 14 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.691 | 4.425 | 0.77 | 4.921 | 3.655 | 0.014758 | 0.019904 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR75
EN720 | 20240908 | L11 | 11 | 14 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.203 | 1.811 | 0.77 | 1.433 | 1.041 | 0.0045697 | 0.0053025 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR76 | Phaeo = Chl
EN720 | 20240908 | L11 | 11 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 17.98 | 11.88 | 0.69667 | 17.2833 | 11.1833 | 0.13545 | 0.06405 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 18.75 | 12.41 | 0.7975 | 17.9525 | 11.6125 | 0.14042 | 0.068815 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 16.88 | 11.23 | 0.5965 | 16.2835 | 10.6335 | 0.12546 | 0.064234 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 18.64 | 12.43 | 0.68367 | 17.9563 | 11.7463 | 0.13754 | 0.074106 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 15.09 | 10.05 | 0.69667 | 14.3933 | 9.3533 | 0.11191 | 0.054942 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.39 | 10.72 | 0.68367 | 14.7063 | 10.0363 | 0.10344 | 0.077403 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240908 | L11 | 11 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.401 | 2.346 | 0.77 | 2.631 | 1.576 | 0.023128 | 0.0049784 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR65
EN720 | 20240908 | L11 | 11 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.649 | 2.101 | 0.77 | 1.879 | 1.331 | 0.012013 | 0.011724 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR66 | Phaeo = Chl
EN720 | 20240908 | L11 | 11 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.18 | 1.495 | 0.77 | 1.41 | 0.725 | 0.0079998 | -0.0011118 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR73
EN720 | 20240908 | L11 | 11 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.06 | 1.643 | 0.77 | 1.29 | 0.873 | 0.0048523 | 0.0034117 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR74
EN720 | 20240909 | L6 | 14 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 18.24 | 14.23 | 0.68367 | 17.5563 | 13.5463 | 0.088817 | 0.15527 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 66
EN720 | 20240909 | L6 | 14 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 19.01 | 14.82 | 0.68367 | 18.3263 | 14.1363 | 0.092804 | 0.16191 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 66
EN720 | 20240909 | L6 | 14 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 75.02 | 54.69 | 0.69667 | 74.3233 | 53.9933 | 0.45142 | 0.51176 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 52
EN720 | 20240909 | L6 | 14 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 77.46 | 55.5 | 0.68367 | 76.7763 | 54.8163 | 0.48639 | 0.50131 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 52
EN720 | 20240909 | L6 | 14 | 5 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 65.96 | 48.45 | 0.69667 | 65.2633 | 47.7533 | 0.3888 | 0.46306 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 52
EN720 | 20240909 | L6 | 14 | 5 |  |  | b | 285 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 65.55 | 48.63 | 0.5965 | 64.9535 | 48.0335 | 0.38229 | 0.4896 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | spilled ~5mL of C14N5 <10b; Phaeo>Chl; depth=52
EN720 | 20240909 | L6 | 14 | 5 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13.18 | 9.827 | 0.77 | 12.41 | 9.057 | 0.073505 | 0.088017 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR87 | Phaeo>Chl; depth = 52
EN720 | 20240909 | L6 | 14 | 5 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.35 | 9.272 | 0.77 | 11.58 | 8.502 | 0.067476 | 0.084148 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR88 | Phaeo>Chl; depth = 52
EN720 | 20240909 | L6 | 14 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.362 | 1.882 | 0.59 | 1.772 | 1.292 | 0.0055955 | 0.006657 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR95 | Phaeo>Chl; depth = 52
EN720 | 20240909 | L6 | 14 | 5 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.354 | 1.958 | 0.59 | 1.764 | 1.368 | 0.0046163 | 0.008357 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR96 | Phaeo>Chl; depth = 52
EN720 | 20240909 | L6 | 14 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 278.3 | 180.8 | 0.7975 | 277.5025 | 180.0025 | 2.1595 | 1.0838 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L6 | 14 | 8 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 136 | 89.04 | 0.68367 | 135.3163 | 88.3563 | 2.0802 | 1.1039 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L6 | 14 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 264.9 | 173 | 0.68367 | 264.2163 | 172.3163 | 2.0355 | 1.0694 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L6 | 14 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 246 | 162.5 | 0.7975 | 245.2025 | 161.7025 | 1.8494 | 1.0642 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L6 | 14 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 24.76 | 16.68 | 0.59 | 24.17 | 16.09 | 0.17713 | 0.10982 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR85
EN720 | 20240909 | L6 | 14 | 8 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 22.96 | 15.29 | 0.77 | 22.19 | 14.52 | 0.16757 | 0.090497 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR86
EN720 | 20240909 | L6 | 14 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.582 | 2.029 | 0.77 | 1.812 | 1.259 | 0.0064582 | 0.0055031 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR93 | Phaeo = Chl; depth = 35
EN720 | 20240909 | L6 | 14 | 8 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.678 | 1.998 | 0.77 | 1.908 | 1.228 | 0.0079414 | 0.0037254 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR94
EN720 | 20240909 | L6 | 14 | 14 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 82.15 | 53.21 | 0.68367 | 81.4663 | 52.5263 | 0.64099 | 0.30545 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L6 | 14 | 14 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 78.01 | 50.77 | 0.68367 | 77.3263 | 50.0863 | 0.60334 | 0.29914 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L6 | 14 | 14 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.43 | 7.718 | 0.59 | 10.84 | 7.128 | 0.081098 | 0.04559 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR83
EN720 | 20240909 | L6 | 14 | 14 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.55 | 8.521 | 0.59 | 11.96 | 7.931 | 0.088024 | 0.052936 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR84
EN720 | 20240909 | L6 | 14 | 14 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.229 | 6.176 | 0.77 | 8.459 | 5.406 | 0.03559 | 0.015677 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR91
EN720 | 20240909 | L6 | 14 | 14 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.7 | 5.333 | 0.77 | 6.93 | 4.563 | 0.027543 | 0.015651 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR92
EN720 | 20240909 | L6 | 14 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.5 | 16.89 | 0.68367 | 24.8163 | 16.2063 | 0.1907 | 0.10131 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L6 | 14 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 24.7 | 15.77 | 0.5965 | 24.1035 | 15.1735 | 0.19829 | 0.072391 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L6 | 14 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.18 | 16.32 | 0.68367 | 24.4963 | 15.6363 | 0.19624 | 0.085503 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L6 | 14 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 28.16 | 18.14 | 0.7975 | 27.3625 | 17.3425 | 0.22193 | 0.090552 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L6 | 14 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 21.25 | 13.83 | 0.68367 | 20.5663 | 13.1463 | 0.16434 | 0.072531 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L6 | 14 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 20.58 | 13.5 | 0.7975 | 19.7825 | 12.7025 | 0.15681 | 0.072065 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L6 | 14 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.128 | 4.703 | 0.59 | 6.538 | 4.113 | 0.053161 | 0.02019 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR81
EN720 | 20240909 | L6 | 14 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.684 | 3.923 | 0.59 | 5.094 | 3.333 | 0.038605 | 0.020836 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR82
EN720 | 20240909 | L6 | 14 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.914 | 5.474 | 0.77 | 7.144 | 4.704 | 0.028496 | 0.016195 | 20240307 | 15-01 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR89
EN720 | 20240909 | L6 | 14 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.575 | 4.814 | 0.77 | 6.805 | 4.044 | 0.032128 | 0.0061538 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR90
EN720 | 20240909 | L8 | 15 | 3 |  |  | a | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN720 | 20240909 | L8 | 15 | 3 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN720 | 20240909 | L8 | 15 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 44.71 | 35.15 | 0.7975 | 43.9125 | 34.3525 | 0.21174 | 0.40723 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 50
EN720 | 20240909 | L8 | 15 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 44.79 | 34.8 | 0.69617 | 44.0938 | 34.1038 | 0.22127 | 0.39323 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L8 | 15 | 5 |  |  | a | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13.69 | 10.5 | 0.59 | 13.1 | 9.91 | 0.069694 | 0.10644 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR103 | Phaeo>Chl; depth = 50
EN720 | 20240909 | L8 | 15 | 5 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 12.06 | 10.04 | 0.59 | 11.47 | 9.45 | 0.044132 | 0.12383 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR104 | Phaeo>Chl; depth = 50
EN720 | 20240909 | L8 | 15 | 5 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.376 | 3.352 | 0.59 | 3.786 | 2.762 | 0.011937 | 0.014256 | 20240307 | 15-23 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR111 | Phaeo>Chl; depth = 50
EN720 | 20240909 | L8 | 15 | 5 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.151 | 2.905 | 0.59 | 3.561 | 2.315 | 0.014499 | 0.0074155 | 20240307 | 15-23 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR112
EN720 | 20240909 | L8 | 15 | 8 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 50.18 | 34.04 | 0.69617 | 49.4838 | 33.3438 | 0.35748 | 0.24332 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L8 | 15 | 8 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 47.69 | 32.76 | 0.7975 | 46.8925 | 31.9625 | 0.33068 | 0.24523 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L8 | 15 | 8 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 41.83 | 27.79 | 0.7975 | 41.0325 | 26.9925 | 0.31097 | 0.17539 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L8 | 15 | 8 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 39.41 | 26.27 | 0.68367 | 38.7263 | 25.5863 | 0.29104 | 0.16999 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L8 | 15 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.06 | 10.63 | 0.59 | 13.47 | 10.04 | 0.075193 | 0.10386 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR101 | Phaeo>Chl; depth = 38
EN720 | 20240909 | L8 | 15 | 8 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13.9 | 10.49 | 0.59 | 13.31 | 9.9 | 0.074754 | 0.1018 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR102 | Phaeo>Chl; depth = 38
EN720 | 20240909 | L8 | 15 | 8 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.823 | 4.932 | 0.59 | 6.233 | 4.342 | 0.022044 | 0.019133 | 20240307 | 15-23 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR109
EN720 | 20240909 | L8 | 15 | 8 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.569 | 5.147 | 0.59 | 6.979 | 4.557 | 0.028234 | 0.014982 | 20240307 | 15-23 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR110
EN720 | 20240909 | L8 | 15 | 11 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 197.6 | 127.6 | 0.68367 | 196.9163 | 126.9163 | 1.5504 | 0.73641 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L8 | 15 | 11 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 201.6 | 131.9 | 0.68367 | 200.9163 | 131.2163 | 1.5438 | 0.82053 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L8 | 15 | 11 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 179.1 | 115.7 | 0.7975 | 178.3025 | 114.9025 | 1.4042 | 0.66612 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L8 | 15 | 11 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 187.7 | 121.9 | 0.69617 | 187.0038 | 121.2038 | 1.4574 | 0.7265 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L8 | 15 | 11 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 42.04 | 27.05 | 0.59 | 41.45 | 26.46 | 0.32861 | 0.14327 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR99
EN720 | 20240909 | L8 | 15 | 11 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 40.69 | 25.9 | 0.59 | 40.1 | 25.31 | 0.32312 | 0.12672 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR100
EN720 | 20240909 | L8 | 15 | 11 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.457 | 3.2 | 0.59 | 3.867 | 2.61 | 0.01468 | 0.010117 | 20240307 | 15-23 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR107
EN720 | 20240909 | L8 | 15 | 11 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.064 | 3.373 | 0.59 | 4.474 | 2.783 | 0.019748 | 0.0066919 | 20240307 | 15-23 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR108
EN720 | 20240909 | L8 | 15 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 31.55 | 20.88 | 0.5965 | 30.9535 | 20.2835 | 0.23692 | 0.12491 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L8 | 15 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 30.64 | 20.17 | 0.5965 | 30.0435 | 19.5735 | 0.23248 | 0.11669 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L8 | 15 | 18 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13.74 | 9.568 | 0.7975 | 12.9425 | 8.7705 | 0.092405 | 0.065625 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Did the filters go in in the wrong tissue capsules; <10 bigger
EN720 | 20240909 | L8 | 15 | 18 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 17.15 | 11.51 | 0.7975 | 16.3525 | 10.7125 | 0.12492 | 0.068102 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Did the filters go in in the wrong tissue capsules; <10 bigger
EN720 | 20240909 | L8 | 15 | 18 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.52 | 17.08 | 0.68367 | 24.8363 | 16.3963 | 0.18694 | 0.1085 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | Did the filters go in in the wrong tissue capsules; <10 bigger
EN720 | 20240909 | L8 | 15 | 18 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 26.21 | 17.6 | 0.68367 | 25.5263 | 16.9163 | 0.1907 | 0.1141 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | C15N18 <10b pick zoop; was the mesh bad; Did the filters go in in the wrong tissue capsules; <10 bigger
EN720 | 20240909 | L8 | 15 | 18 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.393 | 3.666 | 0.59 | 4.803 | 3.076 | 0.037859 | 0.016998 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR97
EN720 | 20240909 | L8 | 15 | 18 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.294 | 3.806 | 0.59 | 4.704 | 3.216 | 0.03262 | 0.024734 | 20240307 | 15-22 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR98
EN720 | 20240909 | L8 | 15 | 18 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.253 | 3.686 | 0.59 | 4.663 | 3.096 | 0.0183 | 0.011114 | 20240307 | 15-23 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR105
EN720 | 20240909 | L8 | 15 | 18 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.732 | 3.131 | 0.59 | 4.142 | 2.541 | 0.01863 | 0.0054241 | 20240307 | 15-23 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR106
EN720 | 20240909 | L7 | 16 | 2 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.291 | 7.944 | 0.69617 | 8.5948 | 7.2478 | 0.029835 | 0.10076 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phao>Chl; depth = 116
EN720 | 20240909 | L7 | 16 | 2 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.938 | 7.635 | 0.7975 | 8.1405 | 6.8375 | 0.02886 | 0.094341 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phao>Chl; depth = 116
EN720 | 20240909 | L7 | 16 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 32.96 | 24.81 | 0.69617 | 32.2638 | 24.1138 | 0.18051 | 0.25398 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phao>Chl; depth = 65
EN720 | 20240909 | L7 | 16 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 31.91 | 23.92 | 0.69617 | 31.2138 | 23.2238 | 0.17697 | 0.24149 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phao>Chl; depth = 65
EN720 | 20240909 | L7 | 16 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 34.21 | 23.86 | 0.7975 | 33.4125 | 23.0625 | 0.22924 | 0.18631 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L7 | 16 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 33.84 | 23.78 | 0.69617 | 33.1438 | 23.0838 | 0.22282 | 0.19312 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L7 | 16 | 6 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 33.68 | 23.35 | 0.69617 | 32.9838 | 22.6538 | 0.2288 | 0.17939 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L7 | 16 | 6 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 32.58 | 22.75 | 0.68367 | 31.8963 | 22.0663 | 0.21772 | 0.17988 | 20240307 | 15-03 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L7 | 16 | 8 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 235.7 | 152.2 | 0.69617 | 235.0038 | 151.5038 | 3.6989 | 1.7609 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L7 | 16 | 8 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 237 | 152.4 | 0.7975 | 236.2025 | 151.6025 | 3.7476 | 1.7157 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L7 | 16 | 8 |  |  | a | 290 | 10 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 202.4 | 131.2 | 0.7975 | 201.6025 | 130.4025 | 3.154 | 1.5453 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L7 | 16 | 8 |  |  | b | 290 | 10 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 225.2 | 144.8 | 0.68367 | 224.5163 | 144.1163 | 3.5615 | 1.632 | 20240307 | 15-04 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L7 | 16 | 8 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 72.86 | 48.52 | 0.77 | 72.09 | 47.75 | 0.53358 | 0.31799 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR115
EN720 | 20240909 | L7 | 16 | 8 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 80.51 | 52.07 | 0.59 | 79.92 | 51.48 | 0.62134 | 0.29362 | 20240307 | 15-23 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR116
EN720 | 20240909 | L7 | 16 | 8 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.277 | 4.266 | 0.59 | 5.687 | 3.676 | 0.023485 | 0.011439 | 20240307 | 15-23 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR121
EN720 | 20240909 | L7 | 16 | 8 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.142 | 3.411 | 0.77 | 4.372 | 2.641 | 0.020215 | 0.0048756 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR122
EN720 | 20240909 | L7 | 16 | 16 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 47.27 | 30.48 | 0.7975 | 46.4725 | 29.6825 | 0.37188 | 0.16295 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L7 | 16 | 16 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 49.04 | 31.64 | 0.69617 | 48.3438 | 30.9438 | 0.38539 | 0.17217 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L7 | 16 | 16 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 10.94 | 6.779 | 0.59 | 10.35 | 6.189 | 0.091218 | 0.019156 | 20240307 | 15-23 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR117; TAR chl says N17
EN720 | 20240909 | L7 | 16 | 16 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 10.24 | 6.669 | 0.59 | 9.65 | 6.079 | 0.078284 | 0.030129 | 20240307 | 15-23 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR118; TAR chl says N17
EN720 | 20240909 | L7 | 16 | 16 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.674 | 5.836 | 0.77 | 7.904 | 5.066 | 0.033083 | 0.014959 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR123; TAR chl says N17
EN720 | 20240909 | L7 | 16 | 16 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.113 | 5.535 | 0.77 | 7.343 | 4.765 | 0.030053 | 0.015136 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR124; TAR chl says N17
EN720 | 20240909 | L7 | 16 | 19 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 37.39 | 23.77 | 0.68367 | 36.7063 | 23.0863 | 0.30167 | 0.11431 | 20240307 | 15-04 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L7 | 16 | 19 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 41.36 | 26.31 | 0.69617 | 40.6638 | 25.6138 | 0.33334 | 0.12818 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L7 | 16 | 19 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 31.54 | 20.18 | 0.68367 | 30.8563 | 19.4963 | 0.25161 | 0.099682 | 20240307 | 15-04 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L7 | 16 | 19 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 31.95 | 20.6 | 0.69667 | 31.2533 | 19.9033 | 0.25202 | 0.10303 | 20250529 | 15-32 | 20250715 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240909 | L7 | 16 | 19 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.909 | 5.869 | 0.77 | 8.139 | 5.099 | 0.066643 | 0.024292 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR113
EN720 | 20240909 | L7 | 16 | 19 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.27 | 6.281 | 0.77 | 8.5 | 5.511 | 0.065525 | 0.032758 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR114
EN720 | 20240909 | L7 | 16 | 19 |  |  | a | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 11.64 | 7.479 | 0.59 | 11.05 | 6.889 | 0.048418 | 0.016794 | 20240307 | 15-23 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR119
EN720 | 20240909 | L7 | 16 | 19 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 9.827 | 6.434 | 0.59 | 9.237 | 5.844 | 0.039625 | 0.015896 | 20240307 | 15-23 | 20250325 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR120
EN720 | 20240910 | L9 | 19 | 3 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 23.54 | 18.12 | 0.68367 | 22.8563 | 17.4363 | 0.12005 | 0.19413 | 20240307 | 15-04 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 92
EN720 | 20240910 | L9 | 19 | 3 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 23.63 | 18.54 | 0.7975 | 22.8325 | 17.7425 | 0.11274 | 0.20695 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 92
EN720 | 20240910 | L9 | 19 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.72 | 11.56 | 0.7975 | 13.9225 | 10.7625 | 0.06999 | 0.12393 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 79
EN720 | 20240910 | L9 | 19 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 14.79 | 11.44 | 0.68367 | 14.1063 | 10.7563 | 0.074199 | 0.11961 | 20240307 | 15-04 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 79
EN720 | 20240910 | L9 | 19 | 7 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 53.48 | 38.72 | 0.69617 | 52.7838 | 38.0238 | 0.32692 | 0.35821 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 60
EN720 | 20240910 | L9 | 19 | 7 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 51.83 | 37.41 | 0.69617 | 51.1338 | 36.7138 | 0.31939 | 0.34214 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | Phaeo>Chl; depth = 60
EN720 | 20240910 | L9 | 19 | 10 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 173.8 | 113.6 | 0.5965 | 173.2035 | 113.0035 | 1.3367 | 0.67914 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L9 | 19 | 10 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 174.1 | 114.9 | 0.69617 | 173.4038 | 114.2038 | 1.3112 | 0.74656 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L9 | 19 | 10 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 158.3 | 104 | 0.7975 | 157.5025 | 103.2025 | 1.2027 | 0.65686 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L9 | 19 | 10 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 152.7 | 101.3 | 0.69617 | 152.0038 | 100.6038 | 1.1385 | 0.67427 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L9 | 19 | 10 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 23.35 | 15.92 | 0.77 | 22.58 | 15.15 | 0.16288 | 0.1073 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR129
EN720 | 20240910 | L9 | 19 | 10 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 24.84 | 16.68 | 0.77 | 24.07 | 15.91 | 0.17828 | 0.1045 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR130
EN720 | 20240910 | L9 | 19 | 10 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.345 | 3.695 | 0.77 | 4.575 | 2.925 | 0.01927 | 0.0085198 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR135
EN720 | 20240910 | L9 | 19 | 10 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.64 | 4.055 | 0.77 | 4.87 | 3.285 | 0.01851 | 0.012699 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR136
EN720 | 20240910 | L9 | 19 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 52.62 | 35.49 | 0.69617 | 51.9238 | 34.7938 | 0.37941 | 0.24752 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L9 | 19 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 50.04 | 33.51 | 0.69617 | 49.3438 | 32.8138 | 0.36612 | 0.22513 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L9 | 19 | 15 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.771 | 5.044 | 0.77 | 6.001 | 4.274 | 0.037859 | 0.038363 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR127 | Phaeo=Chl; depth = 32
EN720 | 20240910 | L9 | 19 | 15 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.635 | 4.913 | 0.77 | 5.865 | 4.143 | 0.03775 | 0.036136 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR128
EN720 | 20240910 | L9 | 19 | 15 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.237 | 2.951 | 0.77 | 3.467 | 2.181 | 0.014991 | 0.0056919 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR133
EN720 | 20240910 | L9 | 19 | 15 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 4.786 | 3.118 | 0.77 | 4.016 | 2.348 | 0.019444 | 0.0028226 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR134
EN720 | 20240910 | L9 | 19 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.9 | 11.08 | 0.68367 | 15.2163 | 10.3963 | 0.10676 | 0.080568 | 20240307 | 15-04 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L9 | 19 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.81 | 11.03 | 0.7975 | 16.0125 | 10.2325 | 0.12802 | 0.056353 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L9 | 19 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.83 | 11.22 | 0.69617 | 16.1338 | 10.5238 | 0.12426 | 0.065367 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L9 | 19 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 16.65 | 11.29 | 0.69617 | 15.9538 | 10.5938 | 0.11872 | 0.072166 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L9 | 19 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 13.8 | 9.389 | 0.7975 | 13.0025 | 8.5915 | 0.097699 | 0.057106 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L9 | 19 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 15.33 | 10.29 | 0.69617 | 14.6338 | 9.5938 | 0.11163 | 0.061235 | 20240307 | 15-07 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L9 | 19 | 20 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 3.154 | 2.482 | 0.77 | 2.384 | 1.712 | 0.014732 | 0.0158 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR125 | Phaeo>Chl; depth = 3m
EN720 | 20240910 | L9 | 19 | 20 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.871 | 2.109 | 0.77 | 2.101 | 1.339 | 0.016705 | 0.007175 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR126
EN720 | 20240910 | L9 | 19 | 20 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.134 | 1.641 | 0.77 | 1.364 | 0.871 | 0.0057575 | 0.0025175 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR131
EN720 | 20240910 | L9 | 19 | 20 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 2.509 | 1.879 | 0.77 | 1.739 | 1.109 | 0.0073308 | 0.0031672 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR132
EN720 | 20240910 | L3 | 20 | 1 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 57.27 | 39.58 | 0.7975 | 56.4725 | 38.7825 | 0.39181 | 0.30699 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 1 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 49.81 | 37.16 | 0.69617 | 49.1138 | 36.4638 | 0.28018 | 0.37684 | 20240307 | 15-08 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 5 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 113.2 | 77.39 | 0.69617 | 112.5038 | 76.6938 | 0.79315 | 0.58875 | 20240307 | 15-08 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 5 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 102.2 | 68.61 | 0.69617 | 101.5038 | 67.9138 | 0.74398 | 0.47972 | 20240307 | 15-08 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 9 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 136.3 | 91.33 | 0.7975 | 135.5025 | 90.5325 | 0.99603 | 0.63522 | 20240307 | 15-05 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 9 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 141.7 | 93.41 | 0.7975 | 140.9025 | 92.6125 | 1.0696 | 0.59916 | 20240307 | 15-06 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 9 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 116.4 | 75.49 | 0.69617 | 115.7038 | 74.7938 | 0.90611 | 0.44155 | 20240307 | 15-08 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 9 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 116.7 | 77.11 | 0.68367 | 116.0163 | 76.4263 | 0.87687 | 0.50021 | 20240307 | 15-04 | 20250219 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 9 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 35.75 | 25.32 | 0.77 | 34.98 | 24.55 | 0.22865 | 0.20918 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR141
EN720 | 20240910 | L3 | 20 | 9 |  |  | b | 294 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 34.09 | 24.44 | 0.77 | 33.32 | 23.67 | 0.21083 | 0.20986 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR142
EN720 | 20240910 | L3 | 20 | 9 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 20.2 | 15.07 | 0.77 | 19.43 | 14.3 | 0.059911 | 0.075948 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR147 | Phaeo>Chl; depth = 34m
EN720 | 20240910 | L3 | 20 | 9 |  |  | b | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.14 | 18.81 | 0.77 | 24.37 | 18.04 | 0.073925 | 0.097466 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR148 | Phaeo>Chl; depth = 34m
EN720 | 20240910 | L3 | 20 | 15 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 110.9 | 73.6 | 0.69617 | 110.2038 | 72.9038 | 0.82615 | 0.48746 | 20240307 | 15-08 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 15 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 104.9 | 69.63 | 0.69617 | 104.2038 | 68.9338 | 0.78119 | 0.46089 | 20240307 | 15-08 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 17 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 35.54 | 22.63 | 0.69617 | 34.8438 | 21.9338 | 0.28594 | 0.10927 | 20240307 | 15-08 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 17 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 37.16 | 24.19 | 0.7975 | 36.3625 | 23.3925 | 0.28727 | 0.13422 | 20240307 | 15-06 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 17 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.743 | 6.154 | 0.77 | 7.973 | 5.384 | 0.056756 | 0.039261 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR139
EN720 | 20240910 | L3 | 20 | 17 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.194 | 5.857 | 0.77 | 7.424 | 5.087 | 0.051232 | 0.039489 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR140
EN720 | 20240910 | L3 | 20 | 17 |  |  | a | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 8.084 | 5.808 | 0.77 | 7.314 | 5.038 | 0.026532 | 0.021245 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR145
EN720 | 20240910 | L3 | 20 | 17 |  |  | b | 551 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.792 | 4.227 | 0.77 | 5.022 | 3.457 | 0.018244 | 0.01454 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR146
EN720 | 20240910 | L3 | 20 | 20 |  |  | a | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 27.35 | 17.94 | 0.69617 | 26.6538 | 17.2438 | 0.20842 | 0.10229 | 20240307 | 15-08 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 20 |  |  | b | 290 | 0 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 25.47 | 16.78 | 0.7975 | 24.6725 | 15.9825 | 0.19247 | 0.095505 | 20240307 | 15-06 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 20 |  |  | a | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 21.68 | 13.69 | 0.7975 | 20.8825 | 12.8925 | 0.17697 | 0.055333 | 20240307 | 15-06 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 20 |  |  | b | 290 | 10 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 22.39 | 13.79 | 0.69617 | 21.6938 | 13.0938 | 0.19048 | 0.04545 | 20240307 | 15-08 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240910 | L3 | 20 | 20 |  |  | a | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 7.471 | 5.415 | 0.77 | 6.701 | 4.645 | 0.045072 | 0.037767 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR137
EN720 | 20240910 | L3 | 20 | 20 |  |  | b | 293 | 5 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.644 | 4.058 | 0.77 | 4.874 | 3.288 | 0.034768 | 0.023869 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR138
EN720 | 20240910 | L3 | 20 | 20 |  |  | a | 550 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 6.127 | 4.556 | 0.77 | 5.357 | 3.786 | 0.018347 | 0.017622 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 3 |  | TAR143 | Phaeo=Chl; depth = 4
EN720 | 20240910 | L3 | 20 | 20 |  |  | b | 552 | 20 | 5 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 5.311 | 3.781 | 0.77 | 4.541 | 3.011 | 0.017803 | 0.010699 | 20240307 | 15-02 | 20250114 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1 |  | TAR144
EN720 | 20240911 | MVCO | 24 | 2 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 204.1 | 131.1 | 0.7975 | 203.3025 | 130.3025 | 3.2337 | 1.4619 | 20240307 | 15-06 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240911 | MVCO | 24 | 2 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 195.3 | 126.3 | 0.69617 | 194.6038 | 125.6038 | 3.0565 | 1.4698 | 20240307 | 15-08 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240911 | MVCO | 24 | 2 |  |  | a | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN720 | 20240911 | MVCO | 24 | 2 |  |  | b | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN720 | 20240911 | MVCO | 24 | 4 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 301.2 | 193 | 0.7975 | 300.4025 | 192.2025 | 4.793 | 2.1333 | 20240307 | 15-06 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240911 | MVCO | 24 | 4 |  |  | b | 290 | 0 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN720 | 20240911 | MVCO | 24 | 4 |  |  | a | 290 | 10 | [10 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
EN720 | 20240911 | MVCO | 24 | 4 |  |  | b | 290 | 10 | 15 | 1 | 0 | 1 | Cal_052825 | 1.8034 | 0.57373 | 216.1 | 135.5 | 0.5965 | 215.5035 | 134.9035 | 5.3691 | 1.8505 | 20250529 | 15-26 | 20250612 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240911 | MVCO | 24 | 7 |  |  | a | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 307.3 | 193 | 0.7975 | 306.5025 | 192.2025 | 5.0632 | 1.8631 | 20240307 | 15-06 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240911 | MVCO | 24 | 7 |  |  | b | 290 | 0 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 318.3 | 200.2 | 0.69617 | 317.6038 | 199.5038 | 5.2316 | 1.9579 | 20240307 | 15-08 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240911 | MVCO | 24 | 7 |  |  | a | 290 | 10 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 162.7 | 106 | 0.7975 | 161.9025 | 105.2025 | 2.5117 | 1.2795 | 20240307 | 15-06 | 20250225 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1
EN720 | 20240911 | MVCO | 24 | 7 |  |  | b | 290 | 10 | 10 | 1 | 0 | 1 | Cal_030724 | 1.8135 | 0.57627 | 150.1 | 100 | 0.69617 | 149.4038 | 99.3038 | 2.2193 | 1.3593 | 20240307 | 15-08 | 20250227 |  |  |  | Turner Designs Aquafluor Handheld Serial Number 800446 | 1

## missing
AR24B | 20171027 | L1 | 1 | 4 | a | 282 | 0
AR24B | 20171027 | L1 | 1 | 4 | b | 282 | 0
AR24B | 20171027 | L1 | 1 | 5 | a | 282 | 0
AR24B | 20171027 | L1 | 1 | 5 | b | 282 | 0

## cruise_proj_type
Number | Project
1 | LTER
2 | OOI
3 | JPstudentcruise
4 | LTER_OOI
5 | LTER_JPstudentcruise

## Cal_052925
Calibration Date | Cruise_Site | Fluorometer | Personell | CALIBRATION | Blank_Acetone | AvgBlank | Dilution | ug_l | Flurometer_Before_Acid | Fluorometer_After_Acid | Fluorometer_Before_Acid-AvgBlank | Fluorometer_After_Acid-AvgBlank | Acid_Ratio__Before-AvgBlank/After-AvgBlank | Fd__Chl/Before-AvgBlank | tau | Fd
2025-05-28 | mvco | Turner Designs Aquafluor Handheld Serial Number 800446 | taylor | Cal_052825 | 0.552 | 0.535 | full(on spec) | 10917.4746
20250529 |  |  |  |  | 0.51 |  | 20 | 545.8737 | 952 | 523.1 | 951.465 | 522.565 | 1.8208 | 0.57372 | 1.8034 | 0.57373 | 1.8034 | 1.9
[5 empty] | 0.356 |  | 40 | 272.9369 | 471.5 | 271 | 470.965 | 270.465 | 1.7413 | 0.57953 |  |  | 1.8034 | 1.9
[5 empty] | 0.72 |  | 80 | 136.4684 | 245.7 | 136.9 | 245.165 | 136.365 | 1.7979 | 0.55664 | old averages |  | 1.8034 | 1.9
[5 empty] | 0.555 |  | 160 | 68.2342 | 122.2 | 68.85 | 121.665 | 68.315 | 1.7809 | 0.56084 | 1.7608 | 0.56365 | 1.8034 | 1.9
[5 empty] | 0.517 |  | 320 | 34.1171 | 61.41 | 34.15 | 60.875 | 33.615 | 1.8109 | 0.56045 |  |  | 1.8034 | 1.9
[7 empty] | 640 | 17.0586 | 30.95 | 17.32 | 30.415 | 16.785 | 1.812 | 0.56086 |  |  | 1.8034 | 1.9
[7 empty] | 1280 | 8.5293 | 15.31 | 9.134 | 14.775 | 8.599 | 1.7182 | 0.57728 |  |  | 1.8034 | 1.9
[7 empty] | 2560 | 4.2646 | 7.925 | 4.568 | 7.39 | 4.033 | 1.8324 | 0.57708 |  |  | 1.8034 | 1.9
[7 empty] | 5120 | 2.1323 | 4.405 | 2.767 | 3.87 | 2.232 | 1.7339 | 0.55099 |  |  | 1.8034 | 1.9
[7 empty] | 10240 | 1.0662 | 2.337 | 1.662 | 1.802 | 1.127 | 1.5989 | 0.59165 |  |  | 1.8034 | 1.9
[7 empty] | 20480 | 0.53308 | 1.578 | 1.141 | 1.043 | 0.606 | 1.7211 | 0.5111 |  |  | 1.8034 | 1.9
[7 empty] | 557.226 | spinach_5percent | [7 empty] | *** The 1/20 dilution was done using a 5ml volumetric flask of chl-a stock into 100ml volumetric diluted with 90% acetone
0.05104
19.5925
[7 empty] | 559.4958 | spinach_5percent
0.051248
19.5131

## Cal_030724
Calibration Date | Cruise_Site | Fluorometer | Personell | CALIBRATION | Blank_Acetone | AvgBlank | Dilution | ug_l | Flurometer_Before_Acid | Fluorometer_After_Acid | Fluorometer_Before_Acid-AvgBlank | Fluorometer_After_Acid-AvgBlank | Acid_Ratio__Before-AvgBlank/After-AvgBlank | Fd__Chl/Before-AvgBlank | tau | Fd
2024-03-07 | mvco | Turner Designs Aquafluor Handheld Serial Number 800446 | taylor | Cal_030724 | 0.645 | 0.67067 | full(on spec) | 11920.908
20240307 |  |  |  |  | 0.708 |  | 20 | 596.0454 | 1021 | 560.5 | 1020.3293 | 559.8293 | 1.8226 | 0.58417 | 1.8135 | 0.57627 | 1.8135 | 1.9
[5 empty] | 0.546 |  | 40 | 298.0227 | 535.8 | 299.8 | 535.1293 | 299.1293 | 1.789 | 0.55692 |  |  | 1.8135 | 1.9
[5 empty] | 0.89 |  | 80 | 149.0113 | 271.3 | 151.5 | 270.6293 | 150.8293 | 1.7943 | 0.55061 | old averages |  | 1.8135 | 1.9
[5 empty] | 0.613 |  | 160 | 74.5057 | 137.2 | 77.14 | 136.5293 | 76.4693 | 1.7854 | 0.54571 | 1.8098 | 0.54646 | 1.8135 | 1.9
[5 empty] | 0.622 |  | 320 | 37.2528 | 69.99 | 39.17 | 69.3193 | 38.4993 | 1.8005 | 0.53741 |  |  | 1.8135 | 1.9
[7 empty] | 640 | 18.6264 | 34.94 | 19.98 | 34.2693 | 19.3093 | 1.7748 | 0.54353 |  |  | 1.8135 | 1.9
[7 empty] | 1280 | 9.3132 | 17.8 | 10.27 | 17.1293 | 9.5993 | 1.7844 | 0.5437 |  |  | 1.8135 | 1.9
[7 empty] | 2560 | 4.6566 | 9.253 | 5.76 | 8.5823 | 5.0893 | 1.6863 | 0.54258 |  |  | 1.8135 | 1.9
[7 empty] | 5120 | 2.3283 | 4.913 | 3.058 | 4.2423 | 2.3873 | 1.777 | 0.54883 |  |  | 1.8135 | 1.9
[7 empty] | 10240 | 1.1642 | 2.695 | 1.836 | 2.0243 | 1.1653 | 1.7371 | 0.57508 |  |  | 1.8135 | 1.9
[7 empty] | 20480 | 0.58208 | 1.877 | 1.23 | 1.2063 | 0.55933 | 2.1567 | 0.48252 |  |  | 1.8135 | 1.9
[7 empty] | 619.2084 | spinach_5percent | [7 empty] | *** The 1/20 dilution was done using a 5ml volumetric flask of chl-a stock into 100ml volumetric diluted with 90% acetone
0.051943
19.2519
[7 empty] | 599.9544 | spinach_5percent
0.050328
19.8697

## Cal_051922a
Calibration Date | Cruise_Site | Fluorometer | Personell | CALIBRATION | Blank_Acetone | AvgBlank | Dilution | ug_l | Flurometer_Before_Acid | Fluorometer_After_Acid | Fluorometer_Before_Acid-AvgBlank | Fluorometer_After_Acid-AvgBlank | Acid_Ratio__Before-AvgBlank/After-AvgBlank | Fd__Chl/Before-AvgBlank | tau | Fd
2022-05-19 |  | Turner Designs Aquafluor Handheld Serial Number 800446 | taylor | Cal_051922a | 0.59 | 0.54817 | full(on spec) | 12937.9605
20220519 |  |  |  |  | 0.501 |  | 20 | 646.898 | 1145 | 616.3 | 1144.4518 | 615.7518 | 1.8586 | 0.56525 | 1.8382 | 0.56187 | 1.8382 | 1.9
[5 empty] | 0.602 |  | 40 | 323.449 | 583 | 327 | 582.4518 | 326.4518 | 1.7842 | 0.55532 |  |  | 1.8382 | 1.9
[5 empty] | 0.356 |  | 80 | 161.7245 | 298 | 166.8 | 297.4518 | 166.2518 | 1.7892 | 0.5437 | old averages |  | 1.8382 | 1.9
[5 empty] | 0.598 |  | 160 | 80.8623 | 149 | 85.09 | 148.4518 | 84.5418 | 1.756 | 0.5447 | 1.8249 | 0.55957 | 1.8382 | 1.9
[5 empty] | 0.642 |  | 320 | 40.4311 | 76.53 | 42.11 | 75.9818 | 41.5618 | 1.8282 | 0.53212 |  |  | 1.8382 | 1.9
[7 empty] | 640 | 20.2156 | 38.1 | 21.23 | 37.5518 | 20.6818 | 1.8157 | 0.53834 |  |  | 1.8382 | 1.9
[7 empty] | 1280 | 10.1078 | 19.2 | 10.49 | 18.6518 | 9.9418 | 1.8761 | 0.54192 |  |  | 1.8382 | 1.9
[7 empty] | 2560 | 5.0539 | 9.529 | 5.53 | 8.9808 | 4.9818 | 1.8027 | 0.56274 |  |  | 1.8382 | 1.9
[7 empty] | 5120 | 2.5269 | 5.299 | 3.104 | 4.7508 | 2.5558 | 1.8588 | 0.5319 |  |  | 1.8382 | 1.9
[7 empty] | 10240 | 1.2635 | 2.88 | 1.572 | 2.3318 | 1.0238 | 2.2776 | 0.54184 |  |  | 1.8382 | 1.9
[7 empty] | 20480 | 0.63174 | 1.454 | 1.183 | 0.90583 | 0.63483 | 1.4269 | 0.69741 |  |  | 1.8382 | 1.9
[7 empty] | 651.5342 | spinach_5percent | [7 empty] | *** The 1/20 dilution was done using a 5ml volumetric flask of chl-a stock into 100ml volumetric diluted with 90% acetone
0.050358
19.8577

## Cal_111320b_12mm
Calibration Date | Cruise_Site | Fluorometer | Personell | CALIBRATION | Blank_Acetone | AvgBlank | Dilution | ug_l | Flurometer_Before_Acid | Fluorometer_After_Acid | Fluorometer_Before_Acid-AvgBlank | Fluorometer_After_Acid-AvgBlank | Acid_Ratio__Before-AvgBlank/After-AvgBlank | Fd__Chl/Before-AvgBlank | tau | Fd
2020-11-13 | mvco | Turner Designs Aquafluor Handheld 800446 | taylor | Cal_111320b_12mm | 0.659 | 0.507 | full(on spec) | 10942.7512
20201113 |  |  |  |  | 0.468 |  | 20 | 547.1376 | 991.7 | 539.2 | 991.193 | 538.693 | 1.84 | 0.552 | 1.8343 | 0.5442 | 1.8343 | 1.9
[5 empty] | 0.543 |  | 40 | 273.5688 | 520.9 | 286.4 | 520.393 | 285.893 | 1.8202 | 0.5257 |  |  | 1.8343 | 1.9
[5 empty] | 0.276 |  | 80 | 136.7844 | 264.2 | 145.4 | 263.693 | 144.893 | 1.8199 | 0.51873 | old averages |  | 1.8343 | 1.9
[5 empty] | 0.498 |  | 160 | 68.3922 | 135.4 | 75.03 | 134.893 | 74.523 | 1.8101 | 0.50701 | 1.719 | 0.55187 | 1.8343 | 1.9
[5 empty] | 0.598 |  | 320 | 34.1961 | 68.3 | 37.96 | 67.793 | 37.453 | 1.8101 | 0.50442 |  |  | 1.8343 | 1.9
[7 empty] | 640 | 17.098 | 33.08 | 18.72 | 32.573 | 18.213 | 1.7884 | 0.52491 |  |  | 1.8343 | 1.9
[7 empty] | 1280 | 8.549 | 16.95 | 9.641 | 16.443 | 9.134 | 1.8002 | 0.51992 |  |  | 1.8343 | 1.9
[7 empty] | 2560 | 4.2745 | 8.432 | 5.149 | 7.925 | 4.642 | 1.7072 | 0.53937 |  |  | 1.8343 | 1.9
[7 empty] | 5120 | 2.1373 | 4.268 | 2.642 | 3.761 | 2.135 | 1.7616 | 0.56827 |  |  | 1.8343 | 1.9
[7 empty] | 10240 | 1.0686 | 2.355 | 1.724 | 1.848 | 1.217 | 1.5185 | 0.57826 |  |  | 1.8343 | 1.9
[7 empty] | 20480 | 0.53431 | 1.237 | 1.099 | 0.73 | 0.592 | 1.2331 | 0.73194 |  |  | 1.8343 | 1.9
[6 empty] | 2020-11-13 | 554.169 | spinach_5percent | [7 empty] | *** The 1/20 dilution was done using a 5ml volumetric flask of chl-a stock into 100ml volumetric diluted with 90% acetone
0.050643
19.7462
[6 empty] | 2018-01-26 |  | spinach_5percent2
0
#DIV/0!

## Cal_111320
Calibration Date | Cruise_Site | Fluorometer | Personell | CALIBRATION | Blank_Acetone | AvgBlank | Dilution | ug_l | Flurometer_Before_Acid | Fluorometer_After_Acid | Fluorometer_Before_Acid-AvgBlank | Fluorometer_After_Acid-AvgBlank | Acid_Ratio__Before-AvgBlank/After-AvgBlank | Fd__Chl/Before-AvgBlank | tau | Fd
2020-11-13 | mvco | Turner Designs Aquafluor Handheld 800446 | taylor | Cal_111320_12mm | 0.58 | 0.59871 | full(on spec) | 10218.7065
[5 empty] | 0.612 |  | 20 | 510.9353 | 935.8 | 494.5 | 935.2013 | 493.9013 | 1.8935 | 0.54634 | 1.8663 | 0.53886 | 1.8663 | 1.9
[5 empty] | 0.584 |  | 40 | 255.4677 | 490.6 | 272.2 | 490.0013 | 271.6013 | 1.8041 | 0.52136 |  |  | 1.8663 | 1.9
[5 empty] | 0.562 |  | 80 | 127.7338 | 249.8 | 140.6 | 249.2013 | 140.0013 | 1.78 | 0.51257 | old averages |  | 1.8663 | 1.9
[5 empty] | 0.359 |  | 160 | 63.8669 | 126.6 | 70.54 | 126.0013 | 69.9413 | 1.8015 | 0.50688 | 1.8207 | 0.53587 | 1.8663 | 1.9
[5 empty] | 0.775 |  | 320 | 31.9335 | 64.16 | 35.03 | 63.5613 | 34.4313 | 1.846 | 0.5024 |  |  | 1.8663 | 1.9
[5 empty] | 0.719 |  | 640 | 15.9667 | 32.45 | 18.05 | 31.8513 | 17.4513 | 1.8252 | 0.50129 |  |  | 1.8663 | 1.9
[7 empty] | 1280 | 7.9834 | 16.52 | 9.203 | 15.9213 | 8.6043 | 1.8504 | 0.50143 |  |  | 1.8663 | 1.9
[7 empty] | 2560 | 3.9917 | 8.466 | 4.757 | 7.8673 | 4.1583 | 1.892 | 0.50738 |  |  | 1.8663 | 1.9
[7 empty] | 5120 | 1.9958 | 4.617 | 2.565 | 4.0183 | 1.9663 | 2.0436 | 0.49669 |  |  | 1.8663 | 1.9
[7 empty] | 10240 | 0.99792 | 2.559 | 1.812 | 1.9603 | 1.2133 | 1.6157 | 0.50907 |  |  | 1.8663 | 1.9
[7 empty] | 20480 | 0.49896 | 1.231 | 0.976 | 0.63229 | 0.37729 | 1.6759 | 0.78914 |  |  | 1.8663 | 1.9
[6 empty] | 2020-11-13 | 508.2126 | spinach_5percent | [7 empty] | *** The 1/20 dilution was done using a 5ml volumetric flask of chl-a stock into 100ml volumetric diluted with 90% acetone
0.049734
20.1071
[6 empty] | 2018-01-26 |  | spinach_5percent2
0
#DIV/0!

## Cal_110119
Calibration name = Cal_110119 | Cal_date | Calibration name | tau | Fd
| 20181101 | Cal_110119 | 1.8682 | 0.5122
For Tau, average 2018 and 2020 used cablibration taus
2019 Tau = 1.8682
Plot average solid standard of each calibration day vs Fd of that calibration. Linear line fit to all calibrations that were used for this Aquafluor
Take the grouping of 2019 data that needs this calibration and average the solid standard for all of the days samples were analyzed
Caliculate the Fd from putting the averaged standard value into the line fit.
2019 average solid standard = 27.6778
line fit is y = -106.0398*x + 81.9895
2019 Fd = 0.5122
Apply this cal to LabNB 9-37 to 9-100 and 10-01 to 10-45
This cal will apply to samples run between 11/1/2018 and 12/22/2019

## Cal_012618_12mm
Calibration Date | Cruise_Site | Fluorometer | Personell | CALIBRATION | Blank_Acetone | AvgBlank | Dilution | ug_l | Flurometer_Before_Acid | Fluorometer_After_Acid | Fluorometer_Before_Acid-AvgBlank | Fluorometer_After_Acid-AvgBlank | Acid_Ratio__Before-AvgBlank/After-AvgBlank | Fd__Chl/Before-AvgBlank | tau | Fd
2018-01-26 | mvco | Turner Designs Aquafluor Handheld 800446 | taylor | Cal_012618_12mm | 0.635 | 0.67883 | full(on spec) | 11439.8312
[5 empty] | 0.891 |  | 20 | 571.9916 | 1166 | 609.5 | 1165.3212 | 608.8212 | 1.9141 | 0.49084 | 1.902 | 0.48424 | 1.902 | 1.9
[5 empty] | 0.532 |  | 40 | 285.9958 | 610.3 | 326.5 | 609.6212 | 325.8212 | 1.871 | 0.46914 |  |  | 1.902 | 1.9
[5 empty] | 0.768 |  | 80 | 142.9979 | 312 | 167.3 | 311.3212 | 166.6212 | 1.8684 | 0.45933 | old averages |  | 1.902 | 1.9
[5 empty] | 0.561 |  | 160 | 71.4989 | 157.4 | 84.59 | 156.7212 | 83.9112 | 1.8677 | 0.45622 | 1.8812 | 0.45619 | 1.902 | 1.9
[5 empty] | 0.686 |  | 320 | 35.7495 | 79.69 | 41.8 | 79.0112 | 41.1212 | 1.9214 | 0.45246 |  |  | 1.902 | 1.9
[7 empty] | 640 | 17.8747 | 39.86 | 22.04 | 39.1812 | 21.3612 | 1.8342 | 0.45621 |  |  | 1.902 | 1.9
[7 empty] | 1280 | 8.9374 | 20.43 | 11.28 | 19.7512 | 10.6012 | 1.8631 | 0.4525 |  |  | 1.902 | 1.9
[7 empty] | 2560 | 4.4687 | 10.53 | 5.816 | 9.8512 | 5.1372 | 1.9176 | 0.45362 |  |  | 1.902 | 1.9
[7 empty] | 5120 | 2.2343 | 5.945 | 3.38 | 5.2662 | 2.7012 | 1.9496 | 0.42428 |  |  | 1.902 | 1.9
[7 empty] | 10240 | 1.1172 | 3.209 | 1.994 | 2.5302 | 1.3152 | 1.9238 | 0.44154 |  |  | 1.902 | 1.9
[7 empty] | 20480 | 0.55859 | 1.888 | 1.365 | 1.2092 | 0.68617 | 1.7622 | 0.46196 |  |  | 1.902 | 1.9
[6 empty] | 2018-01-26 | 586.1982 | spinach_5percent | [7 empty] | *** The 1/20 dilution was done using a 5ml volumetric flask of chl-a stock into 100ml volumetric diluted with 90% acetone
0.051242
19.5153
[6 empty] | 2018-01-26 | 595.9393 | spinach_5percent2
0.052093
19.1963